## Final Report
Final Report
Job 2021-0048
|  | Isotope(s) requested: | δ13C, δ15N
|  | Instrument Used: | Carlo Erba 1110 Elemental Analyzer coupled to a Thermo Delta V IRMS
|  | Analytical Code: | 010 (combustion/reduction organics)
|  | Units: | δ13C and δ15N values are reported w.r.t. VPDB and AIR-N2 respectively, expressed in per mil.
|  | Principal Investigator: | Nanette Nelson
|  | Job submission contact: | Tyler Tappenbeck
|  | Sample Material(s): | Fish tissue and Zooplankton
|  | Project: | FLBS EPA MeHg and isotopes
|  | Date Submited: | 2021-05-06
|  | Number of unknown samples analyzed: | 240
|  | Quality Assurance Reference Material 1: | 36-UWSIF-Glutamic 1
|  | Quality Assurance Reference Material 2: | 39-UWSIF-Glutamic 2
|  | Quality Control Reference Material 3: | 01-UWSIF-Liver
|  | Number of reference samples analyzed: | 94
Quality Assurance Data |  |  |  |  | Quality Control Data
Reference Material 1 | Known |  | Known |  | Reference Material 3 | Known | Normalized | Known | Normalized
36-UWSIF-glutamic acid | δ13C VPDB |  | δ15N AIR-N2 |  | 01-UWSIF-Liver | δ13C VPDB | δ13C VPDB | δ15N AIR-N2 | δ15N AIR-N2
average (N=34) | -28.28 |  | -4.62 |  | average (N=30) | -17.75 | -17.7778 | 6.81 | 6.7422
standard uncertainty | 0.10131 |  | 0.035905 |  | standard uncertainty |  | 0.080795 |  | 0.036087
Reference Material 2 | Known |  | Known |  | long-term | δ13C VPDB | -18.05 | δ15N AIR-N2 | 6.41
39-UWSIF-glutamic acid | δ13C VPDB |  | δ15N AIR-N2 |  | acceptable range | 2 σ = 0.3 | -17.45 | 2 σ = 0.4 | 7.21
average (N=30) | 24.36 |  | 27.89 |  | Weight Percent | Known | Measured | Known | Measured
standard uncertainty | 0.085722 |  | 0.05799 |  |  | Wt% C | Wt% C | Wt% N | Wt% N
[5 empty] | average (N=30) | 49.64 | 50.0348 | 12.88 | 12.9341
[5 empty] | standard uncertainty |  | 0.38647 |  | 0.11584
[5 empty] | percent error (%) |  | 0.94433 |  | 0.85812
Record Keeping |  |  |  |  | Quality Assurance Approval
Date Reported: | 2021-05-25 |  |  |  | Reviewer: | Craig Cook
Date Invoiced: | 2021-05-25 |  |  |  | Title: | Facility Director
Initial: | cjm |  |  |  | Date Reviewed: | 2021-05-24
Comments: |  |  |  |  | Comments: | Data meet all QAQC criteria in the SOP.
Analytical Comments:
*Sample weight percents are calculated using the sample weights reported by the user and are dependent upon the accuracy of these values.
Quality Control Color Legend
Yellow | Peak amplitudes too low for reliable results. Use with extreme caution or rerun the sample.
Green | Weight percent values wrong. Use with caution.
Pink | Sample missing from tray.
Blue | Sample lost during the analysis. Reload and rerun the sample.
Olive | Possible sample ID problem. Check loading documents.
SIF ID | Sample ID | δ15N | δ13C | Wt% N* | Wt% C* | C:N ratio | Comments
20210048.001 | 501 | 11.7878 | -30.974 | 14.0808 | 47.9424 | 3.4048
20210048.002 | 503 | 10.9693 | -31.7902 | 14.082 | 49.4339 | 3.5104
20210048.002 | Replicate A11 | 11.0043 | -31.7327 | 14.2457 | 49.695 | 3.4884
20210048.003 | 505 | 8.4082 | -31.3503 | 13.9989 | 49.2138 | 3.5155
20210048.004 | 507 | 10.7258 | -31.7367 | 13.3339 | 50.086 | 3.7563
20210048.005 | 509 | 10.3685 | -31.9789 | 14.2681 | 48.839 | 3.4229
20210048.006 | 511 | 7.1506 | -31.0103 | 14.5578 | 48.1578 | 3.308
20210048.007 | 513 | 8.3553 | -31.9758 | 11.8983 | 52.4939 | 4.4119
20210048.008 | 515 | 10.5372 | -31.2484 | 14.3434 | 47.5313 | 3.3138
20210048.009 | 517 | 10.644 | -31.6984 | 13.7 | 49.7421 | 3.6308
20210048.010 | 519 | 7.1556 | -32.0727 | 13.7089 | 49.1462 | 3.585
20210048.011 | 521 | 11.0981 | -31.9829 | 13.4193 | 49.7891 | 3.7102
20210048.012 | 523 | 10.9035 | -31.5249 | 13.87 | 49.208 | 3.5478
20210048.013 | 525 | 10.2846 | -30.1457 | 14.7831 | 47.2054 | 3.1932
20210048.014 | 527 | 11.4784 | -32.1585 | 12.9262 | 49.6798 | 3.8434
20210048.015 | 529 | 7.3532 | -29.8733 | 14.3938 | 48.7743 | 3.3886
20210048.016 | 531 | 11.2568 | -31.2212 | 14.3933 | 49.0016 | 3.4045
20210048.017 | 533 | 10.6869 | -29.2881 | 13.431 | 51.2161 | 3.8133
20210048.018 | 535 | 7.6736 | -31.4129 | 14.7386 | 48.7378 | 3.3068
20210048.019 | 537 | 8.9362 | -30.8196 | 14.552 | 48.3745 | 3.3243
20210048.020 | 541 | 10.0661 | -27.9816 | 13.6453 | 49.5063 | 3.6281
20210048.021 | 543 | 8.1906 | -29.726 | 14.5024 | 48.2931 | 3.33
20210048.022 | 545 | 7.0468 | -31.6258 | 13.3568 | 50.0971 | 3.7507
20210048.022 | Replicate C8 | 7.0468 | -31.6056 | 13.3526 | 50.7997 | 3.8045
20210048.023 | 547 | 10.1908 | -29.2185 | 14.6859 | 49.0475 | 3.3398
20210048.024 | 549 | 10.7098 | -29.7129 | 14.7625 | 46.1032 | 3.123
20210048.025 | 551 | 11.2897 | -32.5257 | 12.6304 | 52.3387 | 4.1439
20210048.026 | 553 | 11.0292 | -31.31 | 14.6545 | 48.5703 | 3.3144
20210048.027 | 555 | 10.7238 | -31.6954 | 13.8575 | 50.5809 | 3.6501
20210048.028 | 557 | 10.7128 | -31.7902 | 13.9678 | 50.5731 | 3.6207
20210048.029 | 559 | 10.7488 | -30.3818 | 14.5706 | 49.0964 | 3.3696
20210048.030 | 561 | 7.6167 | -31.1294 | 14.9226 | 47.5889 | 3.189
20210048.031 | 563 | 11.169 | -30.1296 | 14.8483 | 48.2539 | 3.2498
20210048.032 | 565 | 7.7674 | -31.2272 | 14.9346 | 48.8057 | 3.268
20210048.033 | 567 | 11.4075 | -31.1274 | 13.7048 | 50.9533 | 3.7179
20210048.034 | 569 | 7.7944 | -30.8983 | 14.7879 | 46.8087 | 3.1653
20210048.035 | 571 | 7.5778 | -31.196 | 14.7386 | 47.2049 | 3.2028
20210048.036 | 573 | 10.5252 | -30.3253 | 14.3844 | 49.0499 | 3.4099
20210048.037 | 575 | 8.4771 | -31.4058 | 13.9285 | 48.9967 | 3.5177
20210048.038 | 577 | 8.3164 | -30.8257 | 14.704 | 47.7977 | 3.2507
20210048.039 | 579 | 10.6809 | -30.9669 | 10.6961 | 55.5989 | 5.198
20210048.040 | 581 | 9.551 | -31.2383 | 13.9399 | 48.9235 | 3.5096
20210048.041 | 583 | 11.146 | -33.2521 | 11.773 | 54.319 | 4.6139
20210048.042 | 585 | 8.4392 | -30.862 | 14.4545 | 47.8951 | 3.3135
20210048.042 | Replicate E11 | 8.4491 | -30.7783 | 14.4219 | 47.8063 | 3.3148
20210048.043 | 587 | 11.0302 | -31.4714 | 8.8494 | 56.7517 | 6.4131
20210048.044 | 589 | 10.5761 | -31.8871 | 13.5258 | 50.5618 | 3.7382
20210048.045 | 591 | 11.4584 | -32.0192 | 12.9334 | 49.6396 | 3.8381
20210048.046 | 593 | 10.9294 | -29.958 | 14.7289 | 48.7632 | 3.3107
20210048.047 | 595 | 7.2085 | -31.9345 | 14.0177 | 49.6868 | 3.5446
20210048.048 | 597 | 7.3492 | -30.3535 | 14.6869 | 48.0233 | 3.2698
20210048.049 | 599 | 8.5829 | -31.0709 | 14.3735 | 48.5247 | 3.376
20210048.050 | 601 | 11.0811 | -31.4502 | 14.0076 | 49.0704 | 3.5031
20210048.051 | 603 | 9.1598 | -31.2696 | 13.9009 | 49.1306 | 3.5344
20210048.052 | 605 | 8.2645 | -31.0275 | 14.4587 | 48.2495 | 3.3371
20210048.053 | 607 | 10.596 | -30.0438 | 14.4838 | 47.6062 | 3.2869
20210048.054 | 609 | 10.605 | -31.7539 | 12.7149 | 53.3003 | 4.1919
20210048.055 | 611 | 10.8975 | -31.654 | 11.4487 | 53.3883 | 4.6633
20210048.056 | 613 | 7.5658 | -31.1425 | 14.6584 | 47.9841 | 3.2735
20210048.057 | 615 | 7.5479 | -32.5449 | 12.4959 | 50.9964 | 4.081
20210048.058 | 617 | 8.8055 | -31.6036 | 13.3194 | 49.8357 | 3.7416
20210048.059 | 619 | 11.9066 | -30.4756 | 14.0637 | 47.241 | 3.3591
20210048.060 | 623 | 10.8306 | -30.8882 | 14.5562 | 48.5349 | 3.3343
20210048.061 | 625 | 10.9903 | -30.974 | 14.3776 | 48.2924 | 3.3589
20210048.062 | 627 | 10.7038 | -31.5501 | 13.5573 | 48.4113 | 3.5709
20210048.063 | 629 | 8.4002 | -30.7541 | 14.5186 | 47.5388 | 3.2743
20210048.064 | 631 | 10.8436 | -30.8771 | 12.8806 | 50.1564 | 3.8939
20210048.065 | 635 | 9.5091 | -30.5926 | 14.4124 | 49.1291 | 3.4088
20210048.066 | 637 | 11.1171 | -31.1869 | 14.2609 | 48.8186 | 3.4233
20210048.067 | 639 | 11.3806 | -31.0658 | 13.8572 | 48.4888 | 3.4992
20210048.068 | 641 | 9.1997 | -30.5281 | 14.6138 | 47.157 | 3.2269
20210048.068 | Replicate H8 | 9.2277 | -30.4928 | 14.7989 | 47.6548 | 3.2202
20210048.069 | 643 | 10.7442 | -30.099 | 14.7932 | 49.2795 | 3.3312
20210048.070 | 645 | 10.1068 | -30.094 | 15.1732 | 48.9549 | 3.2264
20210048.070 | Replicate A11 | 10.1268 | -30.0657 | 15.1423 | 48.9229 | 3.2309
20210048.071 | 651 | 7.0176 | -30.1617 | 15.0998 | 47.8608 | 3.1696
20210048.072 | 653 | 6.845 | -30.7487 | 14.6697 | 48.9291 | 3.3354
20210048.073 | 655 | 10.7652 | -29.8444 | 14.7152 | 49.6143 | 3.3716
20210048.074 | 657 | 6.8759 | -30.2728 | 14.6388 | 49.0862 | 3.3531
20210048.075 | 659 | 10.1028 | -27.7317 | 13.9163 | 50.7763 | 3.6487
20210048.076 | 661 | 10.3492 | -29.417 | 14.4542 | 47.9373 | 3.3165
20210048.077 | 663 | 11.0894 | -30.7952 | 15.1015 | 48.2594 | 3.1957
20210048.078 | 665 | 11.0076 | -32.9363 | 11.3278 | 50.2376 | 4.4349
20210048.079 | 667 | 8.2295 | -31.2216 | 14.7922 | 47.7844 | 3.2304
20210048.080 | 669 | 10.5866 | -30.0697 | 15.0876 | 47.9493 | 3.1781
20210048.081 | 671 | 8.0779 | -31.1165 | 15.131 | 47.909 | 3.1663
20210048.082 | 673 | 10.4639 | -30.9064 | 14.8596 | 47.5687 | 3.2012
20210048.083 | 703 | 8.2465 | -32.23 | 12.8267 | 50.8828 | 3.9669
20210048.084 | 705 | 11.6131 | -31.1984 | 13.3644 | 50.2173 | 3.7576
20210048.085 | 709 | 7.1323 | -30.5436 | 14.928 | 48.1847 | 3.2278
20210048.086 | 713 | 10.5387 | -29.5241 | 13.1582 | 50.1952 | 3.8148
20210048.087 | 715 | 8.8021 | -34.6024 | 11.7339 | 54.2136 | 4.6202
20210048.088 | 719 | 10.6445 | -31.1064 | 14.3995 | 48.3685 | 3.359
20210048.089 | 721 | 11.1602 | -31.4318 | 14.3432 | 48.6707 | 3.3933
20210048.090 | 723 | 7.2181 | -31.3408 | 14.9333 | 47.6555 | 3.1912
20210048.090 | Replicate C8 | 7.3029 | -31.3944 | 15.1184 | 47.9481 | 3.1715
20210048.091 | 725 | 10.6445 | -31.8814 | 13.9065 | 49.9696 | 3.5933
20210048.092 | 727 | 8.4201 | -31.3297 | 14.6505 | 48.413 | 3.3045
20210048.093 | 729 | 11.8564 | -30.6426 | 14.3059 | 49.1498 | 3.4356
20210048.094 | 731 | 10.9876 | -32.4159 | 13.0812 | 51.6628 | 3.9494
20210048.095 | 733 | 10.9876 | -30.3415 | 15.1021 | 47.654 | 3.1555
20210048.096 | 735 | 8.1727 | -31.4914 | 14.1989 | 49.1359 | 3.4606
20210048.097 | 739 | 9.7118 | -30.5264 | 14.522 | 47.1583 | 3.2474
20210048.098 | 741 | 11.2819 | -30.3304 | 14.2719 | 48.2899 | 3.3836
20210048.099 | 747 | 8.6385 | -31.2529 | 14.3789 | 49.1898 | 3.421
20210048.100 | 847 | 10.1058 | -33.8193 | 8.8645 | 55.6604 | 6.279
20210048.101 | 849 | 10.6046 | -32.7483 | 12.4706 | 51.908 | 4.1624
20210048.102 | 851 | 8.6305 | -29.4059 | 14.3991 | 46.5687 | 3.2341
20210048.103 | 853 | 9.8824 | -30.9296 | 13.4686 | 49.7576 | 3.6943
20210048.104 | 855 | 10.9168 | -30.4355 | 13.534 | 49.44 | 3.653
20210048.105 | 857 | 10.2445 | -30.9286 | 13.0178 | 51.0336 | 3.9203
20210048.106 | 859 | 10.2215 | -31.3964 | 13.4805 | 49.4608 | 3.6691
20210048.107 | 861 | 10.6964 | -30.6295 | 13.8992 | 49.1831 | 3.5386
20210048.108 | 863 | 8.0879 | -27.3892 | 14.4835 | 47.5784 | 3.285
20210048.109 | 865 | 10.2704 | -29.9505 | 13.7352 | 49.596 | 3.6109
20210048.110 | 867 | 11.6689 | -34.2801 | 8.254 | 59.7406 | 7.2378
20210048.110 | Replicate E11 | 11.6998 | -33.4172 | 9.7175 | 57.5684 | 5.9242
20210048.111 | 869 | [5 empty] | Sample missing from tray.
20210048.112 | 871 | 10.1737 | -30.1536 | 14.389 | 47.8622 | 3.3263
20210048.113 | 873 | 9.4904 | -30.2314 | 14.638 | 47.4881 | 3.2442
20210048.114 | 875 | 7.4984 | -29.9263 | 14.5257 | 48.1374 | 3.314
20210048.115 | 877 | 10.2954 | -33.6304 | 11.6242 | 53.5877 | 4.61
20210048.116 | 879 | 10.2894 | -29.7606 | 14.7611 | 47.49 | 3.2172
20210048.117 | 881 | 9.7906 | -30.1779 | 14.7341 | 47.6085 | 3.2312
20210048.118 | 883 | 10.5278 | -29.7878 | 14.3076 | 47.2728 | 3.304
20210048.119 | 885 | 12.051 | -31.3934 | 12.2251 | 52.5804 | 4.301
20210048.120 | 887 | 9.8644 | -28.8886 | 14.8365 | 47.3762 | 3.1932
20210048.121 | 889 | 10.055 | -28.9866 | 14.6396 | 47.2728 | 3.2291
20210048.122 | 891 | 11.1353 | -30.0435 | 13.8194 | 47.9241 | 3.4679
20210048.123 | 893 | 9.4106 | -30.9175 | 13.8323 | 46.8097 | 3.3841
20210048.124 | 895 | 10.457 | -31.0721 | 13.6862 | 50.3854 | 3.6815
20210048.125 | 897 | 9.8016 | -30.7366 | 14.041 | 47.2587 | 3.3658
20210048.126 | 899 | 7.8325 | -30.4567 | 13.1978 | 49.2717 | 3.7333
20210048.127 | 901 | 9.643 | -32.6382 | 11.5576 | 52.5355 | 4.5455
20210048.128 | 903 | 10.0829 | -32.2158 | 13.1859 | 50.5094 | 3.8306
20210048.129 | 905 | 9.655 | -32.6816 | 11.8749 | 52.6395 | 4.4328
20210048.130 | 907 | 10.1397 | -32.717 | 12.4451 | 52.6011 | 4.2267
20210048.131 | 909 | 9.9203 | -30.9538 | 14.3419 | 48.2331 | 3.3631
20210048.132 | 911 | 10.3512 | -31.0256 | 14.3359 | 48.0809 | 3.3539
20210048.133 | 913 | 10.0579 | -30.7558 | 14.2056 | 48.8886 | 3.4415
20210048.134 | 915 | 10.1507 | -31.8481 | 13.6658 | 49.6924 | 3.6363
20210048.135 | 917 | 10.2195 | -30.5921 | 14.6687 | 47.4433 | 3.2343
20210048.136 | 919 | 9.7158 | -31.5571 | 13.5156 | 49.081 | 3.6314
20210048.136 | Replicate H8 | 9.6719 | -31.6652 | 13.6416 | 49.174 | 3.6047
20210048.137 | 921 | 10.2223 | -29.4226 | 12.7789 | 40.8389 | 3.1958
20210048.138 | 923 | 10.3302 | -31.965 | 13.7796 | 50.7746 | 3.6848
20210048.138 | Replicate A11 | 10.3571 | -32.591 | 12.8465 | 51.782 | 4.0308
20210048.139 | 925 | 9.799 | -32.1034 | 12.4896 | 51.4176 | 4.1168
20210048.140 | 927 | 10.2683 | -32.0983 | 13.4429 | 50.6637 | 3.7688
20210048.141 | 929 | 9.3268 | -31.7792 | 13.053 | 51.6324 | 3.9556
20210048.142 | 931 | 8.9953 | -28.5967 | 14.3203 | 47.4286 | 3.312
20210048.143 | 933 | 7.7992 | -30.713 | 14.2201 | 49.3048 | 3.4673
20210048.144 | 935 | 10.0346 | -30.5646 | 14.6815 | 48.495 | 3.3031
20210048.145 | 937 | 10.2882 | -34.36 | 9.6718 | 57.4919 | 5.9443
20210048.146 | 939 | 10.2044 | -32.0418 | 13.4212 | 50.7269 | 3.7796
20210048.147 | 941 | 10.8743 | -32.6132 | 12.2121 | 53.8261 | 4.4076
20210048.148 | 943 | 10.9731 | -29.8537 | 14.432 | 49.1264 | 3.404
20210048.149 | 945 | 10.2882 | -32.6971 | 12.8917 | 52.589 | 4.0793
20210048.150 | 947 | 10.2673 | -32.2114 | 13.2954 | 51.1734 | 3.8489
20210048.151 | 949 | 10.3322 | -31.1916 | 14.1459 | 49.5768 | 3.5047
20210048.152 | 951 | 10.5538 | -29.5064 | 14.7276 | 48.1699 | 3.2707
20210048.153 | 953 | 11.6061 | -31.0724 | 13.2715 | 50.7079 | 3.8208
20210048.154 | 955 | 8.2275 | -30.3051 | 14.0425 | 49.3257 | 3.5126
20210048.155 | 957 | 9.8729 | -30.3909 | 14.9942 | 47.4616 | 3.1653
20210048.156 | 959 | 10.3621 | -28.2776 | 15.0442 | 48.3398 | 3.2132
20210048.157 | 961 | 11.3176 | -31.3208 | 12.7703 | 51.5068 | 4.0333
20210048.158 | 963 | 8.0578 | -29.5125 | 14.4959 | 48.8611 | 3.3707
20210048.158 | Replicate C8 | 8.0898 | -29.4014 | 14.3791 | 47.9662 | 3.3358
20210048.159 | 965 | 11.1339 | -29.0702 | 14.2494 | 49.4185 | 3.4681
20210048.160 | 967 | 9.7062 | -32.4578 | 13.0246 | 53.0575 | 4.0736
20210048.161 | 969 | 10.2193 | -31.5036 | 13.9622 | 49.3035 | 3.5312
20210048.162 | 971 | 9.8939 | -31.9256 | 12.9792 | 51.5328 | 3.9704
20210048.163 | 973 | 9.7631 | -32.4648 | 12.4862 | 52.9922 | 4.2441
20210048.164 | 975 | 10.0676 | -31.2754 | 14.1257 | 49.1767 | 3.4814
20210048.165 | 749 | 10.1814 | -30.2657 | 15.2703 | 47.7808 | 3.129
20210048.166 | 751 | 9.3318 | -29.9507 | 13.6797 | 49.4367 | 3.6139
20210048.167 | 753 | 10.4849 | -31.3451 | 13.5426 | 49.0354 | 3.6208
20210048.168 | 755 | 10.7505 | -30.6756 | 14.1902 | 48.687 | 3.431
20210048.169 | 757 | 9.9608 | -31.1078 | 13.308 | 50.3187 | 3.7811
20210048.170 | 759 | 10.3082 | -27.7354 | 14.7876 | 48.3501 | 3.2696
20210048.171 | 761 | 10.3332 | -31.9145 | 10.4086 | 54.979 | 5.2821
20210048.172 | 763 | 10.6267 | -28.0484 | 13.8225 | 48.9367 | 3.5404
20210048.173 | 765 | 9.5215 | -30.3404 | 14.1585 | 47.9658 | 3.3878
20210048.174 | 767 | 10.1115 | -30.611 | 14.2313 | 47.7739 | 3.3569
20210048.175 | 769 | 10.1075 | -31.8156 | 12.6803 | 49.5222 | 3.9054
20210048.176 | 771 | 10.3332 | -28.3644 | 14.2434 | 47.4758 | 3.3332
20210048.177 | 773 | 9.5304 | -30.2768 | 14.3383 | 47.9216 | 3.3422
20210048.178 | 775 | 10.2323 | -32.173 | 12.3557 | 53.0346 | 4.2923
20210048.178 | Replicate E11 | 10.2962 | -31.0997 | 14.0262 | 50.7223 | 3.6163
20210048.179 | 777 | 9.816 | -28.9763 | 14.5829 | 47.0771 | 3.2282
20210048.180 | 779 | 9.3467 | -30.9432 | 12.2078 | 51.0175 | 4.1791
20210048.181 | 781 | 11.4813 | -31.031 | 14.1266 | 48.5731 | 3.4384
20210048.182 | 783 | 10.2812 | -28.1332 | 13.6847 | 48.3642 | 3.5342
20210048.183 | 785 | 10.413 | -30.8483 | 11.8994 | 52.4117 | 4.4046
20210048.184 | 787 | 9.7321 | -31.3168 | 13.1625 | 49.6648 | 3.7732
20210048.185 | 789 | 8.9813 | -29.2065 | 14.5529 | 47.3997 | 3.2571
20210048.186 | 791 | 8.7177 | -30.4434 | 13.3536 | 47.4225 | 3.5513
20210048.187 | 793 | 10.2812 | -30.6847 | 13.3527 | 49.1869 | 3.6837
20210048.188 | 795 | 10.1005 | -30.7493 | 13.9805 | 49.2479 | 3.5226
20210048.189 | 797 | 10.5957 | -30.6746 | 13.8267 | 48.9535 | 3.5405
20210048.190 | 799 | 8.7707 | -29.4075 | 14.1741 | 45.9723 | 3.2434
20210048.191 | 801 | 10.1235 | -31.5016 | 12.787 | 51.4494 | 4.0236
20210048.192 | 803 | 10.1774 | -29.5337 | 14.2302 | 47.403 | 3.3312
20210048.193 | 805 | 10.4769 | -30.2899 | 14.7759 | 48.0091 | 3.2491
20210048.194 | 807 | 10.9252 | -31.3703 | 13.2718 | 49.2225 | 3.7088
20210048.195 | 809 | 10.2743 | -29.5215 | 13.9075 | 48.8085 | 3.5095
20210048.196 | 811 | 9.8969 | -29.5145 | 14.6545 | 48.807 | 3.3305
20210048.197 | 813 | 7.5217 | -28.734 | 13.8199 | 47.8041 | 3.4591
20210048.198 | 815 | 9.3677 | -32.588 | 10.4787 | 55.5948 | 5.3055
20210048.199 | 817 | 10.446 | -30.6605 | 14.0605 | 47.9044 | 3.407
20210048.200 | 819 | 9.8759 | -30.8665 | 14.1569 | 48.1518 | 3.4013
20210048.201 | 821 | 9.3068 | -30.2253 | 13.0082 | 51.3104 | 3.9445
20210048.202 | 823 | 11.3026 | -30.7695 | 13.7622 | 48.4566 | 3.521
20210048.203 | 825 | 11.6011 | -30.5716 | 13.6014 | 48.4742 | 3.5639
20210048.204 | 827 | 10.2593 | -30.82 | 13.9885 | 48.3495 | 3.4564
20210048.204 | Replicate H8 | 10.2473 | -30.819 | 13.8373 | 48.5015 | 3.5051
20210048.205 | 829 | 10.0768 | -30.2236 | 13.3758 | 49.778 | 3.7215
20210048.206 | 831 | 9.7486 | -32.8444 | 11.1422 | 54.8929 | 4.9266
20210048.206 | Replicate A11 | 9.7746 | -32.8938 | 10.8072 | 53.451 | 4.9459
20210048.207 | 833 | 10.0808 | -29.8344 | 14.2282 | 50.0282 | 3.5161
20210048.208 | 835 | 10.3128 | -31.1836 | 13.1082 | 50.5279 | 3.8547
20210048.209 | 837 | 10.715 | -31.8975 | 13.0509 | 51.1799 | 3.9216
20210048.210 | 839 | 10.0107 | -30.8871 | 12.9241 | 54.3287 | 4.2037
20210048.211 | 841 | 10.0878 | -30.0986 | 13.1665 | 51.4145 | 3.9049
20210048.212 | 843 | 9.9257 | -30.7006 | 13.6307 | 49.4706 | 3.6294
20210048.213 | 845 | 8.7722 | -29.4461 | 14.7286 | 46.7683 | 3.1753
20210048.214 | 002A | 3.9313 | -33.4565 | 9.6423 | 43.0279 | 4.4624
20210048.215 | 002B | 3.9123 | -33.273 | 9.7709 | 43.3295 | 4.4345
20210048.216 | 002C | 3.8052 | -33.5281 | 9.9807 | 43.8315 | 4.3916
20210048.217 | 004A | 6.7234 | -33.5352 | 11.1254 | 49.5493 | 4.4537
20210048.218 | 004B | 6.7274 | -34.5577 | 9.59 | 52.393 | 5.4633
20210048.219 | 004C | 6.1532 | -35.4794 | 8.2043 | 53.5824 | 6.531
20210048.220 | 006A | 3.4221 | -33.0703 | 9.7748 | 45.5939 | 4.6644
20210048.221 | 006B | 3.5361 | -33.1671 | 9.8236 | 45.3654 | 4.618
20210048.222 | 006C | 3.3631 | -33.3254 | 9.595 | 44.6923 | 4.6579
20210048.223 | 008A | 6.4583 | -33.1641 | 11.1724 | 49.3266 | 4.4151
20210048.224 | 008B | 6.0031 | -33.3567 | 10.5121 | 48.6797 | 4.6308
20210048.225 | 008C | 6.0151 | -32.9544 | 10.7626 | 47.9594 | 4.4561
20210048.226 | 010A | 4.1134 | -33.9617 | 9.3856 | 46.0383 | 4.9052
20210048.226 | Replicate C8 | 4.0413 | -33.9991 | 9.4198 | 46.4396 | 4.93
20210048.227 | 010B | 4.1184 | -33.8175 | 9.8526 | 46.9475 | 4.765
20210048.228 | 010C | 4.0824 | -33.6824 | 9.595 | 45.8262 | 4.776
20210048.229 | 012A | 7.0485 | -33.149 | 11.1343 | 48.1186 | 4.3216
20210048.230 | 012B | 6.3363 | -32.7708 | 10.5502 | 46.4063 | 4.3986
20210048.231 | 012C | 6.6984 | -32.8938 | 11.3471 | 48.2655 | 4.2536
20210048.232 | 014A | 6.7374 | -32.4391 | 11.3395 | 45.282 | 3.9933
20210048.233 | 014B | 6.5784 | -33.4404 | 10.1419 | 49.7424 | 4.9046
20210048.234 | 014C | 6.3943 | -32.5762 | 11.0009 | 45.6241 | 4.1473
20210048.235 | 016A | 6.6154 | -32.9493 | 10.5088 | 46.3416 | 4.4098
20210048.236 | 016B | 7.1036 | -32.4784 | 11.1154 | 45.3065 | 4.076
20210048.237 | 016C | 7.0365 | -33.0996 | 10.9536 | 48.2336 | 4.4034
20210048.238 | 018A | 7.0906 | -33.1722 | 11.2371 | 48.1816 | 4.2877
20210048.239 | 018B | 7.2966 | -33.7923 | 10.1029 | 49.9366 | 4.9428
20210048.240 | 018C | 7.8929 | -33.6451 | 10.4402 | 49.882 | 4.7779
Stable Isotope Facility

## QAQC, calculations
Line | QA Reference Material 1 | δ15N | δ13C | Line | QA Reference Material 2 | δ15N | δ13C | Line | QC Reference Material 3 | δ15N | δ13C |  | wt% Nitrogen | wt'% Carbon |  | wt% Nitrogen | wt'% Carbon
|  | normalized | normalized |  |  | normalized | normalized |  |  | normalized | normalized |  | measured values | measured values |  | absolute error | absolute error
3 | 36-UWSIF-UT Glut 1-20210048.13 | -4.623 | -28.0321 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 27.7733 | 24.3783 | 8 | 01-UWSIF-Liver-20210048.11 | 6.7823 | -17.7201 |  | 13.137 | 50.4859 |  | 0.25701 | 0.84593
4 | 36-UWSIF-UT Glut 1-20210048.14 | -4.607 | -28.3458 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 27.9051 | 24.2703 | 9 | 01-UWSIF-Liver-20210048.12 | 6.8123 | -17.829 |  | 13.0696 | 50.263 |  | 0.18956 | 0.62295
5 | 36-UWSIF-UT Glut 1-20210048.15 | -4.6709 | -28.4215 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 27.8781 | 24.2956 | 38 | 01-UWSIF-Liver-20210048.13 | 6.7763 | -17.8058 |  | 12.9628 | 50.4109 |  | 0.082762 | 0.77088
34 | 36-UWSIF-UT Glut 1-20210048.16 | -4.619 | -28.2661 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 27.946 | 24.4489 | 39 | 01-UWSIF-Liver-20210048.14 | 6.7344 | -17.7847 |  | 13.0363 | 50.7242 |  | 0.15634 | 1.0842
35 | 36-UWSIF-UT Glut 1-20210048.17 | -4.5481 | -28.3226 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 27.9001 | 24.5599 | 68 | 01-UWSIF-Liver-20210048.15 | 6.7304 | -17.6888 |  | 12.9184 | 50.5363 |  | 0.038439 | 0.89634
64 | 36-UWSIF-UT Glut 1-20210048.18 | -4.624 | -28.2137 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 27.8991 | 24.2875 | 69 | 01-UWSIF-Liver-20210048.16 | 6.7234 | -17.7443 |  | 12.9316 | 50.6414 |  | 0.051557 | 1.0014
65 | 36-UWSIF-UT Glut 1-20210048.19 | -4.612 | -28.3317 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 27.946 | 24.2855 | 98 | 01-UWSIF-Liver-20210048.17 | 6.6955 | -17.8886 |  | 12.7475 | 49.8914 |  | 0.13253 | 0.25138
94 | 36-UWSIF-UT Glut 1-20210048.111 | -4.629 | -28.3156 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 27.8592 | 24.3702 | 99 | 01-UWSIF-Liver-20210048.18 | 6.7464 | -17.7897 |  | 12.7605 | 50.002 |  | 0.11949 | 0.36202
95 | 36-UWSIF-UT Glut 1-20210048.112 | -4.611 | -28.2621 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 27.9232 | 24.1948 | 8 | 01-UWSIF-Liver-20210048.21 | 6.8091 | -17.6994 |  | 13.059 | 50.3669 |  | 0.17896 | 0.72691
3 | 36-UWSIF-UT Glut 1-20210048.23 | -4.6422 | -28.2763 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 27.9481 | 24.4353 | 9 | 01-UWSIF-Liver-20210048.22 | 6.7901 | -17.754 |  | 13.0774 | 50.4634 |  | 0.19737 | 0.82344
4 | 36-UWSIF-UT Glut 1-20210048.24 | -4.5774 | -28.3308 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 27.9431 | 24.3181 | 38 | 01-UWSIF-Liver-20210048.23 | 6.7532 | -17.9288 |  | 12.9558 | 49.8053 |  | 0.075825 | 0.1653
5 | 36-UWSIF-UT Glut 1-20210048.25 | -4.6302 | -28.542 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 27.8204 | 24.4312 | 39 | 01-UWSIF-Liver-20210048.24 | 6.7542 | -17.9156 |  | 12.9834 | 49.8134 |  | 0.10336 | 0.17339
34 | 36-UWSIF-UT Glut 1-20210048.26 | -4.5973 | -28.1823 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 27.8484 | 24.3878 | 68 | 01-UWSIF-Liver-20210048.25 | 6.7443 | -17.6812 |  | 12.8421 | 50.1619 |  | 0.037886 | 0.52187
35 | 36-UWSIF-UT Glut 1-20210048.27 | -4.5664 | -28.2712 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 27.8763 | 24.41 | 69 | 01-UWSIF-Liver-20210048.26 | 6.7353 | -17.8075 |  | 12.9159 | 50.2422 |  | 0.035922 | 0.60222
64 | 36-UWSIF-UT Glut 1-20210048.28 | -4.5923 | -28.1793 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 27.9202 | 24.2958 | 98 | 01-UWSIF-Liver-20210048.27 | 6.6824 | -17.7277 |  | 12.8169 | 49.8267 |  | 0.06312 | 0.18674
65 | 36-UWSIF-UT Glut 1-20210048.29 | -4.6133 | -28.2429 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 27.8174 | 24.4141 | 99 | 01-UWSIF-Liver-20210048.28 | 6.6944 | -17.6732 |  | 12.8068 | 49.7227 |  | 0.07319 | 0.082721
94 | 36-UWSIF-UT Glut 1-20210048.211 | -4.6532 | -28.1389 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 27.9409 | 24.247 | 8 | 01-UWSIF-Liver-20210048.31 | 6.7599 | -17.8161 |  | 13.101 | 50.3372 |  | 0.22097 | 0.69716
95 | 36-UWSIF-UT Glut 1-20210048.212 | -4.6711 | -28.3339 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 27.9789 | 24.3359 | 9 | 01-UWSIF-Liver-20210048.32 | 6.703 | -17.8514 |  | 13.0377 | 50.1599 |  | 0.15774 | 0.51991
3 | 36-UWSIF-UT Glut 1-20210048.33 | -4.562 | -28.1302 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 27.8621 | 24.3762 | 38 | 01-UWSIF-Liver-20210048.33 | 6.712 | -17.6162 |  | 13.0074 | 49.7454 |  | 0.12737 | 0.10543
4 | 36-UWSIF-UT Glut 1-20210048.34 | -4.5779 | -28.2786 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 28.0068 | 24.3106 | 39 | 01-UWSIF-Liver-20210048.34 | 6.69 | -17.6091 |  | 12.9985 | 49.7138 |  | 0.11845 | 0.073787
5 | 36-UWSIF-UT Glut 1-20210048.35 | -4.6239 | -28.5189 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 27.877 | 24.3853 | 68 | 01-UWSIF-Liver-20210048.35 | 6.7549 | -17.7707 |  | 12.8204 | 49.4922 |  | 0.059608 | 0.14779
34 | 36-UWSIF-UT Glut 1-20210048.36 | -4.6738 | -28.2685 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 27.8072 | 24.4237 | 69 | 01-UWSIF-Liver-20210048.36 | 6.7279 | -17.698 |  | 12.8336 | 49.4161 |  | 0.046434 | 0.22395
35 | 36-UWSIF-UT Glut 1-20210048.37 | -4.6059 | -28.2564 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 27.8481 | 24.451 | 98 | 01-UWSIF-Liver-20210048.37 | 6.7589 | -17.8807 |  | 12.6994 | 49.2395 |  | 0.18062 | 0.40048
64 | 36-UWSIF-UT Glut 1-20210048.38 | -4.6279 | -28.2099 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 27.7822 | 24.3641 | 99 | 01-UWSIF-Liver-20210048.38 | 6.681 | -17.7989 |  | 12.7474 | 49.3033 |  | 0.13264 | 0.33672
65 | 36-UWSIF-UT Glut 1-20210048.39 | -4.5889 | -28.2927 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 27.9059 | 24.4408 | 8 | 01-UWSIF-Liver-20210048.41 | 6.7054 | -17.8203 |  | 12.9981 | 50.1285 |  | 0.1181 | 0.48854
94 | 36-UWSIF-UT Glut 1-20210048.311 | -4.6189 | -28.3038 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 27.9719 | 24.3792 | 9 | 01-UWSIF-Liver-20210048.42 | 6.7754 | -17.8243 |  | 12.9926 | 50.145 |  | 0.1126 | 0.50502
95 | 36-UWSIF-UT Glut 1-20210048.312 | -4.6648 | -28.2493 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 27.8819 | 24.1917 | 38 | 01-UWSIF-Liver-20210048.43 | 6.7454 | -17.8334 |  | 12.9923 | 50.1568 |  | 0.11231 | 0.51676
3 | 36-UWSIF-UT Glut 1-20210048.43 | -4.6281 | -28.2905 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 27.8699 | 24.2562 | 39 | 01-UWSIF-Liver-20210048.44 | 6.7814 | -17.7961 |  | 12.9464 | 49.9661 |  | 0.066442 | 0.32609
4 | 36-UWSIF-UT Glut 1-20210048.44 | -4.5661 | -28.1392 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 27.8629 | 24.4468 | 58 | 01-UWSIF-Liver-20210048.47 | 6.7414 | -17.8233 |  | 12.9481 | 50.0556 |  | 0.068127 | 0.41556
5 | 36-UWSIF-UT Glut 1-20210048.45 | -4.6892 | -28.4024 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 27.8399 | 24.4559 | 59 | 01-UWSIF-Liver-20210048.48 | 6.7644 | -17.7557 |  | 12.8789 | 49.828 |  | 0.001052 | 0.188
34 | 36-UWSIF-UT Glut 1-20210048.46 | -4.6371 | -28.2431 | 30 | δ15N known | 27.89 |  | 30 | δ15N known | 6.81 |  | wt% N known | 12.88
35 | 36-UWSIF-UT Glut 1-20210048.47 | -4.5491 | -28.2764 |  | δ13C known |  | 24.36 |  | δ13C known |  | -17.75 | wt% C known |  | 49.64
54 | 36-UWSIF-UT Glut 1-20210048.411 | -4.6321 | -28.2592 |  | average | 27.888 | 24.3616 |  | average | 6.7422 | -17.7778 | average | 12.9341 | 50.0348 | average | 0.11053 | 0.46876
55 | 36-UWSIF-UT Glut 1-20210048.412 | -4.6111 | -28.3399 |  | std uncertainty | 0.05799 | 0.085722 |  | std uncertainty | 0.036087 | 0.080795 | std uncertainty | 0.11584 | 0.38647 | percent error | 0.85812 | 0.94433
34 | δ15N known | -4.62 | [6 empty] | long-term | 7.21 | -17.45
| δ13C known |  | -28.28 | [5 empty] | acceptable range | 6.41 | -18.05
| average | -4.616 | -28.2785
| std uncertainty | 0.035905 | 0.10131

## Analysis Information
Principal of operation
●The Finnigan DeltaPlus XP is run in continuous flow mode and connected to a Costech 4010 elemental analyzer via a Finnigan ConFlo III interface.
●The Finnigan Delta V is run in continuous flow mode and connected to either a Carlo Erba 1110 or a Thermo Flash isolink elemental analyzer via a Finnigan ConFlo IV interface.
●A sample contained within a tin capsule is dropped into a combustion reactor held at 1020°C. For nitrogen and carbon analysis, the combustion reactor contains chromium oxide for oxidation and silvered cobaltous/cobaltic oxide for removal of sulfur. When the tin capsule is exposed to a gas flow temporarily enriched with ultra-high purity oxygen (25-30 mL min-1), flash combustion occurs which raises the temperature of the sample to >1700°C. The encapsulated sample, depending on its composition, combusts generating one or more of these gases: N2, NxOx, CO2, and H2O. The reduction reactor contains reduced copper wires for the reduction of nitrogen oxides to N2 and the removal of excess O2. An adsorption trap containing magnesium perchlorate removes the H2O. The remaining N2 and CO2 gases travel through a Porapak Q chromatographic column (80-100 mL min-1 at 50˚C) as the carrier gas and then moves to the ConFlo III open-split interface.
●The gas then moves through the ConFlo III to the Finnigan DeltaPlus XP mass spectrometer. Reference gases are pulsed by the ConFlo III into the mass spectrometer for proper mass balance.
QA/QC
Quality assessment of carbon and nitrogen isotope composition of solids is based on the standard uncertainty of the known value of the secondary laboratory reference material calculated on multiple analyses. For carbon isotope composition, if the standard uncertainty is greater than 0.15‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.3‰). The carbon isotopic composition is reported in per mil relative to VPDB scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST 8542 sucrose ANU, respectively are -26.24‰, +37.76‰, and -10.05‰. For nitrogen isotope ratio composition, if the standard uncertainty is greater than 0.2‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.4‰). The nitrogen isotopic composition is reported in per mil relative to nitrogen in air on a scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST8549 potassium nitrate, respectively are -4.52‰, +47.57‰, and 4.36‰.
Normalization
Whenever isotopic analyses are performed, reference materials must be included with the unknowns. In general two ‘bracketing’ reference materials should be used for correcting purposes. These reference materials ideally would be similar in chemical complexity to the unknowns but have very different isotopic compositions. Using two such reference materials, you can generate two linear equations, one for each referenece material. Since these linear equations will have the same slope and intercept, combining the two equations will generate a single linear equation that can be used to correct the unknowns.
The linear equation is derived as follows:
Slope (m): m=[δRM 1-known – δRM 2-known] / [δRM 1-measured – δRM 2-measured]
Intercept (b): b= δRM 1-known – [m* δRM 1-measured]
Normalization equation:
δSA-COR = m* δSA-measured + b
Where:
δRM 1-known = the accepted delta value of reference material 1
δRM 2-known = the accepted delta value of reference material 2
δRM 1-measured = the measured delta value of reference material 1
δRM 2-measured = the measured delta value of reference material 2
δSA-measured = measured delta value of sample
δSA-COR = corrected delta value of sample
13C and 15N analyses of solids

## Contact
For questions about the analysis, please contact:
|  | Chandelle Macdonald | Master Technician
| email: | cmacdon1@uwyo.edu
| Phone: | (307) 766-6373
Contact information for UW Stable Isotope Facility:
| Personnel: | Dr. David G. Williams | Faculty Director
dgw@uwyo.edu
|  | Craig Cook | Facility Director
ccook21@uwyo.edu
|  | Chandelle Macdonald | Master Technician
cmacdon1@uwyo.edu
| Address: | UWYO Stable Isotope Facility
University of Wyoming
Berry Biodiversity Center Rm 214
Laramie, WY 82071
| email: | uwyosif@uwyo.edu
| Phone: | (307) 766-6373
| Fax: | (307) 766-6403

## Tray 1
Date | Line | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Nitrogen % | Carbon % | C:N | Corr. d15N | Corr. d13C
05/17/21 | 10 | 20210048.001 | 501 | 0.823 | 3916 | 13.59 | 850 | -32.506 | 14.0808 | 47.9424 | 3.4048 | 11.7878 | -30.974
05/17/21 | 11 | 20210048.002 | 503 | 0.805 | 3833 | 12.77 | 858 | -33.315 | 14.082 | 49.4339 | 3.5104 | 10.9693 | -31.7902
05/17/21 | 12 | 20210048.002 | Replicate A11 | 0.834 | 4021 | 12.805 | 893 | -33.258 | 14.2457 | 49.695 | 3.4884 | 11.0043 | -31.7327
05/17/21 | 13 | 20210048.003 | 505 | 0.841 | 3989 | 10.204 | 892 | -32.879 | 13.9989 | 49.2138 | 3.5155 | 8.4082 | -31.3503
05/17/21 | 14 | 20210048.004 | 507 | 0.846 | 3816 | 12.526 | 914 | -33.262 | 13.3339 | 50.086 | 3.7563 | 10.7258 | -31.7367
05/17/21 | 15 | 20210048.005 | 509 | 0.83 | 4001 | 12.168 | 873 | -33.502 | 14.2681 | 48.839 | 3.4229 | 10.3685 | -31.9789
05/17/21 | 16 | 20210048.006 | 511 | 0.847 | 4174 | 8.944 | 879 | -32.542 | 14.5578 | 48.1578 | 3.308 | 7.1506 | -31.0103
05/17/21 | 17 | 20210048.007 | 513 | 0.832 | 3350 | 10.151 | 941 | -33.499 | 11.8983 | 52.4939 | 4.4119 | 8.3553 | -31.9758
05/17/21 | 18 | 20210048.008 | 515 | 0.839 | 4075 | 12.337 | 862 | -32.778 | 14.3434 | 47.5313 | 3.3138 | 10.5372 | -31.2484
05/17/21 | 19 | 20210048.009 | 517 | 0.84 | 3905 | 12.444 | 903 | -33.224 | 13.7 | 49.7421 | 3.6308 | 10.644 | -31.6984
05/17/21 | 20 | 20210048.010 | 519 | 0.826 | 3839 | 8.949 | 876 | -33.595 | 13.7089 | 49.1462 | 3.585 | 7.1556 | -32.0727
05/17/21 | 21 | 20210048.011 | 521 | 0.832 | 3794 | 12.899 | 896 | -33.506 | 13.4193 | 49.7891 | 3.7102 | 11.0981 | -31.9829
05/17/21 | 22 | 20210048.012 | 523 | 0.841 | 3968 | 12.704 | 896 | -33.052 | 13.87 | 49.208 | 3.5478 | 10.9035 | -31.5249
05/17/21 | 23 | 20210048.013 | 525 | 0.817 | 4102 | 12.084 | 834 | -31.685 | 14.7831 | 47.2054 | 3.1932 | 10.2846 | -30.1457
05/17/21 | 24 | 20210048.014 | 527 | 0.842 | 3703 | 13.28 | 905 | -33.68 | 12.9262 | 49.6798 | 3.8434 | 11.4784 | -32.1585
05/17/21 | 25 | 20210048.015 | 529 | 0.806 | 3956 | 9.147 | 853 | -31.415 | 14.3938 | 48.7743 | 3.3886 | 7.3532 | -29.8733
05/17/21 | 26 | 20210048.016 | 531 | 0.818 | 4029 | 13.058 | 870 | -32.751 | 14.3933 | 49.0016 | 3.4045 | 11.2568 | -31.2212
05/17/21 | 27 | 20210048.017 | 533 | 0.85 | 3912 | 12.487 | 945 | -30.835 | 13.431 | 51.2161 | 3.8133 | 10.6869 | -29.2881
05/17/21 | 28 | 20210048.018 | 535 | 0.811 | 4090 | 9.468 | 859 | -32.941 | 14.7386 | 48.7378 | 3.3068 | 7.6736 | -31.4129
05/17/21 | 29 | 20210048.019 | 537 | 0.828 | 4124 | 10.733 | 871 | -32.353 | 14.552 | 48.3745 | 3.3243 | 8.9362 | -30.8196
05/17/21 | 30 | 20210048.020 | 541 | 0.844 | 3940 | 11.865 | 908 | -29.54 | 13.6453 | 49.5063 | 3.6281 | 10.0661 | -27.9816
05/17/21 | 31 | 20210048.021 | 543 | 0.826 | 4097 | 9.986 | 866 | -31.269 | 14.5024 | 48.2931 | 3.33 | 8.1906 | -29.726
05/17/21 | 32 | 20210048.022 | 545 | 0.831 | 3794 | 8.84 | 905 | -33.152 | 13.3568 | 50.0971 | 3.7507 | 7.0468 | -31.6258
05/17/21 | 33 | 20210048.022 | Replicate C8 | 0.829 | 3797 | 8.84 | 915 | -33.132 | 13.3526 | 50.7997 | 3.8045 | 7.0468 | -31.6056
05/17/21 | 40 | 20210048.023 | 547 | 0.832 | 4182 | 11.99 | 886 | -30.766 | 14.6859 | 49.0475 | 3.3398 | 10.1908 | -29.2185
05/17/21 | 41 | 20210048.024 | 549 | 0.821 | 4132 | 12.51 | 822 | -31.256 | 14.7625 | 46.1032 | 3.123 | 10.7098 | -29.7129
05/17/21 | 42 | 20210048.025 | 551 | 0.83 | 3591 | 13.091 | 945 | -34.044 | 12.6304 | 52.3387 | 4.1439 | 11.2897 | -32.5257
05/17/21 | 43 | 20210048.026 | 553 | 0.826 | 4149 | 12.83 | 873 | -32.839 | 14.6545 | 48.5703 | 3.3144 | 11.0292 | -31.31
05/17/21 | 44 | 20210048.027 | 555 | 0.827 | 3924 | 12.524 | 910 | -33.221 | 13.8575 | 50.5809 | 3.6501 | 10.7238 | -31.6954
05/17/21 | 45 | 20210048.028 | 557 | 0.82 | 3922 | 12.513 | 902 | -33.315 | 13.9678 | 50.5731 | 3.6207 | 10.7128 | -31.7902
05/17/21 | 46 | 20210048.029 | 559 | 0.84 | 4188 | 12.549 | 895 | -31.919 | 14.5706 | 49.0964 | 3.3696 | 10.7488 | -30.3818
05/17/21 | 47 | 20210048.030 | 561 | 0.843 | 4307 | 9.411 | 872 | -32.66 | 14.9226 | 47.5889 | 3.189 | 7.6167 | -31.1294
05/17/21 | 48 | 20210048.031 | 563 | 0.844 | 4286 | 12.97 | 884 | -31.669 | 14.8483 | 48.2539 | 3.2498 | 11.169 | -30.1296
05/17/21 | 49 | 20210048.032 | 565 | 0.847 | 4325 | 9.562 | 895 | -32.757 | 14.9346 | 48.8057 | 3.268 | 7.7674 | -31.2272
05/17/21 | 50 | 20210048.033 | 567 | 0.82 | 3845 | 13.209 | 907 | -32.658 | 13.7048 | 50.9533 | 3.7179 | 11.4075 | -31.1274
05/17/21 | 51 | 20210048.034 | 569 | 0.837 | 4231 | 9.589 | 851 | -32.431 | 14.7879 | 46.8087 | 3.1653 | 7.7944 | -30.8983
05/17/21 | 52 | 20210048.035 | 571 | 0.848 | 4289 | 9.372 | 871 | -32.726 | 14.7386 | 47.2049 | 3.2028 | 7.5778 | -31.196
05/17/21 | 53 | 20210048.036 | 573 | 0.802 | 3945 | 12.325 | 855 | -31.863 | 14.3844 | 49.0499 | 3.4099 | 10.5252 | -30.3253
05/17/21 | 54 | 20210048.037 | 575 | 0.832 | 3966 | 10.273 | 885 | -32.934 | 13.9285 | 48.9967 | 3.5177 | 8.4771 | -31.4058
05/17/21 | 55 | 20210048.038 | 577 | 0.826 | 4164 | 10.112 | 859 | -32.359 | 14.704 | 47.7977 | 3.2507 | 8.3164 | -30.8257
05/17/21 | 56 | 20210048.039 | 579 | 0.836 | 3060 | 12.481 | 1009 | -32.499 | 10.6961 | 55.5989 | 5.198 | 10.6809 | -30.9669
05/17/21 | 57 | 20210048.040 | 581 | 0.836 | 3989 | 11.349 | 889 | -32.768 | 13.9399 | 48.9235 | 3.5096 | 9.551 | -31.2383
05/17/21 | 58 | 20210048.041 | 583 | 0.836 | 3378 | 12.947 | 988 | -34.764 | 11.773 | 54.319 | 4.6139 | 11.146 | -33.2521
05/17/21 | 59 | 20210048.042 | 585 | 0.839 | 4148 | 10.235 | 871 | -32.395 | 14.4545 | 47.8951 | 3.3135 | 8.4392 | -30.862
05/17/21 | 60 | 20210048.042 | Replicate E11 | 0.832 | 4116 | 10.245 | 866 | -32.312 | 14.4219 | 47.8063 | 3.3148 | 8.4491 | -30.7783
05/17/21 | 61 | 20210048.043 | 587 | 0.805 | 2431 | 12.831 | 994 | -32.999 | 8.8494 | 56.7517 | 6.4131 | 11.0302 | -31.4714
05/18/21 | 62 | 20210048.044 | 589 | 0.818 | 3803 | 12.376 | 901 | -33.411 | 13.5258 | 50.5618 | 3.7382 | 10.5761 | -31.8871
05/18/21 | 63 | 20210048.045 | 591 | 0.812 | 3604 | 13.26 | 877 | -33.542 | 12.9334 | 49.6396 | 3.8381 | 11.4584 | -32.0192
05/18/21 | 70 | 20210048.046 | 593 | 0.822 | 4152 | 12.73 | 872 | -31.499 | 14.7289 | 48.7632 | 3.3107 | 10.9294 | -29.958
05/18/21 | 71 | 20210048.047 | 595 | 0.826 | 3983 | 9.002 | 894 | -33.458 | 14.0177 | 49.6868 | 3.5446 | 7.2085 | -31.9345
05/18/21 | 72 | 20210048.048 | 597 | 0.849 | 4276 | 9.143 | 886 | -31.891 | 14.6869 | 48.0233 | 3.2698 | 7.3492 | -30.3535
05/18/21 | 73 | 20210048.049 | 599 | 0.807 | 3983 | 10.379 | 853 | -32.602 | 14.3735 | 48.5247 | 3.376 | 8.5829 | -31.0709
05/18/21 | 74 | 20210048.050 | 601 | 0.833 | 4005 | 12.882 | 890 | -32.978 | 14.0076 | 49.0704 | 3.5031 | 11.0811 | -31.4502
05/18/21 | 75 | 20210048.051 | 603 | 0.822 | 3921 | 10.957 | 880 | -32.799 | 13.9009 | 49.1306 | 3.5344 | 9.1598 | -31.2696
05/18/21 | 76 | 20210048.052 | 605 | 0.85 | 4231 | 10.06 | 893 | -32.559 | 14.4587 | 48.2495 | 3.3371 | 8.2645 | -31.0275
05/18/21 | 77 | 20210048.053 | 607 | 0.818 | 4068 | 12.396 | 848 | -31.584 | 14.4838 | 47.6062 | 3.2869 | 10.596 | -30.0438
05/18/21 | 78 | 20210048.054 | 609 | 0.833 | 3642 | 12.405 | 965 | -33.279 | 12.7149 | 53.3003 | 4.1919 | 10.605 | -31.7539
05/18/21 | 79 | 20210048.055 | 611 | 0.826 | 3246 | 12.698 | 960 | -33.18 | 11.4487 | 53.3883 | 4.6633 | 10.8975 | -31.654
05/18/21 | 80 | 20210048.056 | 613 | 0.801 | 4030 | 9.36 | 838 | -32.673 | 14.6584 | 47.9841 | 3.2735 | 7.5658 | -31.1425
05/18/21 | 81 | 20210048.057 | 615 | 0.801 | 3439 | 9.342 | 890 | -34.063 | 12.4959 | 50.9964 | 4.081 | 7.5479 | -32.5449
05/18/21 | 82 | 20210048.058 | 617 | 0.848 | 3887 | 10.602 | 920 | -33.13 | 13.3194 | 49.8357 | 3.7416 | 8.8055 | -31.6036
05/18/21 | 83 | 20210048.059 | 619 | 0.839 | 4057 | 13.709 | 863 | -32.012 | 14.0637 | 47.241 | 3.3591 | 11.9066 | -30.4756
05/18/21 | 84 | 20210048.060 | 623 | 0.817 | 4097 | 12.631 | 865 | -32.421 | 14.5562 | 48.5349 | 3.3343 | 10.8306 | -30.8882
05/18/21 | 85 | 20210048.061 | 625 | 0.813 | 4005 | 12.791 | 855 | -32.506 | 14.3776 | 48.2924 | 3.3589 | 10.9903 | -30.974
05/18/21 | 86 | 20210048.062 | 627 | 0.816 | 3791 | 12.504 | 861 | -33.077 | 13.5573 | 48.4113 | 3.5709 | 10.7038 | -31.5501
05/18/21 | 87 | 20210048.063 | 629 | 0.843 | 4210 | 10.196 | 873 | -32.288 | 14.5186 | 47.5388 | 3.2743 | 8.4002 | -30.7541
05/18/21 | 88 | 20210048.064 | 631 | 0.825 | 3648 | 12.644 | 901 | -32.41 | 12.8806 | 50.1564 | 3.8939 | 10.8436 | -30.8771
05/18/21 | 89 | 20210048.065 | 635 | 0.821 | 4069 | 11.307 | 878 | -32.128 | 14.4124 | 49.1291 | 3.4088 | 9.5091 | -30.5926
05/18/21 | 90 | 20210048.066 | 637 | 0.836 | 4090 | 12.918 | 888 | -32.717 | 14.2609 | 48.8186 | 3.4233 | 11.1171 | -31.1869
05/18/21 | 91 | 20210048.067 | 639 | 0.83 | 3947 | 13.182 | 875 | -32.597 | 13.8572 | 48.4888 | 3.4992 | 11.3806 | -31.0658
05/18/21 | 92 | 20210048.068 | 641 | 0.848 | 4252 | 10.997 | 870 | -32.064 | 14.6138 | 47.157 | 3.2269 | 9.1997 | -30.5281
05/18/21 | 93 | 20210048.068 | Replicate H8 | 0.828 | 4218 | 11.025 | 860 | -32.029 | 14.7989 | 47.6548 | 3.2202 | 9.2277 | -30.4928
Reference Check
05/17/21 | 8 | 01-UWSIF-Liver-20210048.11 | 01-UWSIF-Liver- | 0.713 | 3148 | 8.575 | 776 | -19.369 | 13.137 | 50.4859 | 3.843 | 6.7823 | -17.7201
05/17/21 | 9 | 01-UWSIF-Liver-20210048.12 | 01-UWSIF-Liver- | 0.715 | 3137 | 8.605 | 775 | -19.477 | 13.0696 | 50.263 | 3.8458 | 6.8123 | -17.829
05/17/21 | 38 | 01-UWSIF-Liver-20210048.13 | 01-UWSIF-Liver- | 0.776 | 3440 | 8.569 | 852 | -19.454 | 12.9628 | 50.4109 | 3.8889 | 6.7763 | -17.8058
05/17/21 | 39 | 01-UWSIF-Liver-20210048.14 | 01-UWSIF-Liver- | 0.752 | 3351 | 8.527 | 832 | -19.433 | 13.0363 | 50.7242 | 3.891 | 6.7344 | -17.7847
05/18/21 | 68 | 01-UWSIF-Liver-20210048.15 | 01-UWSIF-Liver- | 0.728 | 3215 | 8.523 | 803 | -19.338 | 12.9184 | 50.5363 | 3.912 | 6.7304 | -17.6888
05/18/21 | 69 | 01-UWSIF-Liver-20210048.16 | 01-UWSIF-Liver- | 0.738 | 3269 | 8.516 | 816 | -19.393 | 12.9316 | 50.6414 | 3.9161 | 6.7234 | -17.7443
05/18/21 | 98 | 01-UWSIF-Liver-20210048.17 | 01-UWSIF-Liver- | 0.753 | 3279 | 8.488 | 818 | -19.536 | 12.7475 | 49.8914 | 3.9138 | 6.6955 | -17.8886
05/18/21 | 99 | 01-UWSIF-Liver-20210048.18 | 01-UWSIF-Liver- | 0.735 | 3199 | 8.539 | 801 | -19.438 | 12.7605 | 50.002 | 3.9185 | 6.7464 | -17.7897
[5 empty] | average | 8.5427 |  | -19.4297 | 12.9455 | 50.3694 |  | 6.7501 | -17.7814
[5 empty] | std. dev. | 0.037739 |  | 0.062846 | 0.13886 | 0.2971 |  | 0.037667 | 0.063406
Reference Material
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 | 1327 | -2.852 | 370 | -29.59 | 9.6168 | 41.3622 | 4.301 | -4.623 | -28.0321
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 | 3291 | -2.836 | 899 | -29.901 | 9.4625 | 40.513 | 4.2814 | -4.607 | -28.3458
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 | 10284 | -2.9 | 2395 | -29.976 | 9.5207 | 40.833 | 4.2889 | -4.6709 | -28.4215
05/17/21 | 34 | 36-UWSIF-UT Glut 1-20210048.16 | 36-UWSIF-UT Glut 1- | 0.842 | 2708 | -2.848 | 750 | -29.822 | 9.4947 | 40.9385 | 4.3117 | -4.619 | -28.2661
05/17/21 | 35 | 36-UWSIF-UT Glut 1-20210048.17 | 36-UWSIF-UT Glut 1- | 0.766 | 2437 | -2.777 | 677 | -29.878 | 9.451 | 40.6972 | 4.3061 | -4.5481 | -28.3226
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.18 | 36-UWSIF-UT Glut 1- | 0.783 | 2484 | -2.853 | 694 | -29.77 | 9.3773 | 40.6804 | 4.3382 | -4.624 | -28.2137
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.19 | 36-UWSIF-UT Glut 1- | 0.794 | 2519 | -2.841 | 703 | -29.887 | 9.3719 | 40.624 | 4.3347 | -4.612 | -28.3317
05/18/21 | 94 | 36-UWSIF-UT Glut 1-20210048.111 | 36-UWSIF-UT Glut 1- | 0.754 | 2364 | -2.858 | 664 | -29.871 | 9.2965 | 40.3414 | 4.3394 | -4.629 | -28.3156
05/18/21 | 95 | 36-UWSIF-UT Glut 1-20210048.112 | 36-UWSIF-UT Glut 1- | 0.787 | 2474 | -2.84 | 692 | -29.818 | 9.289 | 40.3623 | 4.3452 | -4.611 | -28.2621
[5 empty] | average | -2.845 |  | -29.8348 | 9.4312 | 40.7058 |  | -4.616 | -28.279
[5 empty] | std. dev. | 0.031745 |  | 0.10885 | 0.10763 | 0.31554 |  | 0.031685 | 0.10982
05/17/21 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 39-UWSIF-UW Glut 2- | 0.728 | 2501 | 29.606 | 690 | 22.358 | 10.3379 | 44.1575 | 4.2714 | 27.7733 | 24.3783
05/17/21 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 39-UWSIF-UW Glut 2- | 0.747 | 2580 | 29.738 | 713 | 22.251 | 10.3375 | 44.2951 | 4.2849 | 27.9051 | 24.2703
05/17/21 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 39-UWSIF-UW Glut 2- | 0.788 | 2722 | 29.711 | 754 | 22.276 | 10.2473 | 44.1684 | 4.3102 | 27.8781 | 24.2956
05/17/21 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 39-UWSIF-UW Glut 2- | 0.782 | 2713 | 29.779 | 752 | 22.428 | 10.2817 | 44.2436 | 4.3031 | 27.946 | 24.4489
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 39-UWSIF-UW Glut 2- | 0.735 | 2520 | 29.733 | 703 | 22.538 | 10.1607 | 44.033 | 4.3336 | 27.9001 | 24.5599
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 39-UWSIF-UW Glut 2- | 0.722 | 2487 | 29.732 | 694 | 22.268 | 10.2194 | 44.2906 | 4.334 | 27.8991 | 24.2875
05/18/21 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 39-UWSIF-UW Glut 2- | 0.703 | 2394 | 29.779 | 670 | 22.266 | 10.1099 | 43.7854 | 4.3309 | 27.946 | 24.2855
05/18/21 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 39-UWSIF-UW Glut 2- | 0.83 | 2829 | 29.692 | 787 | 22.35 | 10.0426 | 43.6353 | 4.345 | 27.8592 | 24.3702
[5 empty] | average | 29.7213 |  | 22.3419 | 10.2171 | 44.0761 |  | 27.8884 | 24.362
[5 empty] | std. dev. | 0.055353 |  | 0.10001 | 0.10645 | 0.24431 |  | 0.055248 | 0.10091
15N Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -2.845 | -4.616
|  | UWSIF 39- UW Glutamic 2 | 29.7213 | 27.888
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF01 (Liver) | 6.7501 | 6.81 | 0.059881
Percentage
|  | UWSIF01 (Liver) | 12.9455 | 12.88
13C Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -29.8348 | -28.279
|  | UWSIF 39- UW Glutamic 2 | 22.3419 | 24.362
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF01 (Liver) | -17.7814 | -17.75 | 0.031375
Percentage
|  | UWSIF01 (Liver) | 50.3694 | 49.64
[6 empty] | Date Analyzed: |  | 2021-05-18
[6 empty] | Analyst: |  | sme
[6 empty] | Instrument: | Carlo Erba

## Tray 2
Date | Line | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Nitrogen % | Carbon % | C:N | Corr. d15N | Corr. d13C
05/18/21 | 10 | 20210048.069 | 643 | 0.806 | 3894 | 12.548 | 844 | -31.649 | 14.7932 | 49.2795 | 3.3312 | 10.7442 | -30.099
05/18/21 | 11 | 20210048.070 | 645 | 0.814 | 4036 | 11.909 | 846 | -31.644 | 15.1732 | 48.9549 | 3.2264 | 10.1068 | -30.094
05/18/21 | 12 | 20210048.070 | Replicate A11 | 0.84 | 4158 | 11.929 | 872 | -31.616 | 15.1423 | 48.9229 | 3.2309 | 10.1268 | -30.0657
05/18/21 | 13 | 20210048.071 | 651 | 0.824 | 4064 | 8.812 | 838 | -31.711 | 15.0998 | 47.8608 | 3.1696 | 7.0176 | -30.1617
05/18/21 | 14 | 20210048.072 | 653 | 0.845 | 4071 | 8.639 | 881 | -32.292 | 14.6697 | 48.9291 | 3.3354 | 6.845 | -30.7487
05/18/21 | 15 | 20210048.073 | 655 | 0.84 | 4065 | 12.569 | 888 | -31.397 | 14.7152 | 49.6143 | 3.3716 | 10.7652 | -29.8444
05/18/21 | 16 | 20210048.074 | 657 | 0.837 | 4012 | 8.67 | 873 | -31.821 | 14.6388 | 49.0862 | 3.3531 | 6.8759 | -30.2728
05/18/21 | 17 | 20210048.075 | 659 | 0.838 | 3833 | 11.905 | 903 | -29.306 | 13.9163 | 50.7763 | 3.6487 | 10.1028 | -27.7317
05/18/21 | 18 | 20210048.076 | 661 | 0.846 | 4009 | 12.152 | 862 | -30.974 | 14.4542 | 47.9373 | 3.3165 | 10.3492 | -29.417
05/18/21 | 19 | 20210048.077 | 663 | 0.818 | 4046 | 12.894 | 839 | -32.338 | 15.1015 | 48.2594 | 3.1957 | 11.0894 | -30.7952
05/18/21 | 20 | 20210048.078 | 665 | 0.824 | 3053 | 12.812 | 878 | -34.457 | 11.3278 | 50.2376 | 4.4349 | 11.0076 | -32.9363
05/18/21 | 21 | 20210048.079 | 667 | 0.802 | 3890 | 10.027 | 819 | -32.76 | 14.7922 | 47.7844 | 3.2304 | 8.2295 | -31.2216
05/18/21 | 22 | 20210048.080 | 669 | 0.836 | 4143 | 12.39 | 856 | -31.62 | 15.0876 | 47.9493 | 3.1781 | 10.5866 | -30.0697
05/18/21 | 23 | 20210048.081 | 671 | 0.828 | 4104 | 9.875 | 848 | -32.656 | 15.131 | 47.909 | 3.1663 | 8.0779 | -31.1165
05/18/21 | 24 | 20210048.082 | 673 | 0.816 | 3965 | 12.267 | 828 | -32.448 | 14.8596 | 47.5687 | 3.2012 | 10.4639 | -30.9064
05/18/21 | 25 | 20210048.083 | 703 | 0.832 | 3511 | 10.044 | 901 | -33.758 | 12.8267 | 50.8828 | 3.9669 | 8.2465 | -32.23
05/18/21 | 26 | 20210048.084 | 705 | 0.801 | 3498 | 13.419 | 856 | -32.737 | 13.3644 | 50.2173 | 3.7576 | 11.6131 | -31.1984
05/18/21 | 27 | 20210048.085 | 709 | 0.835 | 4087 | 8.927 | 858 | -32.089 | 14.928 | 48.1847 | 3.2278 | 7.1323 | -30.5436
05/18/21 | 28 | 20210048.086 | 713 | 0.829 | 3576 | 12.342 | 885 | -31.08 | 13.1582 | 50.1952 | 3.8148 | 10.5387 | -29.5241
05/18/21 | 29 | 20210048.087 | 715 | 0.81 | 3136 | 10.601 | 935 | -36.106 | 11.7339 | 54.2136 | 4.6202 | 8.8021 | -34.6024
05/18/21 | 30 | 20210048.088 | 719 | 0.829 | 3940 | 12.448 | 859 | -32.646 | 14.3995 | 48.3685 | 3.359 | 10.6445 | -31.1064
05/18/21 | 31 | 20210048.089 | 721 | 0.814 | 3828 | 12.965 | 845 | -32.968 | 14.3432 | 48.6707 | 3.3933 | 11.1602 | -31.4318
05/18/21 | 32 | 20210048.090 | 723 | 0.825 | 4061 | 9.013 | 841 | -32.878 | 14.9333 | 47.6555 | 3.1912 | 7.2181 | -31.3408
05/18/21 | 33 | 20210048.090 | Replicate C8 | 0.81 | 4024 | 9.098 | 831 | -32.931 | 15.1184 | 47.9481 | 3.1715 | 7.3029 | -31.3944
05/18/21 | 40 | 20210048.091 | 725 | 0.827 | 3783 | 12.448 | 881 | -33.413 | 13.9065 | 49.9696 | 3.5933 | 10.6445 | -31.8814
05/18/21 | 41 | 20210048.092 | 727 | 0.807 | 3914 | 10.218 | 837 | -32.867 | 14.6505 | 48.413 | 3.3045 | 8.4201 | -31.3297
05/18/21 | 42 | 20210048.093 | 729 | 0.85 | 4022 | 13.663 | 894 | -32.187 | 14.3059 | 49.1498 | 3.4356 | 11.8564 | -30.6426
05/18/21 | 43 | 20210048.094 | 731 | 0.823 | 3563 | 12.792 | 908 | -33.942 | 13.0812 | 51.6628 | 3.9494 | 10.9876 | -32.4159
05/18/21 | 44 | 20210048.095 | 733 | 0.837 | 4184 | 12.792 | 854 | -31.889 | 15.1021 | 47.654 | 3.1555 | 10.9876 | -30.3415
05/18/21 | 45 | 20210048.096 | 735 | 0.85 | 3998 | 9.97 | 893 | -33.027 | 14.1989 | 49.1359 | 3.4606 | 8.1727 | -31.4914
05/18/21 | 46 | 20210048.097 | 739 | 0.826 | 3972 | 11.513 | 835 | -32.072 | 14.522 | 47.1583 | 3.2474 | 9.7118 | -30.5264
05/18/21 | 47 | 20210048.098 | 741 | 0.825 | 3878 | 13.087 | 851 | -31.878 | 14.2719 | 48.2899 | 3.3836 | 11.2819 | -30.3304
05/18/21 | 48 | 20210048.099 | 747 | 0.849 | 4041 | 10.437 | 892 | -32.791 | 14.3789 | 49.1898 | 3.421 | 8.6385 | -31.2529
05/18/21 | 49 | 20210048.100 | 847 | 0.815 | 2397 | 11.908 | 964 | -35.331 | 8.8645 | 55.6604 | 6.279 | 10.1058 | -33.8193
05/18/21 | 50 | 20210048.101 | 849 | 0.833 | 3440 | 12.408 | 922 | -34.271 | 12.4706 | 51.908 | 4.1624 | 10.6046 | -32.7483
05/18/21 | 51 | 20210048.102 | 851 | 0.82 | 3904 | 10.429 | 818 | -30.963 | 14.3991 | 46.5687 | 3.2341 | 8.6305 | -29.4059
05/18/21 | 52 | 20210048.103 | 853 | 0.833 | 3722 | 11.684 | 887 | -32.471 | 13.4686 | 49.7576 | 3.6943 | 9.8824 | -30.9296
05/18/21 | 53 | 20210048.104 | 855 | 0.825 | 3690 | 12.721 | 872 | -31.982 | 13.534 | 49.44 | 3.653 | 10.9168 | -30.4355
05/18/21 | 54 | 20210048.105 | 857 | 0.848 | 3649 | 12.047 | 921 | -32.47 | 13.0178 | 51.0336 | 3.9203 | 10.2445 | -30.9286
05/18/21 | 55 | 20210048.106 | 859 | 0.848 | 3771 | 12.024 | 894 | -32.933 | 13.4805 | 49.4608 | 3.6691 | 10.2215 | -31.3964
05/18/21 | 56 | 20210048.107 | 861 | 0.806 | 3701 | 12.5 | 847 | -32.174 | 13.8992 | 49.1831 | 3.5386 | 10.6964 | -30.6295
05/18/21 | 57 | 20210048.108 | 863 | 0.812 | 3878 | 9.885 | 826 | -28.967 | 14.4835 | 47.5784 | 3.285 | 8.0879 | -27.3892
05/18/21 | 58 | 20210048.109 | 865 | 0.827 | 3753 | 12.073 | 874 | -31.502 | 13.7352 | 49.596 | 3.6109 | 10.2704 | -29.9505
05/18/21 | 59 | 20210048.110 | 867 | 0.816 | 2236 | 13.475 | 1032 | -35.787 | 8.254 | 59.7406 | 7.2378 | 11.6689 | -34.2801
05/18/21 | 60 | 20210048.110 | Replicate E11 | 0.845 | 2727 | 13.506 | 1032 | -34.933 | 9.7175 | 57.5684 | 5.9242 | 11.6998 | -33.4172
05/18/21 | 61 | 20210048.111 | 869 | 0.834 | [9 empty] | Sample not in tray.
05/18/21 | 62 | 20210048.112 | 871 | 0.825 | 3917 | 11.976 | 843 | -31.703 | 14.389 | 47.8622 | 3.3263 | 10.1737 | -30.1536
05/18/21 | 63 | 20210048.113 | 873 | 0.839 | 4048 | 11.291 | 850 | -31.78 | 14.638 | 47.4881 | 3.2442 | 9.4904 | -30.2314
05/18/21 | 70 | 20210048.114 | 875 | 0.826 | 3949 | 9.294 | 845 | -31.478 | 14.5257 | 48.1374 | 3.314 | 7.4984 | -29.9263
05/18/21 | 71 | 20210048.115 | 877 | 0.832 | 3197 | 12.098 | 947 | -35.144 | 11.6242 | 53.5877 | 4.61 | 10.2954 | -33.6304
05/18/21 | 72 | 20210048.116 | 879 | 0.824 | 4014 | 12.092 | 836 | -31.314 | 14.7611 | 47.49 | 3.2172 | 10.2894 | -29.7606
05/18/21 | 73 | 20210048.117 | 881 | 0.845 | 4095 | 11.592 | 857 | -31.727 | 14.7341 | 47.6085 | 3.2312 | 9.7906 | -30.1779
05/18/21 | 74 | 20210048.118 | 883 | 0.841 | 3957 | 12.331 | 848 | -31.341 | 14.3076 | 47.2728 | 3.304 | 10.5278 | -29.7878
05/18/21 | 75 | 20210048.119 | 885 | 0.807 | 3250 | 13.858 | 902 | -32.93 | 12.2251 | 52.5804 | 4.301 | 12.051 | -31.3934
05/18/21 | 76 | 20210048.120 | 887 | 0.849 | 4153 | 11.666 | 857 | -30.451 | 14.8365 | 47.3762 | 3.1932 | 9.8644 | -28.8886
05/18/21 | 77 | 20210048.121 | 889 | 0.811 | 3910 | 11.857 | 819 | -30.548 | 14.6396 | 47.2728 | 3.2291 | 10.055 | -28.9866
05/18/21 | 78 | 20210048.122 | 891 | 0.801 | 3644 | 12.94 | 820 | -31.594 | 13.8194 | 47.9241 | 3.4679 | 11.1353 | -30.0435
05/18/21 | 79 | 20210048.123 | 893 | 0.813 | 3699 | 11.211 | 813 | -32.459 | 13.8323 | 46.8097 | 3.3841 | 9.4106 | -30.9175
05/18/21 | 80 | 20210048.124 | 895 | 0.812 | 3650 | 12.26 | 871 | -32.612 | 13.6862 | 50.3854 | 3.6815 | 10.457 | -31.0721
05/18/21 | 81 | 20210048.125 | 897 | 0.841 | 3892 | 11.603 | 848 | -32.28 | 14.041 | 47.2587 | 3.3658 | 9.8016 | -30.7366
05/19/21 | 82 | 20210048.126 | 899 | 0.81 | 3525 | 9.629 | 851 | -32.003 | 13.1978 | 49.2717 | 3.7333 | 7.8325 | -30.4567
05/19/21 | 83 | 20210048.127 | 901 | 0.811 | 3097 | 11.444 | 907 | -34.162 | 11.5576 | 52.5355 | 4.5455 | 9.643 | -32.6382
05/19/21 | 84 | 20210048.128 | 903 | 0.832 | 3624 | 11.885 | 897 | -33.744 | 13.1859 | 50.5094 | 3.8306 | 10.0829 | -32.2158
05/19/21 | 85 | 20210048.129 | 905 | 0.8 | 3140 | 11.456 | 898 | -34.205 | 11.8749 | 52.6395 | 4.4328 | 9.655 | -32.6816
05/19/21 | 86 | 20210048.130 | 907 | 0.846 | 3482 | 11.942 | 945 | -34.24 | 12.4451 | 52.6011 | 4.2267 | 10.1397 | -32.717
05/19/21 | 87 | 20210048.131 | 909 | 0.807 | 3815 | 11.722 | 831 | -32.495 | 14.3419 | 48.2331 | 3.3631 | 9.9203 | -30.9538
05/19/21 | 88 | 20210048.132 | 911 | 0.844 | 3993 | 12.154 | 866 | -32.566 | 14.3359 | 48.0809 | 3.3539 | 10.3512 | -31.0256
05/19/21 | 89 | 20210048.133 | 913 | 0.815 | 3823 | 11.86 | 850 | -32.299 | 14.2056 | 48.8886 | 3.4415 | 10.0579 | -30.7558
05/19/21 | 90 | 20210048.134 | 915 | 0.838 | 3783 | 11.953 | 887 | -33.38 | 13.6658 | 49.6924 | 3.6363 | 10.1507 | -31.8481
05/19/21 | 91 | 20210048.135 | 917 | 0.822 | 3981 | 12.022 | 833 | -32.137 | 14.6687 | 47.4433 | 3.2343 | 10.2195 | -30.5921
05/19/21 | 92 | 20210048.136 | 919 | 0.817 | 3644 | 11.517 | 856 | -33.092 | 13.5156 | 49.081 | 3.6314 | 9.7158 | -31.5571
05/19/21 | 93 | 20210048.136 | Replicate H8 | 0.817 | 3677 | 11.473 | 857 | -33.199 | 13.6416 | 49.174 | 3.6047 | 9.6719 | -31.6652
Reference Check
05/18/21 | 8 | 01-UWSIF-Liver-20210048.21 | 01-UWSIF-Liver- | 0.723 | 3081 | 8.603 | 775 | -19.377 | 13.059 | 50.3669 | 3.8569 | 6.8091 | -17.6994
05/18/21 | 9 | 01-UWSIF-Liver-20210048.22 | 01-UWSIF-Liver- | 0.717 | 3071 | 8.584 | 773 | -19.431 | 13.0774 | 50.4634 | 3.8588 | 6.7901 | -17.754
05/18/21 | 38 | 01-UWSIF-Liver-20210048.23 | 01-UWSIF-Liver- | 0.736 | 3138 | 8.547 | 785 | -19.604 | 12.9558 | 49.8053 | 3.8442 | 6.7532 | -17.9288
05/18/21 | 39 | 01-UWSIF-Liver-20210048.24 | 01-UWSIF-Liver- | 0.731 | 3123 | 8.548 | 781 | -19.591 | 12.9834 | 49.8134 | 3.8367 | 6.7542 | -17.9156
05/18/21 | 68 | 01-UWSIF-Liver-20210048.25 | 01-UWSIF-Liver- | 0.741 | 3140 | 8.538 | 795 | -19.359 | 12.8421 | 50.1619 | 3.906 | 6.7443 | -17.6812
05/18/21 | 69 | 01-UWSIF-Liver-20210048.26 | 01-UWSIF-Liver- | 0.767 | 3266 | 8.529 | 823 | -19.484 | 12.9159 | 50.2422 | 3.8899 | 6.7353 | -17.8075
05/19/21 | 98 | 01-UWSIF-Liver-20210048.27 | 01-UWSIF-Liver- | 0.825 | 3496 | 8.476 | 877 | -19.405 | 12.8169 | 49.8267 | 3.8876 | 6.6824 | -17.7277
05/19/21 | 99 | 01-UWSIF-Liver-20210048.28 | 01-UWSIF-Liver- | 0.736 | 3105 | 8.488 | 783 | -19.351 | 12.8068 | 49.7227 | 3.8825 | 6.6944 | -17.6732
[5 empty] | average | 8.5391 |  | -19.4503 | 12.9322 | 50.0503 |  | 6.7454 | -17.7734
[5 empty] | std. dev. | 0.04299 |  | 0.10039 | 0.10537 | 0.29122 |  | 0.042882 | 0.10144
Reference Material
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 | 1441 | -2.877 | 407 | -29.845 | 9.4495 | 40.5235 | 4.2884 | -4.6422 | -28.2763
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 | 2250 | -2.812 | 634 | -29.899 | 9.5877 | 41.0583 | 4.2824 | -4.5774 | -28.3308
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 | 4838 | -2.865 | 1304 | -30.108 | 9.5163 | 40.7823 | 4.2855 | -4.6302 | -28.542
05/18/21 | 34 | 36-UWSIF-UT Glut 1-20210048.26 | 36-UWSIF-UT Glut 1- | 0.731 | 2255 | -2.832 | 627 | -29.752 | 9.4356 | 39.866 | 4.225 | -4.5973 | -28.1823
05/18/21 | 35 | 36-UWSIF-UT Glut 1-20210048.27 | 36-UWSIF-UT Glut 1- | 0.829 | 2572 | -2.801 | 714 | -29.84 | 9.432 | 39.9942 | 4.2403 | -4.5664 | -28.2712
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.28 | 36-UWSIF-UT Glut 1- | 0.827 | 2538 | -2.827 | 717 | -29.749 | 9.3064 | 40.3644 | 4.3373 | -4.5923 | -28.1793
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.29 | 36-UWSIF-UT Glut 1- | 0.791 | 2411 | -2.848 | 680 | -29.812 | 9.3123 | 40.2144 | 4.3184 | -4.6133 | -28.2429
05/19/21 | 94 | 36-UWSIF-UT Glut 1-20210048.211 | 36-UWSIF-UT Glut 1- | 0.744 | 2271 | -2.888 | 645 | -29.709 | 9.2845 | 40.2746 | 4.3378 | -4.6532 | -28.1389
05/19/21 | 95 | 36-UWSIF-UT Glut 1-20210048.212 | 36-UWSIF-UT Glut 1- | 0.755 | 2305 | -2.906 | 654 | -29.902 | 9.2734 | 40.2066 | 4.3357 | -4.6711 | -28.3339
[5 empty] | average | -2.8507 |  | -29.8462 | 9.3998 | 40.3649 |  | -4.6159 | -28.2775
[5 empty] | std. dev. | 0.035763 |  | 0.11874 | 0.11141 | 0.37456 |  | 0.035674 | 0.11998
05/18/21 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 39-UWSIF-UW Glut 2- | 0.742 | 2482 | 29.77 | 698 | 22.086 | 10.2938 | 44.2125 | 4.2951 | 27.9232 | 24.1948
05/18/21 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 39-UWSIF-UW Glut 2- | 0.736 | 2330 | 29.795 | 657 | 22.324 | 9.7602 | 41.9235 | 4.2954 | 27.9481 | 24.4353
05/18/21 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 39-UWSIF-UW Glut 2- | 0.769 | 2594 | 29.79 | 720 | 22.208 | 10.2865 | 43.6591 | 4.2443 | 27.9431 | 24.3181
05/18/21 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 39-UWSIF-UW Glut 2- | 0.841 | 2808 | 29.667 | 779 | 22.32 | 10.193 | 43.3863 | 4.2565 | 27.8204 | 24.4312
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 39-UWSIF-UW Glut 2- | 0.716 | 2387 | 29.695 | 675 | 22.277 | 10.1644 | 44.0236 | 4.3312 | 27.8484 | 24.3878
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 39-UWSIF-UW Glut 2- | 0.798 | 2665 | 29.723 | 753 | 22.299 | 10.1586 | 44.1332 | 4.3444 | 27.8763 | 24.41
05/19/21 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 39-UWSIF-UW Glut 2- | 0.73 | 2418 | 29.767 | 680 | 22.186 | 10.1409 | 43.6557 | 4.3049 | 27.9202 | 24.2958
05/19/21 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 39-UWSIF-UW Glut 2- | 0.768 | 2523 | 29.664 | 711 | 22.303 | 10.0161 | 43.3049 | 4.3235 | 27.8174 | 24.4141
[5 empty] | average | 29.7339 |  | 22.2504 | 10.1267 | 43.5374 |  | 27.8871 | 24.3609
[5 empty] | std. dev. | 0.053816 |  | 0.083726 | 0.172 | 0.73265 |  | 0.053681 | 0.084596
15N Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -2.8507 | -4.616
|  | UWSIF 39- UW Glutamic 2 | 29.7339 | 27.888
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF01 (Liver) | 6.7454 | 6.81 | 0.064623
Percentage
|  | UWSIF01 (Liver) | 12.9322 | 12.88
13C Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -29.8462 | -28.279
|  | UWSIF 39- UW Glutamic 2 | 22.2504 | 24.362
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF01 (Liver) | -17.7734 | -17.75 | 0.023433
Percentage
|  | UWSIF01 (Liver) | 50.0503 | 49.64
[6 empty] | Date Analyzed: |  | 2021-05-19
[6 empty] | Analyst: |  | sme
[6 empty] | Instrument: | Carlo Erba

## Tray 3
Date | Line | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Nitrogen % | Carbon % | C:N | Corr. d15N | Corr. d13C
05/19/21 | 10 | 20210048.137 | 921 | 0.846 | 3401 | 11.977 | 716 | -31.01 | 12.7789 | 40.8389 | 3.1958 | 10.2223 | -29.4226
05/19/21 | 11 | 20210048.138 | 923 | 0.842 | 3675 | 12.085 | 884 | -33.528 | 13.7796 | 50.7746 | 3.6848 | 10.3302 | -31.965
05/19/21 | 12 | 20210048.138 | Replicate A11 | 0.829 | 3366 | 12.112 | 887 | -34.148 | 12.8465 | 51.782 | 4.0308 | 10.3571 | -32.591
05/19/21 | 13 | 20210048.139 | 925 | 0.801 | 3166 | 11.553 | 852 | -33.665 | 12.4896 | 51.4176 | 4.1168 | 9.799 | -32.1034
05/19/21 | 14 | 20210048.140 | 927 | 0.839 | 3573 | 12.023 | 879 | -33.66 | 13.4429 | 50.6637 | 3.7688 | 10.2683 | -32.0983
05/19/21 | 15 | 20210048.141 | 929 | 0.843 | 3486 | 11.08 | 902 | -33.344 | 13.053 | 51.6324 | 3.9556 | 9.3268 | -31.7792
05/19/21 | 16 | 20210048.142 | 931 | 0.812 | 3675 | 10.748 | 799 | -30.192 | 14.3203 | 47.4286 | 3.312 | 8.9953 | -28.5967
05/19/21 | 17 | 20210048.143 | 933 | 0.83 | 3740 | 9.55 | 849 | -32.288 | 14.2201 | 49.3048 | 3.4673 | 7.7992 | -30.713
05/19/21 | 18 | 20210048.144 | 935 | 0.825 | 3828 | 11.789 | 827 | -32.141 | 14.6815 | 48.495 | 3.3031 | 10.0346 | -30.5646
05/19/21 | 19 | 20210048.145 | 937 | 0.84 | 2573 | 12.043 | 994 | -35.9 | 9.6718 | 57.4919 | 5.9443 | 10.2882 | -34.36
05/19/21 | 20 | 20210048.146 | 939 | 0.832 | 3539 | 11.959 | 876 | -33.604 | 13.4212 | 50.7269 | 3.7796 | 10.2044 | -32.0418
05/19/21 | 21 | 20210048.147 | 941 | 0.814 | 3149 | 12.63 | 909 | -34.17 | 12.2121 | 53.8261 | 4.4076 | 10.8743 | -32.6132
05/19/21 | 22 | 20210048.148 | 943 | 0.805 | 3678 | 12.729 | 821 | -31.437 | 14.432 | 49.1264 | 3.404 | 10.9731 | -29.8537
05/19/21 | 23 | 20210048.149 | 945 | 0.84 | 3436 | 12.043 | 915 | -34.253 | 12.8917 | 52.589 | 4.0793 | 10.2882 | -32.6971
05/19/21 | 24 | 20210048.150 | 947 | 0.829 | 3487 | 12.022 | 881 | -33.772 | 13.2954 | 51.1734 | 3.8489 | 10.2673 | -32.2114
05/19/21 | 25 | 20210048.151 | 949 | 0.826 | 3706 | 12.087 | 851 | -32.762 | 14.1459 | 49.5768 | 3.5047 | 10.3322 | -31.1916
05/19/21 | 26 | 20210048.152 | 951 | 0.843 | 3951 | 12.309 | 846 | -31.093 | 14.7276 | 48.1699 | 3.2707 | 10.5538 | -29.5064
05/19/21 | 27 | 20210048.153 | 953 | 0.827 | 3473 | 13.363 | 869 | -32.644 | 13.2715 | 50.7079 | 3.8208 | 11.6061 | -31.0724
05/19/21 | 28 | 20210048.154 | 955 | 0.803 | 3584 | 9.979 | 824 | -31.884 | 14.0425 | 49.3257 | 3.5126 | 8.2275 | -30.3051
05/19/21 | 29 | 20210048.155 | 957 | 0.846 | 4029 | 11.627 | 835 | -31.969 | 14.9942 | 47.4616 | 3.1653 | 9.8729 | -30.3909
05/19/21 | 30 | 20210048.156 | 959 | 0.846 | 4035 | 12.117 | 849 | -29.876 | 15.0442 | 48.3398 | 3.2132 | 10.3621 | -28.2776
05/19/21 | 31 | 20210048.157 | 961 | 0.824 | 3329 | 13.074 | 881 | -32.89 | 12.7703 | 51.5068 | 4.0333 | 11.3176 | -31.3208
05/19/21 | 32 | 20210048.158 | 963 | 0.827 | 3793 | 9.809 | 840 | -31.099 | 14.4959 | 48.8611 | 3.3707 | 8.0578 | -29.5125
05/19/21 | 33 | 20210048.158 | Replicate C8 | 0.821 | 3738 | 9.841 | 818 | -30.989 | 14.3791 | 47.9662 | 3.3358 | 8.0898 | -29.4014
05/19/21 | 40 | 20210048.159 | 965 | 0.825 | 3728 | 12.89 | 848 | -30.661 | 14.2494 | 49.4185 | 3.4681 | 11.1339 | -29.0702
05/19/21 | 41 | 20210048.160 | 967 | 0.807 | 3330 | 11.46 | 890 | -34.016 | 13.0246 | 53.0575 | 4.0736 | 9.7062 | -32.4578
05/19/21 | 42 | 20210048.161 | 969 | 0.809 | 3572 | 11.974 | 830 | -33.071 | 13.9622 | 49.3035 | 3.5312 | 10.2193 | -31.5036
05/19/21 | 43 | 20210048.162 | 971 | 0.808 | 3329 | 11.648 | 867 | -33.489 | 12.9792 | 51.5328 | 3.9704 | 9.8939 | -31.9256
05/19/21 | 44 | 20210048.163 | 973 | 0.826 | 3271 | 11.517 | 909 | -34.023 | 12.4862 | 52.9922 | 4.2441 | 9.7631 | -32.4648
05/19/21 | 45 | 20210048.164 | 975 | 0.833 | 3733 | 11.822 | 852 | -32.845 | 14.1257 | 49.1767 | 3.4814 | 10.0676 | -31.2754
05/19/21 | 46 | 20210048.165 | 749 | 0.842 | 4079 | 11.936 | 838 | -31.845 | 15.2703 | 47.7808 | 3.129 | 10.1814 | -30.2657
05/19/21 | 47 | 20210048.166 | 751 | 0.808 | 3503 | 11.085 | 832 | -31.533 | 13.6797 | 49.4367 | 3.6139 | 9.3318 | -29.9507
05/19/21 | 48 | 20210048.167 | 753 | 0.827 | 3552 | 12.24 | 844 | -32.914 | 13.5426 | 49.0354 | 3.6208 | 10.4849 | -31.3451
05/19/21 | 49 | 20210048.168 | 755 | 0.808 | 3639 | 12.506 | 821 | -32.251 | 14.1902 | 48.687 | 3.431 | 10.7505 | -30.6756
05/19/21 | 50 | 20210048.169 | 757 | 0.838 | 3540 | 11.715 | 877 | -32.679 | 13.308 | 50.3187 | 3.7811 | 9.9608 | -31.1078
05/19/21 | 51 | 20210048.170 | 759 | 0.84 | 3943 | 12.063 | 846 | -29.339 | 14.7876 | 48.3501 | 3.2696 | 10.3082 | -27.7354
05/19/21 | 52 | 20210048.171 | 761 | 0.816 | 2699 | 12.088 | 932 | -33.478 | 10.4086 | 54.979 | 5.2821 | 10.3332 | -31.9145
05/19/21 | 53 | 20210048.172 | 763 | 0.803 | 3527 | 12.382 | 820 | -29.649 | 13.8225 | 48.9367 | 3.5404 | 10.6267 | -28.0484
05/19/21 | 54 | 20210048.173 | 765 | 0.817 | 3676 | 11.275 | 818 | -31.919 | 14.1585 | 47.9658 | 3.3878 | 9.5215 | -30.3404
05/19/21 | 55 | 20210048.174 | 767 | 0.827 | 3747 | 11.866 | 825 | -32.187 | 14.2313 | 47.7739 | 3.3569 | 10.1115 | -30.611
05/19/21 | 56 | 20210048.175 | 769 | 0.824 | 3328 | 11.862 | 851 | -33.38 | 12.6803 | 49.5222 | 3.9054 | 10.1075 | -31.8156
05/19/21 | 57 | 20210048.176 | 771 | 0.837 | 3801 | 12.088 | 831 | -29.962 | 14.2434 | 47.4758 | 3.3332 | 10.3332 | -28.3644
05/19/21 | 58 | 20210048.177 | 773 | 0.83 | 3793 | 11.284 | 830 | -31.856 | 14.3383 | 47.9216 | 3.3422 | 9.5304 | -30.2768
05/19/21 | 59 | 20210048.178 | 775 | 0.82 | 3233 | 11.987 | 906 | -33.734 | 12.3557 | 53.0346 | 4.2923 | 10.2323 | -32.173
05/19/21 | 60 | 20210048.178 | Replicate E11 | 0.836 | 3739 | 12.051 | 883 | -32.671 | 14.0262 | 50.7223 | 3.6163 | 10.2962 | -31.0997
05/19/21 | 61 | 20210048.179 | 777 | 0.845 | 3927 | 11.57 | 830 | -30.568 | 14.5829 | 47.0771 | 3.2282 | 9.816 | -28.9763
05/19/21 | 62 | 20210048.180 | 779 | 0.85 | 3312 | 11.1 | 903 | -32.516 | 12.2078 | 51.0175 | 4.1791 | 9.3467 | -30.9432
05/19/21 | 63 | 20210048.181 | 781 | 0.81 | 3642 | 13.238 | 822 | -32.603 | 14.1266 | 48.5731 | 3.4384 | 11.4813 | -31.031
05/19/21 | 70 | 20210048.182 | 783 | 0.808 | 3525 | 12.036 | 817 | -29.733 | 13.6847 | 48.3642 | 3.5342 | 10.2812 | -28.1332
05/19/21 | 71 | 20210048.183 | 785 | 0.801 | 3042 | 12.168 | 875 | -32.422 | 11.8994 | 52.4117 | 4.4046 | 10.413 | -30.8483
05/19/21 | 72 | 20210048.184 | 787 | 0.815 | 3420 | 11.486 | 845 | -32.886 | 13.1625 | 49.6648 | 3.7732 | 9.7321 | -31.3168
05/19/21 | 73 | 20210048.185 | 789 | 0.805 | 3737 | 10.734 | 798 | -30.796 | 14.5529 | 47.3997 | 3.2571 | 8.9813 | -29.2065
05/19/21 | 74 | 20210048.186 | 791 | 0.848 | 3610 | 10.47 | 840 | -32.021 | 13.3536 | 47.4225 | 3.5513 | 8.7177 | -30.4434
05/19/21 | 75 | 20210048.187 | 793 | 0.83 | 3535 | 12.036 | 851 | -32.26 | 13.3527 | 49.1869 | 3.6837 | 10.2812 | -30.6847
05/19/21 | 76 | 20210048.188 | 795 | 0.845 | 3770 | 11.855 | 867 | -32.324 | 13.9805 | 49.2479 | 3.5226 | 10.1005 | -30.7493
05/19/21 | 77 | 20210048.189 | 797 | 0.826 | 3642 | 12.351 | 845 | -32.25 | 13.8267 | 48.9535 | 3.5405 | 10.5957 | -30.6746
05/19/21 | 78 | 20210048.190 | 799 | 0.845 | 3824 | 10.523 | 812 | -30.995 | 14.1741 | 45.9723 | 3.2434 | 8.7707 | -29.4075
05/19/21 | 79 | 20210048.191 | 801 | 0.848 | 3465 | 11.878 | 908 | -33.069 | 12.787 | 51.4494 | 4.0236 | 10.1235 | -31.5016
05/19/21 | 80 | 20210048.192 | 803 | 0.833 | 3778 | 11.932 | 824 | -31.12 | 14.2302 | 47.403 | 3.3312 | 10.1774 | -29.5337
05/19/21 | 81 | 20210048.193 | 805 | 0.843 | 3974 | 12.232 | 845 | -31.869 | 14.7759 | 48.0091 | 3.2491 | 10.4769 | -30.2899
05/19/21 | 82 | 20210048.194 | 807 | 0.818 | 3464 | 12.681 | 841 | -32.939 | 13.2718 | 49.2225 | 3.7088 | 10.9252 | -31.3703
05/20/21 | 83 | 20210048.195 | 809 | 0.803 | 3549 | 12.029 | 819 | -31.108 | 13.9075 | 48.8085 | 3.5095 | 10.2743 | -29.5215
05/20/21 | 84 | 20210048.196 | 811 | 0.843 | 3949 | 11.651 | 859 | -31.101 | 14.6545 | 48.807 | 3.3305 | 9.8969 | -29.5145
05/20/21 | 85 | 20210048.197 | 813 | 0.831 | 3664 | 9.272 | 829 | -30.328 | 13.8199 | 47.8041 | 3.4591 | 7.5217 | -28.734
05/20/21 | 86 | 20210048.198 | 815 | 0.814 | 2726 | 11.121 | 940 | -34.145 | 10.4787 | 55.5948 | 5.3055 | 9.3677 | -32.588
05/20/21 | 87 | 20210048.199 | 817 | 0.802 | 3589 | 12.201 | 801 | -32.236 | 14.0605 | 47.9044 | 3.407 | 10.446 | -30.6605
05/20/21 | 88 | 20210048.200 | 819 | 0.825 | 3726 | 11.63 | 830 | -32.44 | 14.1569 | 48.1518 | 3.4013 | 9.8759 | -30.8665
05/20/21 | 89 | 20210048.201 | 821 | 0.831 | 3450 | 11.06 | 887 | -31.805 | 13.0082 | 51.3104 | 3.9445 | 9.3068 | -30.2253
05/20/21 | 90 | 20210048.202 | 823 | 0.81 | 3555 | 13.059 | 819 | -32.344 | 13.7622 | 48.4566 | 3.521 | 11.3026 | -30.7695
05/20/21 | 91 | 20210048.203 | 825 | 0.837 | 3640 | 13.358 | 848 | -32.148 | 13.6014 | 48.4742 | 3.5639 | 11.6011 | -30.5716
05/20/21 | 92 | 20210048.204 | 827 | 0.836 | 3737 | 12.014 | 844 | -32.394 | 13.9885 | 48.3495 | 3.4564 | 10.2593 | -30.82
05/20/21 | 93 | 20210048.204 | Replicate H8 | 0.832 | 3676 | 12.002 | 842 | -32.393 | 13.8373 | 48.5015 | 3.5051 | 10.2473 | -30.819
Reference Check
05/19/21 | 8 | 01-UWSIF-Liver-20210048.31 | 01-UWSIF-Liver- | 0.811 | 3340 | 8.509 | 844 | -19.515 | 13.101 | 50.3372 | 3.8422 | 6.7599 | -17.8161
05/19/21 | 9 | 01-UWSIF-Liver-20210048.32 | 01-UWSIF-Liver- | 0.719 | 2940 | 8.452 | 747 | -19.55 | 13.0377 | 50.1599 | 3.8473 | 6.703 | -17.8514
05/19/21 | 38 | 01-UWSIF-Liver-20210048.33 | 01-UWSIF-Liver- | 0.75 | 3087 | 8.461 | 778 | -19.317 | 13.0074 | 49.7454 | 3.8244 | 6.712 | -17.6162
05/19/21 | 39 | 01-UWSIF-Liver-20210048.34 | 01-UWSIF-Liver- | 0.79 | 3251 | 8.439 | 817 | -19.31 | 12.9985 | 49.7138 | 3.8246 | 6.69 | -17.6091
05/19/21 | 68 | 01-UWSIF-Liver-20210048.35 | 01-UWSIF-Liver- | 0.771 | 3149 | 8.504 | 798 | -19.47 | 12.8204 | 49.4922 | 3.8604 | 6.7549 | -17.7707
05/19/21 | 69 | 01-UWSIF-Liver-20210048.36 | 01-UWSIF-Liver- | 0.765 | 3126 | 8.477 | 790 | -19.398 | 12.8336 | 49.4161 | 3.8505 | 6.7279 | -17.698
05/20/21 | 98 | 01-UWSIF-Liver-20210048.37 | 01-UWSIF-Liver- | 0.73 | 2946 | 8.508 | 751 | -19.579 | 12.6994 | 49.2395 | 3.8773 | 6.7589 | -17.8807
05/20/21 | 99 | 01-UWSIF-Liver-20210048.38 | 01-UWSIF-Liver- | 0.723 | 2929 | 8.43 | 745 | -19.498 | 12.7474 | 49.3033 | 3.8677 | 6.681 | -17.7989
[5 empty] | average | 8.4725 |  | -19.4546 | 12.9057 | 49.6759 |  | 6.7234 | -17.7551
[5 empty] | std. dev. | 0.031816 |  | 0.10239 | 0.14862 | 0.39791 |  | 0.031766 | 0.10339
Reference Material
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 | 1249 | -2.831 | 357 | -29.73 | 9.5228 | 41.0425 | 4.3099 | -4.562 | -28.1302
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 | 2223 | -2.847 | 628 | -29.877 | 9.5195 | 40.627 | 4.2678 | -4.5779 | -28.2786
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 | 4785 | -2.893 | 1297 | -30.115 | 9.5211 | 40.8337 | 4.2887 | -4.6239 | -28.5189
05/19/21 | 34 | 36-UWSIF-UT Glut 1-20210048.36 | 36-UWSIF-UT Glut 1- | 0.801 | 2382 | -2.943 | 674 | -29.867 | 9.4866 | 40.3482 | 4.2532 | -4.6738 | -28.2685
05/19/21 | 35 | 36-UWSIF-UT Glut 1-20210048.37 | 36-UWSIF-UT Glut 1- | 0.786 | 2342 | -2.875 | 663 | -29.855 | 9.4862 | 40.5048 | 4.2699 | -4.6059 | -28.2564
05/19/21 | 64 | 36-UWSIF-UT Glut 1-20210048.38 | 36-UWSIF-UT Glut 1- | 0.813 | 2426 | -2.897 | 689 | -29.809 | 9.3888 | 40.331 | 4.2957 | -4.6279 | -28.2099
05/19/21 | 65 | 36-UWSIF-UT Glut 1-20210048.39 | 36-UWSIF-UT Glut 1- | 0.777 | 2310 | -2.858 | 654 | -29.891 | 9.3951 | 40.1745 | 4.2761 | -4.5889 | -28.2927
05/20/21 | 94 | 36-UWSIF-UT Glut 1-20210048.311 | 36-UWSIF-UT Glut 1- | 0.755 | 2214 | -2.888 | 632 | -29.902 | 9.2758 | 40.0269 | 4.3152 | -4.6189 | -28.3038
05/20/21 | 95 | 36-UWSIF-UT Glut 1-20210048.312 | 36-UWSIF-UT Glut 1- | 0.756 | 2218 | -2.934 | 635 | -29.848 | 9.2588 | 40.0261 | 4.323 | -4.6648 | -28.2493
[5 empty] | average | -2.8851 |  | -29.8771 | 9.4283 | 40.435 |  | -4.616 | -28.2787
[5 empty] | std. dev. | 0.037368 |  | 0.10315 | 0.10445 | 0.35175 |  | 0.037308 | 0.10415
05/19/21 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 39-UWSIF-UW Glut 2- | 0.8 | 2580 | 29.724 | 728 | 22.144 | 10.3266 | 44.131 | 4.2735 | 27.9409 | 24.247
05/19/21 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 39-UWSIF-UW Glut 2- | 0.829 | 2680 | 29.762 | 755 | 22.232 | 10.341 | 44.1252 | 4.267 | 27.9789 | 24.3359
05/19/21 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 39-UWSIF-UW Glut 2- | 0.779 | 2511 | 29.645 | 706 | 22.272 | 10.2773 | 43.6597 | 4.2482 | 27.8621 | 24.3762
05/19/21 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 39-UWSIF-UW Glut 2- | 0.726 | 2350 | 29.79 | 660 | 22.207 | 10.3461 | 43.753 | 4.229 | 28.0068 | 24.3106
05/19/21 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 39-UWSIF-UW Glut 2- | 0.766 | 2463 | 29.66 | 699 | 22.281 | 10.1668 | 43.6233 | 4.2908 | 27.877 | 24.3853
05/19/21 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 39-UWSIF-UW Glut 2- | 0.757 | 2434 | 29.59 | 691 | 22.319 | 10.1463 | 43.6093 | 4.298 | 27.8072 | 24.4237
05/20/21 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 39-UWSIF-UW Glut 2- | 0.76 | 2426 | 29.631 | 691 | 22.346 | 10.0845 | 43.5184 | 4.3154 | 27.8481 | 24.451
05/20/21 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 39-UWSIF-UW Glut 2- | 0.706 | 2249 | 29.565 | 641 | 22.26 | 10.0794 | 43.4296 | 4.3088 | 27.7822 | 24.3641
[5 empty] | average | 29.6709 |  | 22.2576 | 10.221 | 43.7312 |  | 27.8879 | 24.3617
[5 empty] | std. dev. | 0.080574 |  | 0.06377 | 0.11436 | 0.26278 |  | 0.080445 | 0.064388
15N Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -2.8851 | -4.616
|  | UWSIF 39- UW Glutamic 2 | 29.6709 | 27.888
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF01 (Liver) | 6.7234 | 6.81 | 0.086556
Percentage
|  | UWSIF01 (Liver) | 12.9057 | 12.88
13C Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -29.8771 | -28.279
|  | UWSIF 39- UW Glutamic 2 | 22.2576 | 24.362
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF01 (Liver) | -17.7551 | -17.75 | 0.0051349
Percentage
|  | UWSIF01 (Liver) | 49.6759 | 49.64
[6 empty] | Date Analyzed: |  | 2021-05-20
[6 empty] | Analyst: |  | gh
[6 empty] | Instrument: | Carlo Erba

## Tray 4
Date | Line | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Nitrogen % | Carbon % | C:N | Corr. d15N | Corr. d13C
05/20/21 | 10 | 20210048.205 | 829 | 0.844 | 3445 | 11.884 | 836 | -31.834 | 13.3758 | 49.778 | 3.7215 | 10.0768 | -30.2236
05/20/21 | 11 | 20210048.206 | 831 | 0.807 | 2736 | 11.556 | 880 | -34.433 | 11.1422 | 54.8929 | 4.9266 | 9.7486 | -32.8444
05/20/21 | 12 | 20210048.206 | Replicate A11 | 0.815 | 2690 | 11.582 | 867 | -34.482 | 10.8072 | 53.451 | 4.9459 | 9.7746 | -32.8938
05/20/21 | 13 | 20210048.207 | 833 | 0.839 | 3641 | 11.888 | 837 | -31.448 | 14.2282 | 50.0282 | 3.5161 | 10.0808 | -29.8344
05/20/21 | 14 | 20210048.208 | 835 | 0.828 | 3312 | 12.12 | 833 | -32.786 | 13.1082 | 50.5279 | 3.8547 | 10.3128 | -31.1836
05/20/21 | 15 | 20210048.209 | 837 | 0.831 | 3319 | 12.522 | 849 | -33.494 | 13.0509 | 51.1799 | 3.9216 | 10.715 | -31.8975
05/20/21 | 16 | 20210048.210 | 839 | 0.805 | 3182 | 11.818 | 874 | -32.492 | 12.9241 | 54.3287 | 4.2037 | 10.0107 | -30.8871
05/20/21 | 17 | 20210048.211 | 841 | 0.825 | 3319 | 11.895 | 845 | -31.71 | 13.1665 | 51.4145 | 3.9049 | 10.0878 | -30.0986
05/20/21 | 18 | 20210048.212 | 843 | 0.822 | 3437 | 11.733 | 814 | -32.307 | 13.6307 | 49.4706 | 3.6294 | 9.9257 | -30.7006
05/20/21 | 19 | 20210048.213 | 845 | 0.802 | 3635 | 10.58 | 755 | -31.063 | 14.7286 | 46.7683 | 3.1753 | 8.7722 | -29.4461
05/20/21 | 20 | 20210048.214 | 002A | 0.827 | 2421 | 5.741 | 713 | -35.04 | 9.6423 | 43.0279 | 4.4624 | 3.9313 | -33.4565
05/20/21 | 21 | 20210048.215 | 002B | 0.844 | 2514 | 5.722 | 733 | -34.858 | 9.7709 | 43.3295 | 4.4345 | 3.9123 | -33.273
05/20/21 | 22 | 20210048.216 | 002C | 0.812 | 2470 | 5.615 | 714 | -35.111 | 9.9807 | 43.8315 | 4.3916 | 3.8052 | -33.5281
05/20/21 | 23 | 20210048.217 | 004A | 0.83 | 2824 | 8.532 | 822 | -35.118 | 11.1254 | 49.5493 | 4.4537 | 6.7234 | -33.5352
05/20/21 | 24 | 20210048.218 | 004B | 0.821 | 2407 | 8.536 | 860 | -36.132 | 9.59 | 52.393 | 5.4633 | 6.7274 | -34.5577
05/20/21 | 25 | 20210048.219 | 004C | 0.818 | 2046 | 7.962 | 875 | -37.046 | 8.2043 | 53.5824 | 6.531 | 6.1532 | -35.4794
05/20/21 | 26 | 20210048.220 | 006A | 0.843 | 2515 | 5.232 | 770 | -34.657 | 9.7748 | 45.5939 | 4.6644 | 3.4221 | -33.0703
05/20/21 | 27 | 20210048.221 | 006B | 0.82 | 2457 | 5.346 | 745 | -34.753 | 9.8236 | 45.3654 | 4.618 | 3.5361 | -33.1671
05/20/21 | 28 | 20210048.222 | 006C | 0.841 | 2457 | 5.173 | 753 | -34.91 | 9.595 | 44.6923 | 4.6579 | 3.3631 | -33.3254
05/20/21 | 29 | 20210048.223 | 008A | 0.834 | 2864 | 8.267 | 826 | -34.75 | 11.1724 | 49.3266 | 4.4151 | 6.4583 | -33.1641
05/20/21 | 30 | 20210048.224 | 008B | 0.85 | 2741 | 7.812 | 830 | -34.941 | 10.5121 | 48.6797 | 4.6308 | 6.0031 | -33.3567
05/20/21 | 31 | 20210048.225 | 008C | 0.81 | 2672 | 7.824 | 781 | -34.542 | 10.7626 | 47.9594 | 4.4561 | 6.0151 | -32.9544
05/20/21 | 32 | 20210048.226 | 010A | 0.849 | 2429 | 5.923 | 780 | -35.541 | 9.3856 | 46.0383 | 4.9052 | 4.1134 | -33.9617
05/20/21 | 33 | 20210048.226 | Replicate C8 | 0.843 | 2416 | 5.851 | 782 | -35.578 | 9.4198 | 46.4396 | 4.93 | 4.0413 | -33.9991
05/20/21 | 40 | 20210048.227 | 010B | 0.84 | 2527 | 5.928 | 789 | -35.398 | 9.8526 | 46.9475 | 4.765 | 4.1184 | -33.8175
05/20/21 | 41 | 20210048.228 | 010C | 0.814 | 2384 | 5.892 | 748 | -35.264 | 9.595 | 45.8262 | 4.776 | 4.0824 | -33.6824
05/20/21 | 42 | 20210048.229 | 012A | 0.827 | 2830 | 8.857 | 800 | -34.735 | 11.1343 | 48.1186 | 4.3216 | 7.0485 | -33.149
05/20/21 | 43 | 20210048.230 | 012B | 0.842 | 2717 | 8.145 | 782 | -34.36 | 10.5502 | 46.4063 | 4.3986 | 6.3363 | -32.7708
05/20/21 | 44 | 20210048.231 | 012C | 0.811 | 2829 | 8.507 | 787 | -34.482 | 11.3471 | 48.2655 | 4.2536 | 6.6984 | -32.8938
05/20/21 | 45 | 20210048.232 | 014A | 0.836 | 2905 | 8.546 | 759 | -34.031 | 11.3395 | 45.282 | 3.9933 | 6.7374 | -32.4391
05/20/21 | 46 | 20210048.233 | 014B | 0.843 | 2628 | 8.387 | 842 | -35.024 | 10.1419 | 49.7424 | 4.9046 | 6.5784 | -33.4404
05/20/21 | 47 | 20210048.234 | 014C | 0.815 | 2754 | 8.203 | 749 | -34.167 | 11.0009 | 45.6241 | 4.1473 | 6.3943 | -32.5762
05/20/21 | 48 | 20210048.235 | 016A | 0.817 | 2626 | 8.424 | 760 | -34.537 | 10.5088 | 46.3416 | 4.4098 | 6.6154 | -32.9493
05/20/21 | 49 | 20210048.236 | 016B | 0.835 | 2846 | 8.912 | 761 | -34.07 | 11.1154 | 45.3065 | 4.076 | 7.1036 | -32.4784
05/20/21 | 50 | 20210048.237 | 016C | 0.839 | 2824 | 8.845 | 813 | -34.686 | 10.9536 | 48.2336 | 4.4034 | 7.0365 | -33.0996
05/20/21 | 51 | 20210048.238 | 018A | 0.819 | 2830 | 8.899 | 795 | -34.758 | 11.2371 | 48.1816 | 4.2877 | 7.0906 | -33.1722
05/20/21 | 52 | 20210048.239 | 018B | 0.841 | 2613 | 9.105 | 844 | -35.373 | 10.1029 | 49.9366 | 4.9428 | 7.2966 | -33.7923
05/20/21 | 53 | 20210048.240 | 018C | 0.848 | 2728 | 9.701 | 851 | -35.227 | 10.4402 | 49.882 | 4.7779 | 7.8929 | -33.6451
Reference Check
05/20/21 | 8 | 01-UWSIF-Liver-20210048.41 | 01-UWSIF-Liver- | 0.751 | 2961 | 8.514 | 750 | -19.534 | 12.9981 | 50.1285 | 3.8566 | 6.7054 | -17.8203
05/20/21 | 9 | 01-UWSIF-Liver-20210048.42 | 01-UWSIF-Liver- | 0.809 | 3194 | 8.584 | 808 | -19.538 | 12.9926 | 50.145 | 3.8595 | 6.7754 | -17.8243
05/20/21 | 38 | 01-UWSIF-Liver-20210048.43 | 01-UWSIF-Liver- | 0.783 | 3130 | 8.554 | 790 | -19.547 | 12.9923 | 50.1568 | 3.8605 | 6.7454 | -17.8334
05/20/21 | 39 | 01-UWSIF-Liver-20210048.44 | 01-UWSIF-Liver- | 0.727 | 2885 | 8.59 | 731 | -19.51 | 12.9464 | 49.9661 | 3.8594 | 6.7814 | -17.7961
05/20/21 | 58 | 01-UWSIF-Liver-20210048.47 | 01-UWSIF-Liver- | 0.724 | 2878 | 8.55 | 730 | -19.537 | 12.9481 | 50.0556 | 3.8659 | 6.7414 | -17.8233
05/21/21 | 59 | 01-UWSIF-Liver-20210048.48 | 01-UWSIF-Liver- | 0.723 | 2858 | 8.573 | 726 | -19.47 | 12.8789 | 49.828 | 3.8689 | 6.7644 | -17.7557
[5 empty] | average | 8.5608 |  | -19.5227 | 12.9594 | 50.0467 |  | 6.7523 | -17.8089
[5 empty] | std. dev. | 0.027903 |  | 0.028619 | 0.045717 | 0.12875 |  | 0.027914 | 0.02886
Reference Material
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 | 1339 | -2.815 | 382 | -29.917 | 9.5695 | 40.9744 | 4.2818 | -4.6281 | -28.2905
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 | 2291 | -2.753 | 650 | -29.767 | 9.4851 | 40.6739 | 4.2882 | -4.5661 | -28.1392
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 | 4769 | -2.876 | 1284 | -30.028 | 9.5284 | 40.828 | 4.2849 | -4.6892 | -28.4024
05/20/21 | 34 | 36-UWSIF-UT Glut 1-20210048.46 | 36-UWSIF-UT Glut 1- | 0.73 | 2106 | -2.824 | 598 | -29.87 | 9.5073 | 40.5561 | 4.2658 | -4.6371 | -28.2431
05/20/21 | 35 | 36-UWSIF-UT Glut 1-20210048.47 | 36-UWSIF-UT Glut 1- | 0.798 | 2298 | -2.736 | 651 | -29.903 | 9.4423 | 40.5181 | 4.2911 | -4.5491 | -28.2764
05/20/21 | 54 | 36-UWSIF-UT Glut 1-20210048.411 | 36-UWSIF-UT Glut 1- | 0.851 | 2452 | -2.819 | 696 | -29.886 | 9.4194 | 40.587 | 4.3089 | -4.6321 | -28.2592
05/20/21 | 55 | 36-UWSIF-UT Glut 1-20210048.412 | 36-UWSIF-UT Glut 1- | 0.83 | 2390 | -2.798 | 678 | -29.966 | 9.4571 | 40.6118 | 4.2943 | -4.6111 | -28.3399
[5 empty] | average | -2.803 |  | -29.9053 | 9.487 | 40.6785 |  | -4.6161 | -28.2787
[5 empty] | std. dev. | 0.046876 |  | 0.081269 | 0.052254 | 0.16505 |  | 0.046894 | 0.081951
05/20/21 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 39-UWSIF-UW Glut 2- | 0.717 | 2213 | 29.706 | 629 | 22.375 | 10.2508 | 44.0368 | 4.2959 | 27.9059 | 24.4408
05/20/21 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 39-UWSIF-UW Glut 2- | 0.807 | 2510 | 29.772 | 714 | 22.314 | 10.2806 | 44.2723 | 4.3064 | 27.9719 | 24.3792
05/20/21 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 39-UWSIF-UW Glut 2- | 0.835 | 2618 | 29.682 | 737 | 22.128 | 10.2421 | 43.9303 | 4.2892 | 27.8819 | 24.1917
05/20/21 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 39-UWSIF-UW Glut 2- | 0.722 | 2243 | 29.67 | 636 | 22.192 | 10.2443 | 43.8964 | 4.285 | 27.8699 | 24.2562
05/20/21 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 39-UWSIF-UW Glut 2- | 0.779 | 2415 | 29.663 | 685 | 22.381 | 10.2097 | 43.9331 | 4.3031 | 27.8629 | 24.4468
05/20/21 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 39-UWSIF-UW Glut 2- | 0.855 | 2675 | 29.64 | 758 | 22.39 | 10.2211 | 44.0974 | 4.3143 | 27.8399 | 24.4559
[5 empty] | average | 29.6888 |  | 22.2967 | 10.2414 | 44.0277 |  | 27.8887 | 24.3618
[5 empty] | std. dev. | 0.046184 |  | 0.11105 | 0.024704 | 0.14172 |  | 0.046203 | 0.11199
15N Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -2.803 | -4.616
|  | UWSIF 39- UW Glutamic 2 | 29.6888 | 27.888
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF01 (Liver) | 6.7523 | 6.81 | 0.057742
Percentage
|  | UWSIF01 (Liver) | 12.9594 | 12.88
13C Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -29.9053 | -28.279
|  | UWSIF 39- UW Glutamic 2 | 22.2967 | 24.362
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF01 (Liver) | -17.8089 | -17.75 | 0.058857
Percentage
|  | UWSIF01 (Liver) | 50.0467 | 49.64
[6 empty] | Date Analyzed: |  | 2021-05-21
[6 empty] | Analyst: |  | gh
[6 empty] | Instrument: | Carlo Erba

## Original 1
Date | Row | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Component | Peak Nr | Amt% | Area All | Comment | Preparation | Area 44 | Area 28 | Is Ref _ | rR 45CO2/44CO2 | rR 29N2/28N2 | R 13C/12C | R 15N/14N | AT% 13C/12C | AT% 15N/14N | BGD 28 | BGD 29 | BGD 30 | BGD 44 | BGD 45 | BGD 46 | Time Code | Sample Dilution
05/17/21 | 1 | 36-UWSIF-UT Glut 1-20210048.11 | 36-UWSIF-UT Glut 1- | 0.735 | 6892 | 0.148 |  |  |  | 1 | 26.3728 | 127.2 | JOB: 2021-0048 Tray 1 Nelson |  |  | 126.285 | 0 |  | 0.72432 |  | 0.0036787 |  | 0.36653 | BGD 28: 42.0; BGD 29: 32.2; BGD 30: 958.7 | 2021/05/17 14:07:13 | 0
05/17/21 | 1 | 36-UWSIF-UT Glut 1-20210048.11 | 36-UWSIF-UT Glut 1- | 0.735 | 6900 | 0 |  |  |  | 2 | 26.3761 | 127.217 | JOB: 2021-0048 Tray 1 Nelson |  |  | 126.302 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.0; BGD 29: 31.5; BGD 30: 1066.4 | 2021/05/17 14:07:13 | 0
05/17/21 | 1 | 36-UWSIF-UT Glut 1-20210048.11 | 36-UWSIF-UT Glut 1- | 0.735 | 2403 | -2.762 |  |  | Nitrogen | 3 | 12.3739 | 58.147 | JOB: 2021-0048 Tray 1 Nelson |  |  | 57.73 | 0 |  | 0.72222 |  | 0.003668 |  | 0.36546 | BGD 28: 42.5; BGD 29: 32.5; BGD 30: 1021.0 | 2021/05/17 14:07:13 | 0
05/17/21 | 1 | 36-UWSIF-UT Glut 1-20210048.11 | 36-UWSIF-UT Glut 1- | 0.735 |  |  | 663 | -29.828 | Carbon | 4 | 44.1188 | 17.814 | JOB: 2021-0048 Tray 1 Nelson |  | 17.539 |  | 0 | 1.1517 |  | 0.010847 |  | 1.073 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.2 | 2021/05/17 14:07:13 | 95
05/17/21 | 1 | 36-UWSIF-UT Glut 1-20210048.11 | 36-UWSIF-UT Glut 1- | 0.735 |  |  | 5703 | -43.8 |  | 5 | 137.2088 | 106.686 | JOB: 2021-0048 Tray 1 Nelson |  | 105.066 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 4.4 | 2021/05/17 14:07:13 | 95
05/17/21 | 1 | 36-UWSIF-UT Glut 1-20210048.11 | 36-UWSIF-UT Glut 1- | 0.735 |  |  | 5707 | -44.078 |  | 6 | 137.157 | 107.036 | JOB: 2021-0048 Tray 1 Nelson |  | 105.411 |  | 0 | 1.1348 |  | 0.010687 |  | 1.0574 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 5.6 | 2021/05/17 14:07:13 | 95
05/17/21 | 2 | 36-UWSIF-UT Glut 1-20210048.12 | 36-UWSIF-UT Glut 1- | 0.813 | 6884 | 0.149 |  |  |  | 1 | 23.7898 | 126.904 | ANALYST: GH |  |  | 125.992 | 0 |  | 0.72436 |  | 0.0036787 |  | 0.36653 | BGD 28: 43.7; BGD 29: 33.5; BGD 30: 915.2 | 2021/05/17 14:16:58 | 0
05/17/21 | 2 | 36-UWSIF-UT Glut 1-20210048.12 | 36-UWSIF-UT Glut 1- | 0.813 | 6888 | 0 |  |  |  | 2 | 23.7964 | 126.942 | ANALYST: GH |  |  | 126.029 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.7; BGD 29: 32.8; BGD 30: 1026.7 | 2021/05/17 14:16:58 | 0
05/17/21 | 2 | 36-UWSIF-UT Glut 1-20210048.12 | 36-UWSIF-UT Glut 1- | 0.813 | 2512 | -2.76 |  |  | Nitrogen | 3 | 11.8649 | 61.755 | ANALYST: GH |  |  | 61.312 | 0 |  | 0.72225 |  | 0.003668 |  | 0.36546 | BGD 28: 43.7; BGD 29: 33.4; BGD 30: 992.2 | 2021/05/17 14:16:58 | 0
05/17/21 | 2 | 36-UWSIF-UT Glut 1-20210048.12 | 36-UWSIF-UT Glut 1- | 0.813 |  |  | 695 | -29.841 | Carbon | 4 | 42.2366 | 18.984 | ANALYST: GH |  | 18.691 |  | 0 | 1.1519 |  | 0.010847 |  | 1.073 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.5 | 2021/05/17 14:16:58 | 95
05/17/21 | 2 | 36-UWSIF-UT Glut 1-20210048.12 | 36-UWSIF-UT Glut 1- | 0.813 |  |  | 5692 | -43.8 |  | 5 | 124.07 | 106.488 | ANALYST: GH |  | 104.87 |  | 1 | 1.1356 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 4.0 | 2021/05/17 14:16:58 | 95
05/17/21 | 2 | 36-UWSIF-UT Glut 1-20210048.12 | 36-UWSIF-UT Glut 1- | 0.813 |  |  | 5693 | -44.054 |  | 6 | 124.0357 | 106.756 | ANALYST: GH |  | 105.135 |  | 0 | 1.1353 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.3; BGD 45: 1.6; BGD 46: 5.3 | 2021/05/17 14:16:58 | 95
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 | 6872 | 0.102 |  |  |  | 1 | 9.52 | 126.683 | SAMPLE COUNT: 0009 |  |  | 125.772 | 0 |  | 0.72436 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.2; BGD 29: 33.9; BGD 30: 896.1 | 2021/05/17 14:26:43 | 0
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 | 6883 | 0 |  |  |  | 2 | 9.52 | 126.78 | SAMPLE COUNT: 0009 |  |  | 125.868 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.2; BGD 29: 33.1; BGD 30: 1008.8 | 2021/05/17 14:26:43 | 0
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 | 1327 | -2.852 |  |  | Nitrogen | 3 | 9.52 | 33.041 | SAMPLE COUNT: 0009 |  |  | 32.804 | 0 |  | 0.72222 |  | 0.0036677 |  | 0.36543 | BGD 28: 44.1; BGD 29: 33.7; BGD 30: 974.0 | 2021/05/17 14:26:43 | 0
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 |  |  | 370 | -29.59 | Carbon | 4 | 40.81 | 9.888 | SAMPLE COUNT: 0009 |  | 9.735 |  | 0 | 1.1523 |  | 0.010849 |  | 1.0733 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.2 | 2021/05/17 14:26:43 | 95
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 |  |  | 5692 | -43.8 |  | 5 | 40.81 | 106.395 | SAMPLE COUNT: 0009 |  | 104.779 |  | 1 | 1.1357 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.4; BGD 45: 0.5; BGD 46: 3.6 | 2021/05/17 14:26:43 | 95
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 |  |  | 5689 | -44.068 |  | 6 | 40.81 | 106.838 | SAMPLE COUNT: 0009 |  | 105.216 |  | 0 | 1.1354 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.2; BGD 45: 1.5; BGD 46: 5.0 | 2021/05/17 14:26:43 | 95
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 | 6869 | 0.125 |  |  |  | 1 | 9.52 | 126.633 | HE PSI: 2200 |  |  | 125.722 | 0 |  | 0.72437 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.0; BGD 29: 33.7; BGD 30: 882.2 | 2021/05/17 14:36:27 | 0
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 | 6880 | 0 |  |  |  | 2 | 9.52 | 126.711 | HE PSI: 2200 |  |  | 125.8 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.0; BGD 29: 32.9; BGD 30: 995.4 | 2021/05/17 14:36:27 | 0
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 | 3291 | -2.836 |  |  | Nitrogen | 3 | 9.52 | 80.555 | HE PSI: 2200 |  |  | 79.977 | 0 |  | 0.72222 |  | 0.0036678 |  | 0.36544 | BGD 28: 44.1; BGD 29: 33.6; BGD 30: 967.8 | 2021/05/17 14:36:27 | 0
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 |  |  | 899 | -29.901 | Carbon | 4 | 40.81 | 24.785 | HE PSI: 2200 |  | 24.403 |  | 0 | 1.1518 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 5.3 | 2021/05/17 14:36:27 | 95
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 |  |  | 5687 | -43.8 |  | 5 | 40.81 | 106.454 | HE PSI: 2200 |  | 104.837 |  | 1 | 1.1357 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 4.0 | 2021/05/17 14:36:27 | 95
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 |  |  | 5693 | -44.049 |  | 6 | 40.81 | 106.727 | HE PSI: 2200 |  | 105.107 |  | 0 | 1.1354 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 5.2 | 2021/05/17 14:36:27 | 95
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 | 6865 | 0.125 |  |  |  | 1 | 9.52 | 126.579 | O2 PSI: 500 |  |  | 125.669 | 0 |  | 0.72436 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 887.6 | 2021/05/17 14:46:12 | 0
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 | 6870 | 0 |  |  |  | 2 | 9.52 | 126.687 | O2 PSI: 500 |  |  | 125.776 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.4; BGD 30: 1000.2 | 2021/05/17 14:46:12 | 0
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 | 10284 | -2.9 |  |  | Nitrogen | 3 | 9.52 | 238.593 | O2 PSI: 500 |  |  | 236.882 | 0 |  | 0.72217 |  | 0.0036675 |  | 0.36541 | BGD 28: 44.6; BGD 29: 34.1; BGD 30: 976.2 | 2021/05/17 14:46:12 | 0
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 |  |  | 2395 | -29.976 | Carbon | 4 | 40.81 | 74.631 | O2 PSI: 500 |  | 73.481 |  | 0 | 1.1515 |  | 0.010845 |  | 1.0729 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 5.8 | 2021/05/17 14:46:12 | 95
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 |  |  | 5694 | -43.8 |  | 5 | 40.81 | 106.431 | O2 PSI: 500 |  | 104.815 |  | 1 | 1.1356 |  | 0.010691 |  | 1.0577 |  | BGD 44: 1.2; BGD 45: 1.4; BGD 46: 5.1 | 2021/05/17 14:46:12 | 95
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 |  |  | 5686 | -43.945 |  | 6 | 40.81 | 106.731 | O2 PSI: 500 |  | 105.11 |  | 0 | 1.1354 |  | 0.010689 |  | 1.0576 |  | BGD 44: 1.8; BGD 45: 2.1; BGD 46: 6.1 | 2021/05/17 14:46:12 | 95
05/17/21 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 39-UWSIF-UW Glut 2- | 0.728 | 6868 | 0.127 |  |  |  | 1 | 24.192 | 126.749 | MASS 28: 42 |  |  | 125.838 | 0 |  | 0.72434 |  | 0.0036787 |  | 0.36652 | BGD 28: 46.1; BGD 29: 35.8; BGD 30: 917.3 | 2021/05/17 14:55:57 | 0
05/17/21 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 39-UWSIF-UW Glut 2- | 0.728 | 6866 | 0 |  |  |  | 2 | 24.1774 | 126.673 | MASS 28: 42 |  |  | 125.762 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 44.7; BGD 29: 34.8; BGD 30: 1029.3 | 2021/05/17 14:55:57 | 0
05/17/21 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 39-UWSIF-UW Glut 2- | 0.728 | 2501 | 29.606 |  |  | Nitrogen | 3 | 11.7447 | 61.705 | MASS 28: 42 |  |  | 61.248 | 0 |  | 0.74569 |  | 0.0037871 |  | 0.37728 | BGD 28: 45.5; BGD 29: 35.2; BGD 30: 997.6 | 2021/05/17 14:55:57 | 0
05/17/21 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 39-UWSIF-UW Glut 2- | 0.728 |  |  | 690 | 22.358 | Carbon | 4 | 51.0038 | 18.809 | MASS 28: 42 |  | 18.509 |  | 0 | 1.2095 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 5.7 | 2021/05/17 14:55:57 | 95
05/17/21 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 39-UWSIF-UW Glut 2- | 0.728 |  |  | 5693 | -43.8 |  | 5 | 104.0536 | 106.439 | MASS 28: 42 |  | 104.822 |  | 1 | 1.1357 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 4.1 | 2021/05/17 14:55:57 | 95
05/17/21 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 39-UWSIF-UW Glut 2- | 0.728 |  |  | 5686 | -44.053 |  | 6 | 103.769 | 106.69 | MASS 28: 42 |  | 105.07 |  | 0 | 1.1354 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 5.4 | 2021/05/17 14:55:57 | 95
05/17/21 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 39-UWSIF-UW Glut 2- | 0.747 | 6879 | 0.154 |  |  |  | 1 | 23.5345 | 126.523 | MASS 29: 32 |  |  | 125.613 | 0 |  | 0.72435 |  | 0.0036788 |  | 0.36653 | BGD 28: 44.8; BGD 29: 34.5; BGD 30: 903.2 | 2021/05/17 15:05:42 | 0
05/17/21 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 39-UWSIF-UW Glut 2- | 0.747 | 6871 | 0 |  |  |  | 2 | 23.5502 | 126.607 | MASS 29: 32 |  |  | 125.697 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1016.7 | 2021/05/17 15:05:42 | 0
05/17/21 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 39-UWSIF-UW Glut 2- | 0.747 | 2580 | 29.738 |  |  | Nitrogen | 3 | 11.7518 | 63.349 | MASS 29: 32 |  |  | 62.88 | 0 |  | 0.74578 |  | 0.0037876 |  | 0.37733 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 987.7 | 2021/05/17 15:05:42 | 0
05/17/21 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 39-UWSIF-UW Glut 2- | 0.747 |  |  | 713 | 22.251 | Carbon | 4 | 51.022 | 19.389 | MASS 29: 32 |  | 19.08 |  | 0 | 1.2093 |  | 0.011429 |  | 1.13 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.5 | 2021/05/17 15:05:42 | 95
05/17/21 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 39-UWSIF-UW Glut 2- | 0.747 |  |  | 5699 | -43.8 |  | 5 | 101.2896 | 106.545 | MASS 29: 32 |  | 104.927 |  | 1 | 1.1357 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 4.0 | 2021/05/17 15:05:42 | 95
05/17/21 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 39-UWSIF-UW Glut 2- | 0.747 |  |  | 5698 | -44.043 |  | 6 | 100.927 | 106.872 | MASS 29: 32 |  | 105.249 |  | 0 | 1.1354 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.3; BGD 45: 1.6; BGD 46: 5.4 | 2021/05/17 15:05:42 | 95
05/17/21 | 8 | 01-UWSIF-Liver-20210048.11 | 01-UWSIF-Liver- | 0.713 | 6878 | 0.116 |  |  |  | 1 | 24.6882 | 126.684 | MASS 30: 748 |  |  | 125.773 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.6; BGD 29: 34.3; BGD 30: 899.3 | 2021/05/17 15:15:27 | 0
05/17/21 | 8 | 01-UWSIF-Liver-20210048.11 | 01-UWSIF-Liver- | 0.713 | 6877 | 0 |  |  |  | 2 | 24.7162 | 126.827 | MASS 30: 748 |  |  | 125.915 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.5; BGD 30: 1013.7 | 2021/05/17 15:15:27 | 0
05/17/21 | 8 | 01-UWSIF-Liver-20210048.11 | 01-UWSIF-Liver- | 0.713 | 3148 | 8.575 |  |  | Nitrogen | 3 | 15.0002 | 77.134 | MASS 30: 748 |  |  | 76.575 | 0 |  | 0.73045 |  | 0.0037097 |  | 0.3696 | BGD 28: 44.5; BGD 29: 34.1; BGD 30: 986.5 | 2021/05/17 15:15:27 | 0
05/17/21 | 8 | 01-UWSIF-Liver-20210048.11 | 01-UWSIF-Liver- | 0.713 |  |  | 776 | -19.369 | Carbon | 4 | 57.6263 | 21.18 | MASS 30: 748 |  | 20.851 |  | 0 | 1.1633 |  | 0.010964 |  | 1.0845 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.6 | 2021/05/17 15:15:27 | 95
05/17/21 | 8 | 01-UWSIF-Liver-20210048.11 | 01-UWSIF-Liver- | 0.713 |  |  | 5695 | -43.8 |  | 5 | 106.22 | 106.458 | MASS 30: 748 |  | 104.841 |  | 1 | 1.1356 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 4.1 | 2021/05/17 15:15:27 | 95
05/17/21 | 8 | 01-UWSIF-Liver-20210048.11 | 01-UWSIF-Liver- | 0.713 |  |  | 5697 | -44.03 |  | 6 | 105.745 | 106.867 | MASS 30: 748 |  | 105.245 |  | 0 | 1.1354 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 5.4 | 2021/05/17 15:15:27 | 95
05/17/21 | 9 | 01-UWSIF-Liver-20210048.12 | 01-UWSIF-Liver- | 0.715 | 6869 | 0.121 |  |  |  | 1 | 24.5813 | 126.49 |  |  |  | 125.58 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36652 | BGD 28: 44.7; BGD 29: 34.3; BGD 30: 902.0 | 2021/05/17 15:25:12 | 0
05/17/21 | 9 | 01-UWSIF-Liver-20210048.12 | 01-UWSIF-Liver- | 0.715 | 6857 | 0 |  |  |  | 2 | 24.6001 | 126.587 |  |  |  | 125.676 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.5; BGD 30: 1016.2 | 2021/05/17 15:25:12 | 0
05/17/21 | 9 | 01-UWSIF-Liver-20210048.12 | 01-UWSIF-Liver- | 0.715 | 3137 | 8.605 |  |  | Nitrogen | 3 | 14.9224 | 76.95 |  |  |  | 76.392 | 0 |  | 0.73046 |  | 0.0037099 |  | 0.36961 | BGD 28: 44.6; BGD 29: 34.1; BGD 30: 989.2 | 2021/05/17 15:25:12 | 0
05/17/21 | 9 | 01-UWSIF-Liver-20210048.12 | 01-UWSIF-Liver- | 0.715 |  |  | 775 | -19.477 | Carbon | 4 | 57.3823 | 21.144 |  |  | 20.816 |  | 0 | 1.1632 |  | 0.010962 |  | 1.0844 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.6 | 2021/05/17 15:25:12 | 95
05/17/21 | 9 | 01-UWSIF-Liver-20210048.12 | 01-UWSIF-Liver- | 0.715 |  |  | 5703 | -43.8 |  | 5 | 105.8349 | 106.535 |  |  | 104.917 |  | 1 | 1.1356 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 4.1 | 2021/05/17 15:25:12 | 95
05/17/21 | 9 | 01-UWSIF-Liver-20210048.12 | 01-UWSIF-Liver- | 0.715 |  |  | 5698 | -44.028 |  | 6 | 105.4763 | 106.844 |  |  | 105.222 |  | 0 | 1.1354 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.3; BGD 45: 1.6; BGD 46: 5.5 | 2021/05/17 15:25:12 | 95
05/17/21 | 10 | 20210048.001 | 501 | 0.823 | 6880 | 0.126 |  |  |  | 1 | 21.3795 | 126.631 |  |  |  | 125.72 | 0 |  | 0.72431 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.6; BGD 29: 34.3; BGD 30: 904.6 | 2021/05/17 15:34:57 | 0
05/17/21 | 10 | 20210048.001 | 501 | 0.823 | 6882 | 0 |  |  |  | 2 | 21.3665 | 126.554 |  |  |  | 125.644 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.5; BGD 30: 1019.2 | 2021/05/17 15:34:57 | 0
05/17/21 | 10 | 20210048.001 | 501 | 0.823 | 3916 | 13.59 |  |  | Nitrogen | 3 | 16.1458 | 95.758 |  |  |  | 95.06 | 0 |  | 0.73406 |  | 0.0037282 |  | 0.37143 | BGD 28: 44.5; BGD 29: 34.0; BGD 30: 993.4 | 2021/05/17 15:34:57 | 0
05/17/21 | 10 | 20210048.001 | 501 | 0.823 |  |  | 850 | -32.506 | Carbon | 4 | 54.0767 | 23.311 |  |  | 22.952 |  | 0 | 1.1488 |  | 0.010817 |  | 1.0701 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.7 | 2021/05/17 15:34:57 | 95
05/17/21 | 10 | 20210048.001 | 501 | 0.823 |  |  | 5701 | -43.8 |  | 5 | 91.9462 | 106.535 |  |  | 104.917 |  | 1 | 1.1356 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 4.2 | 2021/05/17 15:34:57 | 95
05/17/21 | 10 | 20210048.001 | 501 | 0.823 |  |  | 5703 | -44.037 |  | 6 | 91.6697 | 106.81 |  |  | 105.188 |  | 0 | 1.1353 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 5.6 | 2021/05/17 15:34:57 | 95
05/17/21 | 11 | 20210048.002 | 503 | 0.805 | 6866 | 0.108 |  |  |  | 1 | 21.8421 | 126.542 |  |  |  | 125.632 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 910.6 | 2021/05/17 15:44:42 | 0
05/17/21 | 11 | 20210048.002 | 503 | 0.805 | 6862 | 0 |  |  |  | 2 | 21.826 | 126.449 |  |  |  | 125.54 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1024.9 | 2021/05/17 15:44:42 | 0
05/17/21 | 11 | 20210048.002 | 503 | 0.805 | 3833 | 12.77 |  |  | Nitrogen | 3 | 16.1405 | 93.641 |  |  |  | 92.959 | 0 |  | 0.73346 |  | 0.0037252 |  | 0.37113 | BGD 28: 44.6; BGD 29: 34.1; BGD 30: 999.4 | 2021/05/17 15:44:42 | 0
05/17/21 | 11 | 20210048.002 | 503 | 0.805 |  |  | 858 | -33.315 | Carbon | 4 | 55.6929 | 23.519 |  |  | 23.158 |  | 0 | 1.1479 |  | 0.010808 |  | 1.0692 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.8 | 2021/05/17 15:44:42 | 95
05/17/21 | 11 | 20210048.002 | 503 | 0.805 |  |  | 5696 | -43.8 |  | 5 | 94.0351 | 106.503 |  |  | 104.885 |  | 1 | 1.1356 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 4.4 | 2021/05/17 15:44:42 | 95
05/17/21 | 11 | 20210048.002 | 503 | 0.805 |  |  | 5700 | -44.041 |  | 6 | 93.6485 | 106.878 |  |  | 105.255 |  | 0 | 1.1353 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 5.6 | 2021/05/17 15:44:42 | 95
05/17/21 | 12 | 20210048.002 | Replicate A11 | 0.834 | 6865 | 0.123 |  |  |  | 1 | 21.0493 | 126.343 |  |  |  | 125.435 | 0 |  | 0.72428 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.7; BGD 29: 34.4; BGD 30: 915.4 | 2021/05/17 15:54:26 | 0
05/17/21 | 12 | 20210048.002 | Replicate A11 | 0.834 | 6866 | 0 |  |  |  | 2 | 21.0844 | 126.553 |  |  |  | 125.643 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.6; BGD 30: 1030.0 | 2021/05/17 15:54:26 | 0
05/17/21 | 12 | 20210048.002 | Replicate A11 | 0.834 | 4021 | 12.805 |  |  | Nitrogen | 3 | 16.3423 | 98.209 |  |  |  | 97.494 | 0 |  | 0.73347 |  | 0.0037253 |  | 0.37115 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 1004.5 | 2021/05/17 15:54:26 | 0
05/17/21 | 12 | 20210048.002 | Replicate A11 | 0.834 |  |  | 893 | -33.258 | Carbon | 4 | 55.6497 | 24.536 |  |  | 24.159 |  | 0 | 1.148 |  | 0.010808 |  | 1.0693 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.9 | 2021/05/17 15:54:26 | 95
05/17/21 | 12 | 20210048.002 | Replicate A11 | 0.834 |  |  | 5701 | -43.8 |  | 5 | 90.7318 | 106.537 |  |  | 104.919 |  | 1 | 1.1356 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 4.4 | 2021/05/17 15:54:26 | 95
05/17/21 | 12 | 20210048.002 | Replicate A11 | 0.834 |  |  | 5700 | -44.029 |  | 6 | 90.432 | 106.838 |  |  | 105.216 |  | 0 | 1.1353 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 5.7 | 2021/05/17 15:54:26 | 95
05/17/21 | 13 | 20210048.003 | 505 | 0.841 | 6863 | 0.088 |  |  |  | 1 | 20.8883 | 126.429 |  |  |  | 125.52 | 0 |  | 0.72426 |  | 0.0036785 |  | 0.3665 | BGD 28: 44.8; BGD 29: 34.5; BGD 30: 920.3 | 2021/05/17 16:04:11 | 0
05/17/21 | 13 | 20210048.003 | 505 | 0.841 | 6865 | 0 |  |  |  | 2 | 20.8717 | 126.329 |  |  |  | 125.42 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1034.6 | 2021/05/17 16:04:11 | 0
05/17/21 | 13 | 20210048.003 | 505 | 0.841 | 3989 | 10.204 |  |  | Nitrogen | 3 | 16.0565 | 97.305 |  |  |  | 96.598 | 0 |  | 0.73159 |  | 0.0037157 |  | 0.3702 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1009.4 | 2021/05/17 16:04:11 | 0
05/17/21 | 13 | 20210048.003 | 505 | 0.841 |  |  | 892 | -32.879 | Carbon | 4 | 55.1212 | 24.501 |  |  | 24.124 |  | 0 | 1.1484 |  | 0.010813 |  | 1.0697 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.0 | 2021/05/17 16:04:11 | 95
05/17/21 | 13 | 20210048.003 | 505 | 0.841 |  |  | 5699 | -43.8 |  | 5 | 90.0019 | 106.511 |  |  | 104.893 |  | 1 | 1.1356 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 4.5 | 2021/05/17 16:04:11 | 95
05/17/21 | 13 | 20210048.003 | 505 | 0.841 |  |  | 5696 | -44.024 |  | 6 | 89.719 | 106.798 |  |  | 105.177 |  | 0 | 1.1353 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 5.8 | 2021/05/17 16:04:11 | 95
05/17/21 | 14 | 20210048.004 | 507 | 0.846 | 6859 | 0.123 |  |  |  | 1 | 20.7697 | 126.459 |  |  |  | 125.549 | 0 |  | 0.72425 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 924.2 | 2021/05/17 16:13:56 | 0
05/17/21 | 14 | 20210048.004 | 507 | 0.846 | 6860 | 0 |  |  |  | 2 | 20.7492 | 126.334 |  |  |  | 125.426 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1039.8 | 2021/05/17 16:13:56 | 0
05/17/21 | 14 | 20210048.004 | 507 | 0.846 | 3816 | 12.526 |  |  | Nitrogen | 3 | 15.2817 | 93.176 |  |  |  | 92.498 | 0 |  | 0.73323 |  | 0.0037243 |  | 0.37105 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1012.6 | 2021/05/17 16:13:56 | 0
05/17/21 | 14 | 20210048.004 | 507 | 0.846 |  |  | 914 | -33.262 | Carbon | 4 | 55.8924 | 25.107 |  |  | 24.721 |  | 0 | 1.1479 |  | 0.010808 |  | 1.0693 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.0 | 2021/05/17 16:13:56 | 95
05/17/21 | 14 | 20210048.004 | 507 | 0.846 |  |  | 5698 | -43.8 |  | 5 | 89.4583 | 106.523 |  |  | 104.905 |  | 1 | 1.1355 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 4.5 | 2021/05/17 16:13:56 | 95
05/17/21 | 14 | 20210048.004 | 507 | 0.846 |  |  | 5689 | -44.031 |  | 6 | 89.2265 | 106.76 |  |  | 105.139 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 5.9 | 2021/05/17 16:13:56 | 95
05/17/21 | 15 | 20210048.005 | 509 | 0.83 | 6841 | 0.139 |  |  |  | 1 | 21.083 | 125.941 |  |  |  | 125.035 | 0 |  | 0.72425 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 927.4 | 2021/05/17 16:23:41 | 0
05/17/21 | 15 | 20210048.005 | 509 | 0.83 | 6861 | 0 |  |  |  | 2 | 21.1591 | 126.393 |  |  |  | 125.484 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1044.4 | 2021/05/17 16:23:41 | 0
05/17/21 | 15 | 20210048.005 | 509 | 0.83 | 4001 | 12.168 |  |  | Nitrogen | 3 | 16.367 | 97.887 |  |  |  | 97.174 | 0 |  | 0.73297 |  | 0.003723 |  | 0.37091 | BGD 28: 44.7; BGD 29: 34.3; BGD 30: 1019.1 | 2021/05/17 16:23:41 | 0
05/17/21 | 15 | 20210048.005 | 509 | 0.83 |  |  | 873 | -33.502 | Carbon | 4 | 54.8749 | 23.976 |  |  | 23.608 |  | 0 | 1.1476 |  | 0.010806 |  | 1.069 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.2 | 2021/05/17 16:23:41 | 95
05/17/21 | 15 | 20210048.005 | 509 | 0.83 |  |  | 5704 | -43.8 |  | 5 | 91.1384 | 106.567 |  |  | 104.949 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 4.6 | 2021/05/17 16:23:41 | 95
05/17/21 | 15 | 20210048.005 | 509 | 0.83 |  |  | 5694 | -44.029 |  | 6 | 90.8382 | 106.868 |  |  | 105.245 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 6.0 | 2021/05/17 16:23:41 | 95
05/17/21 | 16 | 20210048.006 | 511 | 0.847 | 6857 | 0.114 |  |  |  | 1 | 20.7235 | 126.327 |  |  |  | 125.418 | 0 |  | 0.72422 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 933.0 | 2021/05/17 16:33:26 | 0
05/17/21 | 16 | 20210048.006 | 511 | 0.847 | 6855 | 0 |  |  |  | 2 | 20.7358 | 126.401 |  |  |  | 125.493 | 1 |  | 0.72414 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1048.9 | 2021/05/17 16:33:26 | 0
05/17/21 | 16 | 20210048.006 | 511 | 0.847 | 4174 | 8.944 |  |  | Nitrogen | 3 | 16.7116 | 101.977 |  |  |  | 101.237 | 0 |  | 0.73062 |  | 0.0037111 |  | 0.36974 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1023.6 | 2021/05/17 16:33:26 | 0
05/17/21 | 16 | 20210048.006 | 511 | 0.847 |  |  | 879 | -32.542 | Carbon | 4 | 54.0592 | 24.132 |  |  | 23.761 |  | 0 | 1.1487 |  | 0.010816 |  | 1.0701 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.3 | 2021/05/17 16:33:26 | 95
05/17/21 | 16 | 20210048.006 | 511 | 0.847 |  |  | 5711 | -43.8 |  | 5 | 89.2027 | 106.677 |  |  | 105.057 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 4.7 | 2021/05/17 16:33:26 | 95
05/17/21 | 16 | 20210048.006 | 511 | 0.847 |  |  | 5702 | -44.034 |  | 6 | 88.9521 | 106.932 |  |  | 105.308 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 6.1 | 2021/05/17 16:33:26 | 95
05/17/21 | 17 | 20210048.007 | 513 | 0.832 | 6852 | 0.114 |  |  |  | 1 | 21.1178 | 126.45 |  |  |  | 125.541 | 0 |  | 0.72422 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 937.6 | 2021/05/17 16:43:11 | 0
05/17/21 | 17 | 20210048.007 | 513 | 0.832 | 6865 | 0 |  |  |  | 2 | 21.1106 | 126.407 |  |  |  | 125.498 | 1 |  | 0.72413 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1054.3 | 2021/05/17 16:43:11 | 0
05/17/21 | 17 | 20210048.007 | 513 | 0.832 | 3350 | 10.151 |  |  | Nitrogen | 3 | 13.6014 | 81.599 |  |  |  | 81.006 | 0 |  | 0.73148 |  | 0.0037155 |  | 0.37018 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1027.1 | 2021/05/17 16:43:11 | 0
05/17/21 | 17 | 20210048.007 | 513 | 0.832 |  |  | 941 | -33.499 | Carbon | 4 | 58.2892 | 25.909 |  |  | 25.511 |  | 0 | 1.1476 |  | 0.010806 |  | 1.069 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.3 | 2021/05/17 16:43:11 | 95
05/17/21 | 17 | 20210048.007 | 513 | 0.832 |  |  | 5702 | -43.8 |  | 5 | 90.8236 | 106.664 |  |  | 105.044 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 4.7 | 2021/05/17 16:43:11 | 95
05/17/21 | 17 | 20210048.007 | 513 | 0.832 |  |  | 5704 | -44.036 |  | 6 | 90.5374 | 106.95 |  |  | 105.326 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 6.1 | 2021/05/17 16:43:11 | 95
05/17/21 | 18 | 20210048.008 | 515 | 0.839 | 6864 | 0.094 |  |  |  | 1 | 20.9178 | 126.307 |  |  |  | 125.399 | 0 |  | 0.72421 |  | 0.0036785 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.6; BGD 30: 939.2 | 2021/05/17 16:52:56 | 0
05/17/21 | 18 | 20210048.008 | 515 | 0.839 | 6859 | 0 |  |  |  | 2 | 20.9278 | 126.367 |  |  |  | 125.458 | 1 |  | 0.72414 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1055.0 | 2021/05/17 16:52:56 | 0
05/17/21 | 18 | 20210048.008 | 515 | 0.839 | 4075 | 12.337 |  |  | Nitrogen | 3 | 16.4583 | 99.493 |  |  |  | 98.769 | 0 |  | 0.73307 |  | 0.0037236 |  | 0.37098 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1029.3 | 2021/05/17 16:52:56 | 0
05/17/21 | 18 | 20210048.008 | 515 | 0.839 |  |  | 862 | -32.778 | Carbon | 4 | 53.532 | 23.571 |  |  | 23.208 |  | 0 | 1.1485 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/17 16:52:56 | 95
05/17/21 | 18 | 20210048.008 | 515 | 0.839 |  |  | 5705 | -43.8 |  | 5 | 90.1109 | 106.618 |  |  | 104.999 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 4.8 | 2021/05/17 16:52:56 | 95
05/17/21 | 18 | 20210048.008 | 515 | 0.839 |  |  | 5701 | -44.033 |  | 6 | 89.8208 | 106.911 |  |  | 105.288 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 6.2 | 2021/05/17 16:52:56 | 95
05/17/21 | 19 | 20210048.009 | 517 | 0.84 | 6852 | 0.138 |  |  |  | 1 | 20.8826 | 126.245 |  |  |  | 125.338 | 0 |  | 0.72421 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.6; BGD 30: 943.7 | 2021/05/17 17:02:41 | 0
05/17/21 | 19 | 20210048.009 | 517 | 0.84 | 6853 | 0 |  |  |  | 2 | 20.9 | 126.349 |  |  |  | 125.441 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1059.8 | 2021/05/17 17:02:41 | 0
05/17/21 | 19 | 20210048.009 | 517 | 0.84 | 3905 | 12.444 |  |  | Nitrogen | 3 | 15.7071 | 95.083 |  |  |  | 94.391 | 0 |  | 0.73312 |  | 0.003724 |  | 0.37101 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1033.3 | 2021/05/17 17:02:41 | 0
05/17/21 | 19 | 20210048.009 | 517 | 0.84 |  |  | 903 | -33.224 | Carbon | 4 | 55.6319 | 24.744 |  |  | 24.363 |  | 0 | 1.148 |  | 0.010809 |  | 1.0693 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/17 17:02:41 | 95
05/17/21 | 19 | 20210048.009 | 517 | 0.84 |  |  | 5698 | -43.8 |  | 5 | 90.0385 | 106.583 |  |  | 104.964 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 4.8 | 2021/05/17 17:02:41 | 95
05/17/21 | 19 | 20210048.009 | 517 | 0.84 |  |  | 5703 | -44.029 |  | 6 | 89.7555 | 106.869 |  |  | 105.247 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 6.2 | 2021/05/17 17:02:41 | 95
05/17/21 | 20 | 20210048.010 | 519 | 0.826 | 6861 | 0.114 |  |  |  | 1 | 21.2572 | 126.367 |  |  |  | 125.458 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 947.3 | 2021/05/17 17:12:26 | 0
05/17/21 | 20 | 20210048.010 | 519 | 0.826 | 6865 | 0 |  |  |  | 2 | 21.2582 | 126.373 |  |  |  | 125.465 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1064.2 | 2021/05/17 17:12:26 | 0
05/17/21 | 20 | 20210048.010 | 519 | 0.826 | 3839 | 8.949 |  |  | Nitrogen | 3 | 15.7125 | 93.537 |  |  |  | 92.858 | 0 |  | 0.73059 |  | 0.0037111 |  | 0.36974 | BGD 28: 44.8; BGD 29: 34.2; BGD 30: 1037.8 | 2021/05/17 17:12:26 | 0
05/17/21 | 20 | 20210048.010 | 519 | 0.826 |  |  | 876 | -33.595 | Carbon | 4 | 55.2076 | 24.012 |  |  | 23.643 |  | 0 | 1.1475 |  | 0.010805 |  | 1.0689 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/17 17:12:26 | 95
05/17/21 | 20 | 20210048.010 | 519 | 0.826 |  |  | 5697 | -43.8 |  | 5 | 91.6115 | 106.536 |  |  | 104.918 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 4.9 | 2021/05/17 17:12:26 | 95
05/17/21 | 20 | 20210048.010 | 519 | 0.826 |  |  | 5698 | -44.01 |  | 6 | 91.1871 | 106.958 |  |  | 105.334 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 6.3 | 2021/05/17 17:12:26 | 95
05/17/21 | 21 | 20210048.011 | 521 | 0.832 | 6851 | 0.135 |  |  |  | 1 | 21.0678 | 126.152 |  |  |  | 125.245 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 951.7 | 2021/05/17 17:22:11 | 0
05/17/21 | 21 | 20210048.011 | 521 | 0.832 | 6853 | 0 |  |  |  | 2 | 21.0776 | 126.21 |  |  |  | 125.303 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1069.2 | 2021/05/17 17:22:11 | 0
05/17/21 | 21 | 20210048.011 | 521 | 0.832 | 3794 | 12.899 |  |  | Nitrogen | 3 | 15.3766 | 92.207 |  |  |  | 91.535 | 0 |  | 0.73344 |  | 0.0037256 |  | 0.37118 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1042.9 | 2021/05/17 17:22:11 | 0
05/17/21 | 21 | 20210048.011 | 521 | 0.832 |  |  | 896 | -33.506 | Carbon | 4 | 55.7587 | 24.523 |  |  | 24.145 |  | 0 | 1.1476 |  | 0.010806 |  | 1.069 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.6 | 2021/05/17 17:22:11 | 95
05/17/21 | 21 | 20210048.011 | 521 | 0.832 |  |  | 5693 | -43.8 |  | 5 | 91.0236 | 106.462 |  |  | 104.846 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.0 | 2021/05/17 17:22:11 | 95
05/17/21 | 21 | 20210048.011 | 521 | 0.832 |  |  | 5693 | -44.011 |  | 6 | 90.6767 | 106.811 |  |  | 105.19 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 6.3 | 2021/05/17 17:22:11 | 95
05/17/21 | 22 | 20210048.012 | 523 | 0.841 | 6846 | 0.111 |  |  |  | 1 | 20.7961 | 125.874 |  |  |  | 124.969 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 955.7 | 2021/05/17 17:31:56 | 0
05/17/21 | 22 | 20210048.012 | 523 | 0.841 | 6849 | 0 |  |  |  | 2 | 20.853 | 126.216 |  |  |  | 125.309 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1073.5 | 2021/05/17 17:31:56 | 0
05/17/21 | 22 | 20210048.012 | 523 | 0.841 | 3968 | 12.704 |  |  | Nitrogen | 3 | 15.9059 | 96.396 |  |  |  | 95.694 | 0 |  | 0.73331 |  | 0.0037249 |  | 0.37111 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1047.3 | 2021/05/17 17:31:56 | 0
05/17/21 | 22 | 20210048.012 | 523 | 0.841 |  |  | 896 | -33.052 | Carbon | 4 | 55.1163 | 24.498 |  |  | 24.121 |  | 0 | 1.1481 |  | 0.010811 |  | 1.0695 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.6 | 2021/05/17 17:31:56 | 95
05/17/21 | 22 | 20210048.012 | 523 | 0.841 |  |  | 5691 | -43.8 |  | 5 | 90.1282 | 106.382 |  |  | 104.766 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.1 | 2021/05/17 17:31:56 | 95
05/17/21 | 22 | 20210048.012 | 523 | 0.841 |  |  | 5690 | -44.026 |  | 6 | 89.8561 | 106.659 |  |  | 105.04 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 6.5 | 2021/05/17 17:31:56 | 95
05/17/21 | 23 | 20210048.013 | 525 | 0.817 | 6833 | 0.116 |  |  |  | 1 | 21.4108 | 125.896 |  |  |  | 124.991 | 0 |  | 0.72417 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 961.5 | 2021/05/17 17:41:41 | 0
05/17/21 | 23 | 20210048.013 | 525 | 0.817 | 6828 | 0 |  |  |  | 2 | 21.4205 | 125.953 |  |  |  | 125.047 | 1 |  | 0.72409 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1079.4 | 2021/05/17 17:41:41 | 0
05/17/21 | 23 | 20210048.013 | 525 | 0.817 | 4102 | 12.084 |  |  | Nitrogen | 3 | 16.9639 | 99.859 |  |  |  | 99.133 | 0 |  | 0.73284 |  | 0.0037226 |  | 0.37088 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1053.0 | 2021/05/17 17:41:41 | 0
05/17/21 | 23 | 20210048.013 | 525 | 0.817 |  |  | 834 | -31.685 | Carbon | 4 | 53.4135 | 22.763 |  |  | 22.412 |  | 0 | 1.1496 |  | 0.010826 |  | 1.071 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/17 17:41:41 | 95
05/17/21 | 23 | 20210048.013 | 525 | 0.817 |  |  | 5686 | -43.8 |  | 5 | 92.7901 | 106.367 |  |  | 104.752 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.1 | 2021/05/17 17:41:41 | 95
05/17/21 | 23 | 20210048.013 | 525 | 0.817 |  |  | 5691 | -44.024 |  | 6 | 92.4713 | 106.683 |  |  | 105.064 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 6.5 | 2021/05/17 17:41:41 | 95
05/17/21 | 24 | 20210048.014 | 527 | 0.842 | 6840 | 0.123 |  |  |  | 1 | 20.7916 | 125.996 |  |  |  | 125.09 | 0 |  | 0.72418 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 967.5 | 2021/05/17 17:51:26 | 0
05/17/21 | 24 | 20210048.014 | 527 | 0.842 | 6847 | 0 |  |  |  | 2 | 20.8109 | 126.112 |  |  |  | 125.205 | 1 |  | 0.72409 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1085.7 | 2021/05/17 17:51:26 | 0
05/17/21 | 24 | 20210048.014 | 527 | 0.842 | 3703 | 13.28 |  |  | Nitrogen | 3 | 14.8043 | 89.851 |  |  |  | 89.197 | 0 |  | 0.7337 |  | 0.003727 |  | 0.37132 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1059.3 | 2021/05/17 17:51:26 | 0
05/17/21 | 24 | 20210048.014 | 527 | 0.842 |  |  | 905 | -33.68 | Carbon | 4 | 55.5525 | 24.773 |  |  | 24.392 |  | 0 | 1.1474 |  | 0.010804 |  | 1.0688 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/17 17:51:26 | 95
05/17/21 | 24 | 20210048.014 | 527 | 0.842 |  |  | 5701 | -43.8 |  | 5 | 89.9364 | 106.468 |  |  | 104.852 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.2 | 2021/05/17 17:51:26 | 95
05/17/21 | 24 | 20210048.014 | 527 | 0.842 |  |  | 5696 | -44.021 |  | 6 | 89.627 | 106.783 |  |  | 105.162 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 6.6 | 2021/05/17 17:51:26 | 95
05/17/21 | 25 | 20210048.015 | 529 | 0.806 | 6843 | 0.112 |  |  |  | 1 | 21.7294 | 126.048 |  |  |  | 125.142 | 0 |  | 0.72418 |  | 0.0036786 |  | 0.36651 | BGD 28: 45.1; BGD 29: 34.6; BGD 30: 973.3 | 2021/05/17 18:01:10 | 0
05/17/21 | 25 | 20210048.015 | 529 | 0.806 | 6848 | 0 |  |  |  | 2 | 21.759 | 126.219 |  |  |  | 125.312 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 44.0; BGD 29: 33.8; BGD 30: 1092.6 | 2021/05/17 18:01:10 | 0
05/17/21 | 25 | 20210048.015 | 529 | 0.806 | 3956 | 9.147 |  |  | Nitrogen | 3 | 16.505 | 95.866 |  |  |  | 95.17 | 0 |  | 0.73072 |  | 0.0037118 |  | 0.36981 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1066.7 | 2021/05/17 18:01:10 | 0
05/17/21 | 25 | 20210048.015 | 529 | 0.806 |  |  | 853 | -31.415 | Carbon | 4 | 55.0433 | 23.222 |  |  | 22.864 |  | 0 | 1.15 |  | 0.010829 |  | 1.0713 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.0 | 2021/05/17 18:01:10 | 95
05/17/21 | 25 | 20210048.015 | 529 | 0.806 |  |  | 5697 | -43.8 |  | 5 | 93.9386 | 106.483 |  |  | 104.866 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.2 | 2021/05/17 18:01:10 | 95
05/17/21 | 25 | 20210048.015 | 529 | 0.806 |  |  | 5697 | -44.035 |  | 6 | 93.5352 | 106.875 |  |  | 105.253 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 6.7 | 2021/05/17 18:01:10 | 95
05/17/21 | 26 | 20210048.016 | 531 | 0.818 | 6844 | 0.101 |  |  |  | 1 | 21.4153 | 126.076 |  |  |  | 125.169 | 0 |  | 0.72417 |  | 0.0036786 |  | 0.36651 | BGD 28: 45.1; BGD 29: 34.6; BGD 30: 979.4 | 2021/05/17 18:10:55 | 0
05/17/21 | 26 | 20210048.016 | 531 | 0.818 | 6846 | 0 |  |  |  | 2 | 21.4327 | 126.178 |  |  |  | 125.271 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 44.0; BGD 29: 33.8; BGD 30: 1098.7 | 2021/05/17 18:10:55 | 0
05/17/21 | 26 | 20210048.016 | 531 | 0.818 | 4029 | 13.058 |  |  | Nitrogen | 3 | 16.509 | 97.31 |  |  |  | 96.602 | 0 |  | 0.73355 |  | 0.0037262 |  | 0.37124 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1073.1 | 2021/05/17 18:10:55 | 0
05/17/21 | 26 | 20210048.016 | 531 | 0.818 |  |  | 870 | -32.751 | Carbon | 4 | 55.1474 | 23.697 |  |  | 23.332 |  | 0 | 1.1485 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.0 | 2021/05/17 18:10:55 | 95
05/17/21 | 26 | 20210048.016 | 531 | 0.818 |  |  | 5684 | -43.8 |  | 5 | 92.5839 | 106.46 |  |  | 104.843 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.4 | 2021/05/17 18:10:55 | 95
05/17/21 | 26 | 20210048.016 | 531 | 0.818 |  |  | 5691 | -44.011 |  | 6 | 92.2236 | 106.816 |  |  | 105.194 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 6.8 | 2021/05/17 18:10:55 | 95
05/17/21 | 27 | 20210048.017 | 533 | 0.85 | 6843 | 0.105 |  |  |  | 1 | 20.6204 | 126.145 |  |  |  | 125.238 | 0 |  | 0.72417 |  | 0.0036786 |  | 0.36651 | BGD 28: 45.1; BGD 29: 34.7; BGD 30: 984.5 | 2021/05/17 18:20:40 | 0
05/17/21 | 27 | 20210048.017 | 533 | 0.85 | 6841 | 0 |  |  |  | 2 | 20.6043 | 126.047 |  |  |  | 125.14 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 44.0; BGD 29: 33.8; BGD 30: 1104.5 | 2021/05/17 18:20:40 | 0
05/17/21 | 27 | 20210048.017 | 533 | 0.85 | 3912 | 12.487 |  |  | Nitrogen | 3 | 15.3964 | 94.315 |  |  |  | 93.628 | 0 |  | 0.73314 |  | 0.0037241 |  | 0.37103 | BGD 28: 45.0; BGD 29: 34.4; BGD 30: 1077.8 | 2021/05/17 18:20:40 | 0
05/17/21 | 27 | 20210048.017 | 533 | 0.85 |  |  | 945 | -30.835 | Carbon | 4 | 56.901 | 25.822 |  |  | 25.424 |  | 0 | 1.1506 |  | 0.010835 |  | 1.0719 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/17 18:20:40 | 95
05/17/21 | 27 | 20210048.017 | 533 | 0.85 |  |  | 5698 | -43.8 |  | 5 | 89.0106 | 106.55 |  |  | 104.933 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.4 | 2021/05/17 18:20:40 | 95
05/17/21 | 27 | 20210048.017 | 533 | 0.85 |  |  | 5690 | -44.012 |  | 6 | 88.8838 | 106.681 |  |  | 105.061 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 6.9 | 2021/05/17 18:20:40 | 95
05/17/21 | 28 | 20210048.018 | 535 | 0.811 | 6851 | 0.114 |  |  |  | 1 | 21.6187 | 126.183 |  |  |  | 125.276 | 0 |  | 0.72417 |  | 0.0036786 |  | 0.36651 | BGD 28: 45.1; BGD 29: 34.7; BGD 30: 989.0 | 2021/05/17 18:30:25 | 0
05/17/21 | 28 | 20210048.018 | 535 | 0.811 | 6845 | 0 |  |  |  | 2 | 21.621 | 126.197 |  |  |  | 125.29 | 1 |  | 0.72409 |  | 0.0036782 |  | 0.36647 | BGD 28: 44.0; BGD 29: 33.9; BGD 30: 1109.5 | 2021/05/17 18:30:25 | 0
05/17/21 | 28 | 20210048.018 | 535 | 0.811 | 4090 | 9.468 |  |  | Nitrogen | 3 | 16.9096 | 98.813 |  |  |  | 98.096 | 0 |  | 0.73094 |  | 0.003713 |  | 0.36993 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1083.7 | 2021/05/17 18:30:25 | 0
05/17/21 | 28 | 20210048.018 | 535 | 0.811 |  |  | 859 | -32.941 | Carbon | 4 | 54.9599 | 23.354 |  |  | 22.994 |  | 0 | 1.1483 |  | 0.010812 |  | 1.0696 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.2 | 2021/05/17 18:30:25 | 95
05/17/21 | 28 | 20210048.018 | 535 | 0.811 |  |  | 5703 | -43.8 |  | 5 | 93.3388 | 106.503 |  |  | 104.886 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.5 | 2021/05/17 18:30:25 | 95
05/17/21 | 28 | 20210048.018 | 535 | 0.811 |  |  | 5688 | -44.021 |  | 6 | 93.1356 | 106.703 |  |  | 105.083 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 6.9 | 2021/05/17 18:30:25 | 95
05/17/21 | 29 | 20210048.019 | 537 | 0.828 | 6838 | 0.12 |  |  |  | 1 | 21.1334 | 125.938 |  |  |  | 125.032 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36652 | BGD 28: 45.1; BGD 29: 34.7; BGD 30: 992.5 | 2021/05/17 18:40:10 | 0
05/17/21 | 29 | 20210048.019 | 537 | 0.828 | 6844 | 0 |  |  |  | 2 | 21.1529 | 126.054 |  |  |  | 125.147 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 44.0; BGD 29: 33.8; BGD 30: 1112.8 | 2021/05/17 18:40:10 | 0
05/17/21 | 29 | 20210048.019 | 537 | 0.828 | 4124 | 10.733 |  |  | Nitrogen | 3 | 16.698 | 99.618 |  |  |  | 98.894 | 0 |  | 0.73188 |  | 0.0037177 |  | 0.37039 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1087.5 | 2021/05/17 18:40:10 | 0
05/17/21 | 29 | 20210048.019 | 537 | 0.828 |  |  | 871 | -32.353 | Carbon | 4 | 54.4484 | 23.679 |  |  | 23.315 |  | 0 | 1.149 |  | 0.010818 |  | 1.0703 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/17 18:40:10 | 95
05/17/21 | 29 | 20210048.019 | 537 | 0.828 |  |  | 5693 | -43.8 |  | 5 | 91.5895 | 106.335 |  |  | 104.72 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.6 | 2021/05/17 18:40:10 | 95
05/17/21 | 29 | 20210048.019 | 537 | 0.828 |  |  | 5686 | -44.011 |  | 6 | 91.241 | 106.685 |  |  | 105.066 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 6.9 | 2021/05/17 18:40:10 | 95
05/17/21 | 30 | 20210048.020 | 541 | 0.844 | 6838 | 0.094 |  |  |  | 1 | 20.7519 | 126.053 |  |  |  | 125.147 | 0 |  | 0.72418 |  | 0.0036785 |  | 0.36651 | BGD 28: 45.0; BGD 29: 34.6; BGD 30: 996.1 | 2021/05/17 18:49:55 | 0
05/17/21 | 30 | 20210048.020 | 541 | 0.844 | 6833 | 0 |  |  |  | 2 | 20.7438 | 126.005 |  |  |  | 125.099 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 44.0; BGD 29: 33.8; BGD 30: 1117.3 | 2021/05/17 18:49:55 | 0
05/17/21 | 30 | 20210048.020 | 541 | 0.844 | 3940 | 11.865 |  |  | Nitrogen | 3 | 15.6445 | 95.155 |  |  |  | 94.463 | 0 |  | 0.7327 |  | 0.0037218 |  | 0.3708 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1090.6 | 2021/05/17 18:49:55 | 0
05/17/21 | 30 | 20210048.020 | 541 | 0.844 |  |  | 908 | -29.54 | Carbon | 4 | 55.3679 | 24.744 |  |  | 24.362 |  | 0 | 1.1521 |  | 0.01085 |  | 1.0733 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/17 18:49:55 | 95
05/17/21 | 30 | 20210048.020 | 541 | 0.844 |  |  | 5686 | -43.8 |  | 5 | 89.842 | 106.346 |  |  | 104.732 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.7 | 2021/05/17 18:49:55 | 95
05/17/21 | 30 | 20210048.020 | 541 | 0.844 |  |  | 5681 | -44.008 |  | 6 | 89.5525 | 106.643 |  |  | 105.024 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.0 | 2021/05/17 18:49:55 | 95
05/17/21 | 31 | 20210048.021 | 543 | 0.826 | 6832 | 0.114 |  |  |  | 1 | 21.1896 | 125.968 |  |  |  | 125.062 | 0 |  | 0.72418 |  | 0.0036786 |  | 0.36651 | BGD 28: 45.1; BGD 29: 34.7; BGD 30: 999.4 | 2021/05/17 18:59:40 | 0
05/17/21 | 31 | 20210048.021 | 543 | 0.826 | 6837 | 0 |  |  |  | 2 | 21.1981 | 126.018 |  |  |  | 125.112 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 44.0; BGD 29: 33.8; BGD 30: 1120.4 | 2021/05/17 18:59:40 | 0
05/17/21 | 31 | 20210048.021 | 543 | 0.826 | 4097 | 9.986 |  |  | Nitrogen | 3 | 16.6393 | 99.031 |  |  |  | 98.312 | 0 |  | 0.73133 |  | 0.0037149 |  | 0.37012 | BGD 28: 45.0; BGD 29: 34.4; BGD 30: 1094.4 | 2021/05/17 18:59:40 | 0
05/17/21 | 31 | 20210048.021 | 543 | 0.826 |  |  | 866 | -31.269 | Carbon | 4 | 54.3885 | 23.578 |  |  | 23.215 |  | 0 | 1.1502 |  | 0.010831 |  | 1.0715 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.4 | 2021/05/17 18:59:40 | 95
05/17/21 | 31 | 20210048.021 | 543 | 0.826 |  |  | 5681 | -43.8 |  | 5 | 91.8133 | 106.333 |  |  | 104.718 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.7 | 2021/05/17 18:59:40 | 95
05/17/21 | 31 | 20210048.021 | 543 | 0.826 |  |  | 5687 | -44.025 |  | 6 | 91.55 | 106.597 |  |  | 104.979 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.0 | 2021/05/17 18:59:40 | 95
05/17/21 | 32 | 20210048.022 | 545 | 0.831 | 6828 | 0.12 |  |  |  | 1 | 21.0128 | 125.674 |  |  |  | 124.77 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36652 | BGD 28: 45.0; BGD 29: 34.6; BGD 30: 1002.2 | 2021/05/17 19:09:25 | 0
05/17/21 | 32 | 20210048.022 | 545 | 0.831 | 6833 | 0 |  |  |  | 2 | 21.0547 | 125.924 |  |  |  | 125.018 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.9; BGD 29: 33.7; BGD 30: 1122.8 | 2021/05/17 19:09:25 | 0
05/17/21 | 32 | 20210048.022 | 545 | 0.831 | 3794 | 8.84 |  |  | Nitrogen | 3 | 15.3032 | 91.659 |  |  |  | 90.994 | 0 |  | 0.7305 |  | 0.0037107 |  | 0.3697 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1096.0 | 2021/05/17 19:09:25 | 0
05/17/21 | 32 | 20210048.022 | 545 | 0.831 |  |  | 905 | -33.152 | Carbon | 4 | 56.06 | 24.65 |  |  | 24.27 |  | 0 | 1.1481 |  | 0.01081 |  | 1.0694 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.4 | 2021/05/17 19:09:25 | 95
05/17/21 | 32 | 20210048.022 | 545 | 0.831 |  |  | 5685 | -43.8 |  | 5 | 91.2469 | 106.347 |  |  | 104.732 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 5.8 | 2021/05/17 19:09:25 | 95
05/17/21 | 32 | 20210048.022 | 545 | 0.831 |  |  | 5684 | -44.025 |  | 6 | 91.0254 | 106.571 |  |  | 104.953 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.2 | 2021/05/17 19:09:25 | 95
05/17/21 | 33 | 20210048.022 | Replicate C8 | 0.829 | 6819 | 0.129 |  |  |  | 1 | 21.079 | 125.766 |  |  |  | 124.862 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 45.1; BGD 29: 34.6; BGD 30: 1003.7 | 2021/05/17 19:19:10 | 0
05/17/21 | 33 | 20210048.022 | Replicate C8 | 0.829 | 6834 | 0 |  |  |  | 2 | 21.1079 | 125.938 |  |  |  | 125.032 | 1 |  | 0.72409 |  | 0.0036782 |  | 0.36647 | BGD 28: 44.0; BGD 29: 33.8; BGD 30: 1125.1 | 2021/05/17 19:19:10 | 0
05/17/21 | 33 | 20210048.022 | Replicate C8 | 0.829 | 3797 | 8.84 |  |  | Nitrogen | 3 | 15.2976 | 91.406 |  |  |  | 90.743 | 0 |  | 0.73049 |  | 0.0037107 |  | 0.3697 | BGD 28: 44.9; BGD 29: 34.3; BGD 30: 1096.5 | 2021/05/17 19:19:10 | 0
05/17/21 | 33 | 20210048.022 | Replicate C8 | 0.829 |  |  | 915 | -33.132 | Carbon | 4 | 56.7438 | 24.947 |  |  | 24.563 |  | 0 | 1.1481 |  | 0.01081 |  | 1.0694 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/17 19:19:10 | 95
05/17/21 | 33 | 20210048.022 | Replicate C8 | 0.829 |  |  | 5684 | -43.8 |  | 5 | 91.5203 | 106.293 |  |  | 104.679 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 5.7 | 2021/05/17 19:19:10 | 95
05/17/21 | 33 | 20210048.022 | Replicate C8 | 0.829 |  |  | 5679 | -44.008 |  | 6 | 91.2412 | 106.575 |  |  | 104.957 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.2 | 2021/05/17 19:19:10 | 95
05/17/21 | 34 | 36-UWSIF-UT Glut 1-20210048.16 | 36-UWSIF-UT Glut 1- | 0.842 | 6830 | 0.099 |  |  |  | 1 | 20.7625 | 125.82 |  |  |  | 124.916 | 0 |  | 0.72418 |  | 0.0036786 |  | 0.36651 | BGD 28: 45.0; BGD 29: 34.6; BGD 30: 1005.9 | 2021/05/17 19:28:55 | 0
05/17/21 | 34 | 36-UWSIF-UT Glut 1-20210048.16 | 36-UWSIF-UT Glut 1- | 0.842 | 6836 | 0 |  |  |  | 2 | 20.7806 | 125.929 |  |  |  | 125.024 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.9; BGD 29: 33.8; BGD 30: 1127.4 | 2021/05/17 19:28:55 | 0
05/17/21 | 34 | 36-UWSIF-UT Glut 1-20210048.16 | 36-UWSIF-UT Glut 1- | 0.842 | 2708 | -2.848 |  |  | Nitrogen | 3 | 10.8027 | 65.632 |  |  |  | 65.162 | 0 |  | 0.72205 |  | 0.0036677 |  | 0.36543 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1097.9 | 2021/05/17 19:28:55 | 0
05/17/21 | 34 | 36-UWSIF-UT Glut 1-20210048.16 | 36-UWSIF-UT Glut 1- | 0.842 |  |  | 750 | -29.822 | Carbon | 4 | 46.9564 | 20.24 |  |  | 19.927 |  | 0 | 1.1518 |  | 0.010847 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.4 | 2021/05/17 19:28:55 | 95
05/17/21 | 34 | 36-UWSIF-UT Glut 1-20210048.16 | 36-UWSIF-UT Glut 1- | 0.842 |  |  | 5684 | -43.8 |  | 5 | 90.1355 | 106.264 |  |  | 104.65 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.7 | 2021/05/17 19:28:55 | 95
05/17/21 | 34 | 36-UWSIF-UT Glut 1-20210048.16 | 36-UWSIF-UT Glut 1- | 0.842 |  |  | 5681 | -44.02 |  | 6 | 89.8547 | 106.552 |  |  | 104.934 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 7.1 | 2021/05/17 19:28:55 | 95
05/17/21 | 35 | 36-UWSIF-UT Glut 1-20210048.17 | 36-UWSIF-UT Glut 1- | 0.766 | 6823 | 0.125 |  |  |  | 1 | 22.8252 | 125.835 |  |  |  | 124.93 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1003.1 | 2021/05/17 19:38:40 | 0
05/17/21 | 35 | 36-UWSIF-UT Glut 1-20210048.17 | 36-UWSIF-UT Glut 1- | 0.766 | 6825 | 0 |  |  |  | 2 | 22.8296 | 125.859 |  |  |  | 124.954 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1125.3 | 2021/05/17 19:38:40 | 0
05/17/21 | 35 | 36-UWSIF-UT Glut 1-20210048.17 | 36-UWSIF-UT Glut 1- | 0.766 | 2437 | -2.777 |  |  | Nitrogen | 3 | 10.7264 | 59.303 |  |  |  | 58.878 | 0 |  | 0.72209 |  | 0.003668 |  | 0.36546 | BGD 28: 44.7; BGD 29: 34.1; BGD 30: 1094.5 | 2021/05/17 19:38:40 | 0
05/17/21 | 35 | 36-UWSIF-UT Glut 1-20210048.17 | 36-UWSIF-UT Glut 1- | 0.766 |  |  | 677 | -29.878 | Carbon | 4 | 47.1344 | 18.21 |  |  | 17.929 |  | 0 | 1.1518 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.5 | 2021/05/17 19:38:40 | 95
05/17/21 | 35 | 36-UWSIF-UT Glut 1-20210048.17 | 36-UWSIF-UT Glut 1- | 0.766 |  |  | 5683 | -43.8 |  | 5 | 98.9593 | 106.375 |  |  | 104.76 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 5.6 | 2021/05/17 19:38:40 | 95
05/17/21 | 35 | 36-UWSIF-UT Glut 1-20210048.17 | 36-UWSIF-UT Glut 1- | 0.766 |  |  | 5684 | -44.018 |  | 6 | 98.6774 | 106.638 |  |  | 105.019 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 7.1 | 2021/05/17 19:38:40 | 95
05/17/21 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 39-UWSIF-UW Glut 2- | 0.788 | 6828 | 0.111 |  |  |  | 1 | 22.1868 | 125.829 |  |  |  | 124.924 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.7; BGD 29: 34.3; BGD 30: 1000.8 | 2021/05/17 19:48:25 | 0
05/17/21 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 39-UWSIF-UW Glut 2- | 0.788 | 6835 | 0 |  |  |  | 2 | 22.2265 | 126.052 |  |  |  | 125.146 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.5; BGD 30: 1123.3 | 2021/05/17 19:48:25 | 0
05/17/21 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 39-UWSIF-UW Glut 2- | 0.788 | 2722 | 29.711 |  |  | Nitrogen | 3 | 11.6617 | 66.306 |  |  |  | 65.815 | 0 |  | 0.74562 |  | 0.0037875 |  | 0.37732 | BGD 28: 44.6; BGD 29: 34.1; BGD 30: 1093.5 | 2021/05/17 19:48:25 | 0
05/17/21 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 39-UWSIF-UW Glut 2- | 0.788 |  |  | 754 | 22.276 | Carbon | 4 | 50.6087 | 20.446 |  |  | 20.119 |  | 0 | 1.2093 |  | 0.011429 |  | 1.13 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.5 | 2021/05/17 19:48:25 | 95
05/17/21 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 39-UWSIF-UW Glut 2- | 0.788 |  |  | 5689 | -43.8 |  | 5 | 96.176 | 106.395 |  |  | 104.78 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.7 | 2021/05/17 19:48:25 | 95
05/17/21 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 39-UWSIF-UW Glut 2- | 0.788 |  |  | 5692 | -44.034 |  | 6 | 95.8471 | 106.709 |  |  | 105.089 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 7.2 | 2021/05/17 19:48:25 | 95
05/17/21 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 39-UWSIF-UW Glut 2- | 0.782 | 6824 | 0.12 |  |  |  | 1 | 22.3451 | 125.762 |  |  |  | 124.858 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36652 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 1002.0 | 2021/05/17 19:58:10 | 0
05/17/21 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 39-UWSIF-UW Glut 2- | 0.782 | 6836 | 0 |  |  |  | 2 | 22.3777 | 125.944 |  |  |  | 125.039 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1125.2 | 2021/05/17 19:58:10 | 0
05/17/21 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 39-UWSIF-UW Glut 2- | 0.782 | 2713 | 29.779 |  |  | Nitrogen | 3 | 11.6998 | 66.016 |  |  |  | 65.528 | 0 |  | 0.74567 |  | 0.0037877 |  | 0.37734 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1095.8 | 2021/05/17 19:58:10 | 0
05/17/21 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 39-UWSIF-UW Glut 2- | 0.782 |  |  | 752 | 22.428 | Carbon | 4 | 50.7279 | 20.319 |  |  | 19.994 |  | 0 | 1.2095 |  | 0.011431 |  | 1.1302 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.5 | 2021/05/17 19:58:10 | 95
05/17/21 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 39-UWSIF-UW Glut 2- | 0.782 |  |  | 5678 | -43.8 |  | 5 | 97.0498 | 106.265 |  |  | 104.652 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.7 | 2021/05/17 19:58:10 | 95
05/17/21 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 39-UWSIF-UW Glut 2- | 0.782 |  |  | 5688 | -44.007 |  | 6 | 96.7121 | 106.587 |  |  | 104.969 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 7.2 | 2021/05/17 19:58:10 | 95
05/17/21 | 38 | 01-UWSIF-Liver-20210048.13 | 01-UWSIF-Liver- | 0.776 | 6828 | 0.099 |  |  |  | 1 | 22.5219 | 125.784 |  |  |  | 124.88 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 1002.7 | 2021/05/17 20:07:55 | 0
05/17/21 | 38 | 01-UWSIF-Liver-20210048.13 | 01-UWSIF-Liver- | 0.776 | 6834 | 0 |  |  |  | 2 | 22.5571 | 125.98 |  |  |  | 125.074 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1125.5 | 2021/05/17 20:07:55 | 0
05/17/21 | 38 | 01-UWSIF-Liver-20210048.13 | 01-UWSIF-Liver- | 0.776 | 3440 | 8.569 |  |  | Nitrogen | 3 | 14.8232 | 82.938 |  |  |  | 82.337 | 0 |  | 0.73032 |  | 0.0037097 |  | 0.3696 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1096.4 | 2021/05/17 20:07:55 | 0
05/17/21 | 38 | 01-UWSIF-Liver-20210048.13 | 01-UWSIF-Liver- | 0.776 |  |  | 852 | -19.454 | Carbon | 4 | 56.9295 | 23.103 |  |  | 22.744 |  | 0 | 1.1632 |  | 0.010963 |  | 1.0844 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.6 | 2021/05/17 20:07:55 | 95
05/17/21 | 38 | 01-UWSIF-Liver-20210048.13 | 01-UWSIF-Liver- | 0.776 |  |  | 5682 | -43.8 |  | 5 | 97.8004 | 106.265 |  |  | 104.652 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.9 | 2021/05/17 20:07:55 | 95
05/17/21 | 38 | 01-UWSIF-Liver-20210048.13 | 01-UWSIF-Liver- | 0.776 |  |  | 5677 | -44.016 |  | 6 | 97.4029 | 106.64 |  |  | 105.021 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.3 | 2021/05/17 20:07:55 | 95
05/17/21 | 39 | 01-UWSIF-Liver-20210048.14 | 01-UWSIF-Liver- | 0.752 | 6831 | 0.11 |  |  |  | 1 | 23.2375 | 125.767 |  |  |  | 124.863 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1007.0 | 2021/05/17 20:17:40 | 0
05/17/21 | 39 | 01-UWSIF-Liver-20210048.14 | 01-UWSIF-Liver- | 0.752 | 6825 | 0 |  |  |  | 2 | 23.2149 | 125.645 |  |  |  | 124.742 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1129.6 | 2021/05/17 20:17:40 | 0
05/17/21 | 39 | 01-UWSIF-Liver-20210048.14 | 01-UWSIF-Liver- | 0.752 | 3351 | 8.527 |  |  | Nitrogen | 3 | 14.8994 | 80.794 |  |  |  | 80.208 | 0 |  | 0.73028 |  | 0.0037096 |  | 0.36958 | BGD 28: 44.8; BGD 29: 34.2; BGD 30: 1099.2 | 2021/05/17 20:17:40 | 0
05/17/21 | 39 | 01-UWSIF-Liver-20210048.14 | 01-UWSIF-Liver- | 0.752 |  |  | 832 | -19.433 | Carbon | 4 | 57.4793 | 22.503 |  |  | 22.153 |  | 0 | 1.1633 |  | 0.010963 |  | 1.0844 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.7 | 2021/05/17 20:17:40 | 95
05/17/21 | 39 | 01-UWSIF-Liver-20210048.14 | 01-UWSIF-Liver- | 0.752 |  |  | 5687 | -43.8 |  | 5 | 100.7694 | 106.405 |  |  | 104.789 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 5.9 | 2021/05/17 20:17:40 | 95
05/17/21 | 39 | 01-UWSIF-Liver-20210048.14 | 01-UWSIF-Liver- | 0.752 |  |  | 5694 | -44.032 |  | 6 | 100.4235 | 106.72 |  |  | 105.1 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.4 | 2021/05/17 20:17:40 | 95
05/17/21 | 40 | 20210048.023 | 547 | 0.832 | 6827 | 0.115 |  |  |  | 1 | 21.0058 | 125.783 |  |  |  | 124.879 | 0 |  | 0.7242 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1009.2 | 2021/05/17 20:27:25 | 0
05/17/21 | 40 | 20210048.023 | 547 | 0.832 | 6832 | 0 |  |  |  | 2 | 21.0017 | 125.759 |  |  |  | 124.855 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1131.8 | 2021/05/17 20:27:25 | 0
05/17/21 | 40 | 20210048.023 | 547 | 0.832 | 4182 | 11.99 |  |  | Nitrogen | 3 | 16.856 | 101.04 |  |  |  | 100.305 | 0 |  | 0.73279 |  | 0.0037223 |  | 0.37085 | BGD 28: 44.8; BGD 29: 34.2; BGD 30: 1104.0 | 2021/05/17 20:27:25 | 0
05/17/21 | 40 | 20210048.023 | 547 | 0.832 |  |  | 886 | -30.766 | Carbon | 4 | 55.0529 | 24.143 |  |  | 23.77 |  | 0 | 1.1508 |  | 0.010836 |  | 1.072 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/17 20:27:25 | 95
05/17/21 | 40 | 20210048.023 | 547 | 0.832 |  |  | 5694 | -43.8 |  | 5 | 91.034 | 106.452 |  |  | 104.835 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 5.9 | 2021/05/17 20:27:25 | 95
05/17/21 | 40 | 20210048.023 | 547 | 0.832 |  |  | 5683 | -44.026 |  | 6 | 90.7935 | 106.694 |  |  | 105.074 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.4 | 2021/05/17 20:27:25 | 95
05/17/21 | 41 | 20210048.024 | 549 | 0.821 | 6832 | 0.104 |  |  |  | 1 | 21.2827 | 125.756 |  |  |  | 124.852 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 45.0; BGD 29: 34.5; BGD 30: 1013.5 | 2021/05/17 20:37:10 | 0
05/17/21 | 41 | 20210048.024 | 549 | 0.821 | 6823 | 0 |  |  |  | 2 | 21.2956 | 125.832 |  |  |  | 124.927 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.9; BGD 29: 33.7; BGD 30: 1136.6 | 2021/05/17 20:37:10 | 0
05/17/21 | 41 | 20210048.024 | 549 | 0.821 | 4132 | 12.51 |  |  | Nitrogen | 3 | 16.9413 | 100.213 |  |  |  | 99.483 | 0 |  | 0.73317 |  | 0.0037242 |  | 0.37104 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1108.6 | 2021/05/17 20:37:10 | 0
05/17/21 | 41 | 20210048.024 | 549 | 0.821 |  |  | 822 | -31.256 | Carbon | 4 | 52.2968 | 22.322 |  |  | 21.978 |  | 0 | 1.1502 |  | 0.010831 |  | 1.0715 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/17 20:37:10 | 95
05/17/21 | 41 | 20210048.024 | 549 | 0.821 |  |  | 5679 | -43.8 |  | 5 | 92.4408 | 106.264 |  |  | 104.651 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.0 | 2021/05/17 20:37:10 | 95
05/17/21 | 41 | 20210048.024 | 549 | 0.821 |  |  | 5686 | -44.027 |  | 6 | 92.011 | 106.693 |  |  | 105.073 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.4 | 2021/05/17 20:37:10 | 95
05/17/21 | 42 | 20210048.025 | 551 | 0.83 | 6824 | 0.131 |  |  |  | 1 | 21.0412 | 125.692 |  |  |  | 124.789 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1016.9 | 2021/05/17 20:46:55 | 0
05/17/21 | 42 | 20210048.025 | 551 | 0.83 | 6830 | 0 |  |  |  | 2 | 21.0669 | 125.845 |  |  |  | 124.94 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1140.1 | 2021/05/17 20:46:55 | 0
05/17/21 | 42 | 20210048.025 | 551 | 0.83 | 3591 | 13.091 |  |  | Nitrogen | 3 | 14.455 | 86.493 |  |  |  | 85.863 | 0 |  | 0.73358 |  | 0.0037264 |  | 0.37125 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1110.7 | 2021/05/17 20:46:55 | 0
05/17/21 | 42 | 20210048.025 | 551 | 0.83 |  |  | 945 | -34.044 | Carbon | 4 | 58.1679 | 25.765 |  |  | 25.368 |  | 0 | 1.1471 |  | 0.0108 |  | 1.0684 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/17 20:46:55 | 95
05/17/21 | 42 | 20210048.025 | 551 | 0.83 |  |  | 5689 | -43.8 |  | 5 | 91.2827 | 106.422 |  |  | 104.806 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.1 | 2021/05/17 20:46:55 | 95
05/17/21 | 42 | 20210048.025 | 551 | 0.83 |  |  | 5686 | -44.009 |  | 6 | 91.1298 | 106.576 |  |  | 104.958 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.6 | 2021/05/17 20:46:55 | 95
05/17/21 | 43 | 20210048.026 | 553 | 0.826 | 6826 | 0.103 |  |  |  | 1 | 21.1533 | 125.753 |  |  |  | 124.848 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1018.1 | 2021/05/17 20:56:40 | 0
05/17/21 | 43 | 20210048.026 | 553 | 0.826 | 6828 | 0 |  |  |  | 2 | 21.1726 | 125.867 |  |  |  | 124.962 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.9; BGD 29: 33.7; BGD 30: 1141.5 | 2021/05/17 20:56:40 | 0
05/17/21 | 43 | 20210048.026 | 553 | 0.826 | 4149 | 12.83 |  |  | Nitrogen | 3 | 16.8171 | 100.084 |  |  |  | 99.356 | 0 |  | 0.7334 |  | 0.0037254 |  | 0.37116 | BGD 28: 44.8; BGD 29: 34.3; BGD 30: 1113.3 | 2021/05/17 20:56:40 | 0
05/17/21 | 43 | 20210048.026 | 553 | 0.826 |  |  | 873 | -32.839 | Carbon | 4 | 54.6551 | 23.719 |  |  | 23.354 |  | 0 | 1.1485 |  | 0.010813 |  | 1.0697 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/17 20:56:40 | 95
05/17/21 | 43 | 20210048.026 | 553 | 0.826 |  |  | 5697 | -43.8 |  | 5 | 91.7116 | 106.435 |  |  | 104.819 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.1 | 2021/05/17 20:56:40 | 95
05/17/21 | 43 | 20210048.026 | 553 | 0.826 |  |  | 5685 | -44.01 |  | 6 | 91.4573 | 106.69 |  |  | 105.07 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.6 | 2021/05/17 20:56:40 | 95
05/17/21 | 44 | 20210048.027 | 555 | 0.827 | 6830 | 0.118 |  |  |  | 1 | 21.1254 | 125.739 |  |  |  | 124.835 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1021.2 | 2021/05/17 21:06:25 | 0
05/17/21 | 44 | 20210048.027 | 555 | 0.827 | 6823 | 0 |  |  |  | 2 | 21.1322 | 125.779 |  |  |  | 124.875 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1144.5 | 2021/05/17 21:06:25 | 0
05/17/21 | 44 | 20210048.027 | 555 | 0.827 | 3924 | 12.524 |  |  | Nitrogen | 3 | 15.8865 | 94.682 |  |  |  | 93.993 | 0 |  | 0.73318 |  | 0.0037243 |  | 0.37105 | BGD 28: 44.8; BGD 29: 34.2; BGD 30: 1116.1 | 2021/05/17 21:06:25 | 0
05/17/21 | 44 | 20210048.027 | 555 | 0.827 |  |  | 910 | -33.221 | Carbon | 4 | 56.56 | 24.773 |  |  | 24.391 |  | 0 | 1.1481 |  | 0.010809 |  | 1.0693 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/17 21:06:25 | 95
05/17/21 | 44 | 20210048.027 | 555 | 0.827 |  |  | 5698 | -43.8 |  | 5 | 91.6308 | 106.405 |  |  | 104.789 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.2 | 2021/05/17 21:06:25 | 95
05/17/21 | 44 | 20210048.027 | 555 | 0.827 |  |  | 5696 | -44.019 |  | 6 | 91.3582 | 106.678 |  |  | 105.059 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.7 | 2021/05/17 21:06:25 | 95
05/17/21 | 45 | 20210048.028 | 557 | 0.82 | 6826 | 0.115 |  |  |  | 1 | 21.3231 | 125.841 |  |  |  | 124.936 | 0 |  | 0.7242 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1023.1 | 2021/05/17 21:16:10 | 0
05/17/21 | 45 | 20210048.028 | 557 | 0.82 | 6827 | 0 |  |  |  | 2 | 21.3054 | 125.737 |  |  |  | 124.833 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.9; BGD 29: 33.7; BGD 30: 1146.8 | 2021/05/17 21:16:10 | 0
05/17/21 | 45 | 20210048.028 | 557 | 0.82 | 3922 | 12.513 |  |  | Nitrogen | 3 | 16.0126 | 94.627 |  |  |  | 93.938 | 0 |  | 0.73317 |  | 0.0037242 |  | 0.37104 | BGD 28: 44.8; BGD 29: 34.2; BGD 30: 1117.8 | 2021/05/17 21:16:10 | 0
05/17/21 | 45 | 20210048.028 | 557 | 0.82 |  |  | 902 | -33.315 | Carbon | 4 | 56.627 | 24.551 |  |  | 24.173 |  | 0 | 1.1479 |  | 0.010808 |  | 1.0692 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/17 21:16:10 | 95
05/17/21 | 45 | 20210048.028 | 557 | 0.82 |  |  | 5696 | -43.8 |  | 5 | 92.4425 | 106.375 |  |  | 104.76 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.3 | 2021/05/17 21:16:10 | 95
05/17/21 | 45 | 20210048.028 | 557 | 0.82 |  |  | 5690 | -44.013 |  | 6 | 92.188 | 106.629 |  |  | 105.01 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.7 | 2021/05/17 21:16:10 | 95
05/17/21 | 46 | 20210048.029 | 559 | 0.84 | 6826 | 0.111 |  |  |  | 1 | 20.7879 | 125.676 |  |  |  | 124.772 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1025.1 | 2021/05/17 21:25:55 | 0
05/17/21 | 46 | 20210048.029 | 559 | 0.84 | 6828 | 0 |  |  |  | 2 | 20.809 | 125.802 |  |  |  | 124.898 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1148.7 | 2021/05/17 21:25:55 | 0
05/17/21 | 46 | 20210048.029 | 559 | 0.84 | 4188 | 12.549 |  |  | Nitrogen | 3 | 16.7241 | 101.213 |  |  |  | 100.476 | 0 |  | 0.7332 |  | 0.0037244 |  | 0.37105 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1120.8 | 2021/05/17 21:25:55 | 0
05/17/21 | 46 | 20210048.029 | 559 | 0.84 |  |  | 895 | -31.919 | Carbon | 4 | 55.0203 | 24.41 |  |  | 24.033 |  | 0 | 1.1495 |  | 0.010823 |  | 1.0707 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/17 21:25:55 | 95
05/17/21 | 46 | 20210048.029 | 559 | 0.84 |  |  | 5690 | -43.8 |  | 5 | 90.1299 | 106.49 |  |  | 104.873 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 6.3 | 2021/05/17 21:25:55 | 95
05/17/21 | 46 | 20210048.029 | 559 | 0.84 |  |  | 5694 | -44.003 |  | 6 | 89.9529 | 106.67 |  |  | 105.05 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.7 | 2021/05/17 21:25:55 | 95
05/17/21 | 47 | 20210048.030 | 561 | 0.843 | 6820 | 0.127 |  |  |  | 1 | 20.7024 | 125.606 |  |  |  | 124.703 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1027.2 | 2021/05/17 21:35:40 | 0
05/17/21 | 47 | 20210048.030 | 561 | 0.843 | 6829 | 0 |  |  |  | 2 | 20.7344 | 125.799 |  |  |  | 124.895 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1150.3 | 2021/05/17 21:35:40 | 0
05/17/21 | 47 | 20210048.030 | 561 | 0.843 | 4307 | 9.411 |  |  | Nitrogen | 3 | 17.1366 | 104.067 |  |  |  | 103.312 | 0 |  | 0.73092 |  | 0.0037128 |  | 0.36991 | BGD 28: 44.8; BGD 29: 34.2; BGD 30: 1121.7 | 2021/05/17 21:35:40 | 0
05/17/21 | 47 | 20210048.030 | 561 | 0.843 |  |  | 872 | -32.66 | Carbon | 4 | 53.5508 | 23.718 |  |  | 23.352 |  | 0 | 1.1487 |  | 0.010815 |  | 1.0699 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/17 21:35:40 | 95
05/17/21 | 47 | 20210048.030 | 561 | 0.843 |  |  | 5695 | -43.8 |  | 5 | 89.8404 | 106.458 |  |  | 104.841 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.3 | 2021/05/17 21:35:40 | 95
05/17/21 | 47 | 20210048.030 | 561 | 0.843 |  |  | 5688 | -43.999 |  | 6 | 89.5431 | 106.761 |  |  | 105.14 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.7 | 2021/05/17 21:35:40 | 95
05/17/21 | 48 | 20210048.031 | 563 | 0.844 | 6829 | 0.123 |  |  |  | 1 | 20.7036 | 125.762 |  |  |  | 124.858 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1030.0 | 2021/05/17 21:45:26 | 0
05/17/21 | 48 | 20210048.031 | 563 | 0.844 | 6824 | 0 |  |  |  | 2 | 20.7091 | 125.795 |  |  |  | 124.89 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1153.4 | 2021/05/17 21:45:26 | 0
05/17/21 | 48 | 20210048.031 | 563 | 0.844 | 4286 | 12.97 |  |  | Nitrogen | 3 | 17.0501 | 103.666 |  |  |  | 102.912 | 0 |  | 0.73349 |  | 0.0037259 |  | 0.37121 | BGD 28: 44.7; BGD 29: 34.1; BGD 30: 1124.4 | 2021/05/17 21:45:26 | 0
05/17/21 | 48 | 20210048.031 | 563 | 0.844 |  |  | 884 | -31.669 | Carbon | 4 | 54.1792 | 24.093 |  |  | 23.722 |  | 0 | 1.1498 |  | 0.010826 |  | 1.071 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/17 21:45:26 | 95
05/17/21 | 48 | 20210048.031 | 563 | 0.844 |  |  | 5691 | -43.8 |  | 5 | 89.7781 | 106.412 |  |  | 104.796 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.4 | 2021/05/17 21:45:26 | 95
05/17/21 | 48 | 20210048.031 | 563 | 0.844 |  |  | 5686 | -44.009 |  | 6 | 89.5484 | 106.647 |  |  | 105.028 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.8 | 2021/05/17 21:45:26 | 95
05/17/21 | 49 | 20210048.032 | 565 | 0.847 | 6825 | 0.139 |  |  |  | 1 | 20.617 | 125.681 |  |  |  | 124.778 | 0 |  | 0.72421 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1032.0 | 2021/05/17 21:55:11 | 0
05/17/21 | 49 | 20210048.032 | 565 | 0.847 | 6819 | 0 |  |  |  | 2 | 20.628 | 125.747 |  |  |  | 124.843 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1155.5 | 2021/05/17 21:55:11 | 0
05/17/21 | 49 | 20210048.032 | 565 | 0.847 | 4325 | 9.562 |  |  | Nitrogen | 3 | 17.1519 | 104.652 |  |  |  | 103.892 | 0 |  | 0.73103 |  | 0.0037134 |  | 0.36996 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1126.5 | 2021/05/17 21:55:11 | 0
05/17/21 | 49 | 20210048.032 | 565 | 0.847 |  |  | 895 | -32.757 | Carbon | 4 | 54.6746 | 24.47 |  |  | 24.093 |  | 0 | 1.1486 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/17 21:55:11 | 95
05/17/21 | 49 | 20210048.032 | 565 | 0.847 |  |  | 5691 | -43.8 |  | 5 | 89.3855 | 106.489 |  |  | 104.872 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.4 | 2021/05/17 21:55:11 | 95
05/17/21 | 49 | 20210048.032 | 565 | 0.847 |  |  | 5687 | -44.001 |  | 6 | 89.2059 | 106.673 |  |  | 105.054 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.8 | 2021/05/17 21:55:11 | 95
05/17/21 | 50 | 20210048.033 | 567 | 0.82 | 6812 | 0.111 |  |  |  | 1 | 21.2817 | 125.598 |  |  |  | 124.695 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1033.7 | 2021/05/17 22:04:56 | 0
05/17/21 | 50 | 20210048.033 | 567 | 0.82 | 6825 | 0 |  |  |  | 2 | 21.3016 | 125.715 |  |  |  | 124.811 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1157.0 | 2021/05/17 22:04:56 | 0
05/17/21 | 50 | 20210048.033 | 567 | 0.82 | 3845 | 13.209 |  |  | Nitrogen | 3 | 15.7056 | 92.819 |  |  |  | 92.143 | 0 |  | 0.73368 |  | 0.0037268 |  | 0.37129 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1127.9 | 2021/05/17 22:04:56 | 0
05/17/21 | 50 | 20210048.033 | 567 | 0.82 |  |  | 907 | -32.658 | Carbon | 4 | 56.9874 | 24.743 |  |  | 24.362 |  | 0 | 1.1487 |  | 0.010815 |  | 1.0699 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/17 22:04:56 | 95
05/17/21 | 50 | 20210048.033 | 567 | 0.82 |  |  | 5684 | -43.8 |  | 5 | 92.4617 | 106.356 |  |  | 104.741 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/17 22:04:56 | 95
05/17/21 | 50 | 20210048.033 | 567 | 0.82 |  |  | 5684 | -44.016 |  | 6 | 92.1009 | 106.715 |  |  | 105.095 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.8 | 2021/05/17 22:04:56 | 95
05/17/21 | 51 | 20210048.034 | 569 | 0.837 | 6813 | 0.115 |  |  |  | 1 | 20.8356 | 125.515 |  |  |  | 124.613 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1033.3 | 2021/05/17 22:14:41 | 0
05/17/21 | 51 | 20210048.034 | 569 | 0.837 | 6824 | 0 |  |  |  | 2 | 20.8721 | 125.734 |  |  |  | 124.83 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1156.8 | 2021/05/17 22:14:41 | 0
05/17/21 | 51 | 20210048.034 | 569 | 0.837 | 4231 | 9.589 |  |  | Nitrogen | 3 | 16.977 | 102.371 |  |  |  | 101.628 | 0 |  | 0.73105 |  | 0.0037135 |  | 0.36997 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1127.5 | 2021/05/17 22:14:41 | 0
05/17/21 | 51 | 20210048.034 | 569 | 0.837 |  |  | 851 | -32.431 | Carbon | 4 | 52.8508 | 23.14 |  |  | 22.783 |  | 0 | 1.1489 |  | 0.010818 |  | 1.0702 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.4 | 2021/05/17 22:14:41 | 95
05/17/21 | 51 | 20210048.034 | 569 | 0.837 |  |  | 5688 | -43.8 |  | 5 | 90.549 | 106.392 |  |  | 104.776 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/17 22:14:41 | 95
05/17/21 | 51 | 20210048.034 | 569 | 0.837 |  |  | 5691 | -44.002 |  | 6 | 90.1868 | 106.759 |  |  | 105.138 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.8 | 2021/05/17 22:14:41 | 95
05/17/21 | 52 | 20210048.035 | 571 | 0.848 | 6826 | 0.127 |  |  |  | 1 | 20.5683 | 125.533 |  |  |  | 124.63 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1035.1 | 2021/05/17 22:24:26 | 0
05/17/21 | 52 | 20210048.035 | 571 | 0.848 | 6823 | 0 |  |  |  | 2 | 20.6031 | 125.744 |  |  |  | 124.84 | 1 |  | 0.72409 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1158.8 | 2021/05/17 22:24:26 | 0
05/17/21 | 52 | 20210048.035 | 571 | 0.848 | 4289 | 9.372 |  |  | Nitrogen | 3 | 16.9233 | 103.384 |  |  |  | 102.634 | 0 |  | 0.73088 |  | 0.0037127 |  | 0.36989 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1130.8 | 2021/05/17 22:24:26 | 0
05/17/21 | 52 | 20210048.035 | 571 | 0.848 |  |  | 871 | -32.726 | Carbon | 4 | 53.1355 | 23.664 |  |  | 23.299 |  | 0 | 1.1485 |  | 0.010814 |  | 1.0699 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/17 22:24:26 | 95
05/17/21 | 52 | 20210048.035 | 571 | 0.848 |  |  | 5687 | -43.8 |  | 5 | 89.473 | 106.29 |  |  | 104.676 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.4 | 2021/05/17 22:24:26 | 95
05/17/21 | 52 | 20210048.035 | 571 | 0.848 |  |  | 5686 | -44.029 |  | 6 | 89.063 | 106.712 |  |  | 105.092 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.9 | 2021/05/17 22:24:26 | 95
05/17/21 | 53 | 20210048.036 | 573 | 0.802 | 6812 | 0.133 |  |  |  | 1 | 21.746 | 125.521 |  |  |  | 124.619 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.4; BGD 30: 1036.3 | 2021/05/17 22:34:11 | 0
05/17/21 | 53 | 20210048.036 | 573 | 0.802 | 6815 | 0 |  |  |  | 2 | 21.7545 | 125.57 |  |  |  | 124.667 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.7; BGD 30: 1160.2 | 2021/05/17 22:34:11 | 0
05/17/21 | 53 | 20210048.036 | 573 | 0.802 | 3945 | 12.325 |  |  | Nitrogen | 3 | 16.4925 | 95.32 |  |  |  | 94.626 | 0 |  | 0.73303 |  | 0.0037235 |  | 0.37097 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1132.3 | 2021/05/17 22:34:11 | 0
05/17/21 | 53 | 20210048.036 | 573 | 0.802 |  |  | 855 | -31.863 | Carbon | 4 | 55.3508 | 23.238 |  |  | 22.88 |  | 0 | 1.1495 |  | 0.010824 |  | 1.0708 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/17 22:34:11 | 95
05/17/21 | 53 | 20210048.036 | 573 | 0.802 |  |  | 5686 | -43.8 |  | 5 | 94.5352 | 106.358 |  |  | 104.743 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/17 22:34:11 | 95
05/17/21 | 53 | 20210048.036 | 573 | 0.802 |  |  | 5682 | -44.015 |  | 6 | 94.3155 | 106.572 |  |  | 104.954 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.9 | 2021/05/17 22:34:11 | 95
05/17/21 | 54 | 20210048.037 | 575 | 0.832 | 6814 | 0.103 |  |  |  | 1 | 20.9552 | 125.482 |  |  |  | 124.579 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 1037.0 | 2021/05/17 22:43:56 | 0
05/17/21 | 54 | 20210048.037 | 575 | 0.832 | 6814 | 0 |  |  |  | 2 | 20.9751 | 125.6 |  |  |  | 124.697 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1160.8 | 2021/05/17 22:43:56 | 0
05/17/21 | 54 | 20210048.037 | 575 | 0.832 | 3966 | 10.273 |  |  | Nitrogen | 3 | 15.9712 | 95.758 |  |  |  | 95.063 | 0 |  | 0.73155 |  | 0.003716 |  | 0.37022 | BGD 28: 44.7; BGD 29: 34.1; BGD 30: 1131.8 | 2021/05/17 22:43:56 | 0
05/17/21 | 54 | 20210048.037 | 575 | 0.832 |  |  | 885 | -32.934 | Carbon | 4 | 55.0057 | 24.117 |  |  | 23.746 |  | 0 | 1.1483 |  | 0.010812 |  | 1.0696 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.3 | 2021/05/17 22:43:56 | 95
05/17/21 | 54 | 20210048.037 | 575 | 0.832 |  |  | 5682 | -43.8 |  | 5 | 91.2053 | 106.278 |  |  | 104.664 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/17 22:43:56 | 95
05/17/21 | 54 | 20210048.037 | 575 | 0.832 |  |  | 5684 | -44.016 |  | 6 | 90.851 | 106.636 |  |  | 105.017 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.0 | 2021/05/17 22:43:56 | 95
05/17/21 | 55 | 20210048.038 | 577 | 0.826 | 6811 | 0.114 |  |  |  | 1 | 21.1036 | 125.459 |  |  |  | 124.557 | 0 |  | 0.72418 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1037.9 | 2021/05/17 22:53:41 | 0
05/17/21 | 55 | 20210048.038 | 577 | 0.826 | 6811 | 0 |  |  |  | 2 | 21.099 | 125.432 |  |  |  | 124.53 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1161.8 | 2021/05/17 22:53:41 | 0
05/17/21 | 55 | 20210048.038 | 577 | 0.826 | 4164 | 10.112 |  |  | Nitrogen | 3 | 16.875 | 100.427 |  |  |  | 99.698 | 0 |  | 0.73142 |  | 0.0037154 |  | 0.37016 | BGD 28: 44.8; BGD 29: 34.2; BGD 30: 1133.5 | 2021/05/17 22:53:41 | 0
05/17/21 | 55 | 20210048.038 | 577 | 0.826 |  |  | 859 | -32.359 | Carbon | 4 | 53.9097 | 23.326 |  |  | 22.967 |  | 0 | 1.149 |  | 0.010818 |  | 1.0703 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.4 | 2021/05/17 22:53:41 | 95
05/17/21 | 55 | 20210048.038 | 577 | 0.826 |  |  | 5687 | -43.8 |  | 5 | 91.8763 | 106.269 |  |  | 104.656 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/17 22:53:41 | 95
05/17/21 | 55 | 20210048.038 | 577 | 0.826 |  |  | 5675 | -44.015 |  | 6 | 91.634 | 106.513 |  |  | 104.896 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.0 | 2021/05/17 22:53:41 | 95
05/17/21 | 56 | 20210048.039 | 579 | 0.836 | 6815 | 0.101 |  |  |  | 1 | 20.8666 | 125.551 |  |  |  | 124.649 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 1038.9 | 2021/05/17 23:03:26 | 0
05/17/21 | 56 | 20210048.039 | 579 | 0.836 | 6819 | 0 |  |  |  | 2 | 20.8789 | 125.625 |  |  |  | 124.722 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1163.8 | 2021/05/17 23:03:26 | 0
05/17/21 | 56 | 20210048.039 | 579 | 0.836 | 3060 | 12.481 |  |  | Nitrogen | 3 | 12.2008 | 73.574 |  |  |  | 73.038 | 0 |  | 0.73315 |  | 0.0037241 |  | 0.37103 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1132.8 | 2021/05/17 23:03:26 | 0
05/17/21 | 56 | 20210048.039 | 579 | 0.836 |  |  | 1009 | -32.499 | Carbon | 4 | 61.0556 | 27.637 |  |  | 27.211 |  | 0 | 1.1488 |  | 0.010817 |  | 1.0701 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.2 | 2021/05/17 23:03:26 | 95
05/17/21 | 56 | 20210048.039 | 579 | 0.836 |  |  | 5683 | -43.8 |  | 5 | 90.8207 | 106.225 |  |  | 104.612 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 6.6 | 2021/05/17 23:03:26 | 95
05/17/21 | 56 | 20210048.039 | 579 | 0.836 |  |  | 5678 | -44.02 |  | 6 | 90.5638 | 106.487 |  |  | 104.87 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.1 | 2021/05/17 23:03:26 | 95
05/17/21 | 57 | 20210048.040 | 581 | 0.836 | 6805 | 0.132 |  |  |  | 1 | 20.8537 | 125.474 |  |  |  | 124.572 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1037.6 | 2021/05/17 23:13:11 | 0
05/17/21 | 57 | 20210048.040 | 581 | 0.836 | 6813 | 0 |  |  |  | 2 | 20.86 | 125.512 |  |  |  | 124.609 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.9; BGD 29: 33.7; BGD 30: 1162.2 | 2021/05/17 23:13:11 | 0
05/17/21 | 57 | 20210048.040 | 581 | 0.836 | 3989 | 11.349 |  |  | Nitrogen | 3 | 15.9859 | 96.305 |  |  |  | 95.604 | 0 |  | 0.73232 |  | 0.0037199 |  | 0.37062 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1132.8 | 2021/05/17 23:13:11 | 0
05/17/21 | 57 | 20210048.040 | 581 | 0.836 |  |  | 889 | -32.768 | Carbon | 4 | 54.8959 | 24.2 |  |  | 23.827 |  | 0 | 1.1485 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.5 | 2021/05/17 23:13:11 | 95
05/17/21 | 57 | 20210048.040 | 581 | 0.836 |  |  | 5684 | -43.8 |  | 5 | 90.7393 | 106.308 |  |  | 104.694 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.6 | 2021/05/17 23:13:11 | 95
05/17/21 | 57 | 20210048.040 | 581 | 0.836 |  |  | 5674 | -44.014 |  | 6 | 90.6053 | 106.445 |  |  | 104.829 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.0 | 2021/05/17 23:13:11 | 95
05/17/21 | 58 | 20210048.041 | 583 | 0.836 | 6808 | 0.098 |  |  |  | 1 | 20.8439 | 125.416 |  |  |  | 124.514 | 0 |  | 0.72418 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1040.1 | 2021/05/17 23:22:56 | 0
05/17/21 | 58 | 20210048.041 | 583 | 0.836 | 6806 | 0 |  |  |  | 2 | 20.8489 | 125.446 |  |  |  | 124.544 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1164.9 | 2021/05/17 23:22:56 | 0
05/17/21 | 58 | 20210048.041 | 583 | 0.836 | 3378 | 12.947 |  |  | Nitrogen | 3 | 13.4566 | 81.12 |  |  |  | 80.53 | 0 |  | 0.73349 |  | 0.0037258 |  | 0.3712 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1134.1 | 2021/05/17 23:22:56 | 0
05/17/21 | 58 | 20210048.041 | 583 | 0.836 |  |  | 988 | -34.764 | Carbon | 4 | 59.9065 | 26.978 |  |  | 26.563 |  | 0 | 1.1463 |  | 0.010792 |  | 1.0676 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.4 | 2021/05/17 23:22:56 | 95
05/17/21 | 58 | 20210048.041 | 583 | 0.836 |  |  | 5679 | -43.8 |  | 5 | 90.9038 | 106.139 |  |  | 104.528 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 6.7 | 2021/05/17 23:22:56 | 95
05/17/21 | 58 | 20210048.041 | 583 | 0.836 |  |  | 5672 | -44.022 |  | 6 | 90.7035 | 106.345 |  |  | 104.73 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.1 | 2021/05/17 23:22:56 | 95
05/17/21 | 59 | 20210048.042 | 585 | 0.839 | 6802 | 0.114 |  |  |  | 1 | 20.7507 | 125.303 |  |  |  | 124.402 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.9; BGD 29: 34.5; BGD 30: 1039.6 | 2021/05/17 23:32:41 | 0
05/17/21 | 59 | 20210048.042 | 585 | 0.839 | 6794 | 0 |  |  |  | 2 | 20.7516 | 125.308 |  |  |  | 124.408 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.9; BGD 29: 33.7; BGD 30: 1164.4 | 2021/05/17 23:32:41 | 0
05/17/21 | 59 | 20210048.042 | 585 | 0.839 | 4148 | 10.235 |  |  | Nitrogen | 3 | 16.5881 | 100.274 |  |  |  | 99.546 | 0 |  | 0.73152 |  | 0.0037158 |  | 0.37021 | BGD 28: 44.8; BGD 29: 34.2; BGD 30: 1135.5 | 2021/05/17 23:32:41 | 0
05/17/21 | 59 | 20210048.042 | 585 | 0.839 |  |  | 871 | -32.395 | Carbon | 4 | 53.882 | 23.759 |  |  | 23.392 |  | 0 | 1.149 |  | 0.010818 |  | 1.0702 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.5 | 2021/05/17 23:32:41 | 95
05/17/21 | 59 | 20210048.042 | 585 | 0.839 |  |  | 5673 | -43.8 |  | 5 | 90.5931 | 106.125 |  |  | 104.513 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 6.6 | 2021/05/17 23:32:41 | 95
05/17/21 | 59 | 20210048.042 | 585 | 0.839 |  |  | 5676 | -44.017 |  | 6 | 90.2387 | 106.488 |  |  | 104.872 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.1 | 2021/05/17 23:32:41 | 95
05/17/21 | 60 | 20210048.042 | Replicate E11 | 0.832 | 6806 | 0.113 |  |  |  | 1 | 20.9148 | 125.241 |  |  |  | 124.34 | 0 |  | 0.72418 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 1042.4 | 2021/05/17 23:42:27 | 0
05/17/21 | 60 | 20210048.042 | Replicate E11 | 0.832 | 6798 | 0 |  |  |  | 2 | 20.9179 | 125.259 |  |  |  | 124.359 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.8; BGD 29: 33.6; BGD 30: 1166.5 | 2021/05/17 23:42:27 | 0
05/17/21 | 60 | 20210048.042 | Replicate E11 | 0.832 | 4116 | 10.245 |  |  | Nitrogen | 3 | 16.5474 | 99.199 |  |  |  | 98.478 | 0 |  | 0.73152 |  | 0.0037159 |  | 0.37021 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1137.1 | 2021/05/17 23:42:27 | 0
05/17/21 | 60 | 20210048.042 | Replicate E11 | 0.832 |  |  | 866 | -32.312 | Carbon | 4 | 53.8621 | 23.507 |  |  | 23.145 |  | 0 | 1.149 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.6 | 2021/05/17 23:42:27 | 95
05/17/21 | 60 | 20210048.042 | Replicate E11 | 0.832 |  |  | 5676 | -43.8 |  | 5 | 91.3864 | 106.093 |  |  | 104.482 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/17 23:42:27 | 95
05/17/21 | 60 | 20210048.042 | Replicate E11 | 0.832 |  |  | 5669 | -43.998 |  | 6 | 91.178 | 106.306 |  |  | 104.692 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.1 | 2021/05/17 23:42:27 | 95
05/17/21 | 61 | 20210048.043 | 587 | 0.805 | 6798 | 0.121 |  |  |  | 1 | 21.6239 | 125.285 |  |  |  | 124.384 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36652 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 1044.0 | 2021/05/17 23:52:12 | 0
05/17/21 | 61 | 20210048.043 | 587 | 0.805 | 6805 | 0 |  |  |  | 2 | 21.6103 | 125.207 |  |  |  | 124.307 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1168.6 | 2021/05/17 23:52:12 | 0
05/17/21 | 61 | 20210048.043 | 587 | 0.805 | 2431 | 12.831 |  |  | Nitrogen | 3 | 10.0394 | 58.333 |  |  |  | 57.908 | 0 |  | 0.73339 |  | 0.0037254 |  | 0.37116 | BGD 28: 44.6; BGD 29: 34.1; BGD 30: 1134.9 | 2021/05/17 23:52:12 | 0
05/17/21 | 61 | 20210048.043 | 587 | 0.805 |  |  | 994 | -32.999 | Carbon | 4 | 62.5214 | 27.147 |  |  | 26.729 |  | 0 | 1.1483 |  | 0.010811 |  | 1.0696 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/17 23:52:12 | 95
05/17/21 | 61 | 20210048.043 | 587 | 0.805 |  |  | 5671 | -43.8 |  | 5 | 94.478 | 106.066 |  |  | 104.456 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 6.7 | 2021/05/17 23:52:12 | 95
05/17/21 | 61 | 20210048.043 | 587 | 0.805 |  |  | 5665 | -44.024 |  | 6 | 94.3307 | 106.212 |  |  | 104.6 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.1 | 2021/05/17 23:52:12 | 95
05/18/21 | 62 | 20210048.044 | 589 | 0.818 | 6796 | 0.131 |  |  |  | 1 | 21.2654 | 125.198 |  |  |  | 124.298 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 1038.9 | 2021/05/18 00:01:57 | 0
05/18/21 | 62 | 20210048.044 | 589 | 0.818 | 6790 | 0 |  |  |  | 2 | 21.2599 | 125.166 |  |  |  | 124.266 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.6; BGD 30: 1164.1 | 2021/05/18 00:01:57 | 0
05/18/21 | 62 | 20210048.044 | 589 | 0.818 | 3803 | 12.376 |  |  | Nitrogen | 3 | 15.4959 | 91.362 |  |  |  | 90.697 | 0 |  | 0.73306 |  | 0.0037237 |  | 0.37099 | BGD 28: 44.6; BGD 29: 34.1; BGD 30: 1133.5 | 2021/05/18 00:01:57 | 0
05/18/21 | 62 | 20210048.044 | 589 | 0.818 |  |  | 901 | -33.411 | Carbon | 4 | 56.638 | 24.483 |  |  | 24.106 |  | 0 | 1.1478 |  | 0.010807 |  | 1.0691 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.5 | 2021/05/18 00:01:57 | 95
05/18/21 | 62 | 20210048.044 | 589 | 0.818 |  |  | 5674 | -43.8 |  | 5 | 93.0682 | 105.974 |  |  | 104.365 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 6.7 | 2021/05/18 00:01:57 | 95
05/18/21 | 62 | 20210048.044 | 589 | 0.818 |  |  | 5668 | -44.001 |  | 6 | 92.7781 | 106.266 |  |  | 104.653 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.1 | 2021/05/18 00:01:57 | 95
05/18/21 | 63 | 20210048.045 | 591 | 0.812 | 6792 | 0.125 |  |  |  | 1 | 21.4006 | 125.07 |  |  |  | 124.171 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.8; BGD 29: 34.4; BGD 30: 1041.8 | 2021/05/18 00:11:42 | 0
05/18/21 | 63 | 20210048.045 | 591 | 0.812 | 6788 | 0 |  |  |  | 2 | 21.4032 | 125.086 |  |  |  | 124.186 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.5; BGD 30: 1166.9 | 2021/05/18 00:11:42 | 0
05/18/21 | 63 | 20210048.045 | 591 | 0.812 | 3604 | 13.26 |  |  | Nitrogen | 3 | 14.8024 | 86.65 |  |  |  | 86.019 | 0 |  | 0.73371 |  | 0.003727 |  | 0.37131 | BGD 28: 44.7; BGD 29: 34.2; BGD 30: 1137.0 | 2021/05/18 00:11:42 | 0
05/18/21 | 63 | 20210048.045 | 591 | 0.812 |  |  | 877 | -33.542 | Carbon | 4 | 55.8201 | 23.835 |  |  | 23.468 |  | 0 | 1.1477 |  | 0.010805 |  | 1.069 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.6 | 2021/05/18 00:11:42 | 95
05/18/21 | 63 | 20210048.045 | 591 | 0.812 |  |  | 5662 | -43.8 |  | 5 | 93.7664 | 105.963 |  |  | 104.354 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 6.8 | 2021/05/18 00:11:42 | 95
05/18/21 | 63 | 20210048.045 | 591 | 0.812 |  |  | 5669 | -44.028 |  | 6 | 93.4779 | 106.252 |  |  | 104.639 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.2 | 2021/05/18 00:11:42 | 95
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.18 | 36-UWSIF-UT Glut 1- | 0.783 | 6783 | 0.11 |  |  |  | 1 | 22.1796 | 124.994 |  |  |  | 124.095 | 0 |  | 0.7242 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.7; BGD 29: 34.3; BGD 30: 1042.3 | 2021/05/18 00:21:27 | 0
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.18 | 36-UWSIF-UT Glut 1- | 0.783 | 6786 | 0 |  |  |  | 2 | 22.2086 | 125.156 |  |  |  | 124.257 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.7; BGD 29: 33.5; BGD 30: 1167.2 | 2021/05/18 00:21:27 | 0
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.18 | 36-UWSIF-UT Glut 1- | 0.783 | 2484 | -2.853 |  |  | Nitrogen | 3 | 10.6466 | 60.166 |  |  |  | 59.734 | 0 |  | 0.72205 |  | 0.0036677 |  | 0.36543 | BGD 28: 44.6; BGD 29: 34.1; BGD 30: 1135.5 | 2021/05/18 00:21:27 | 0
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.18 | 36-UWSIF-UT Glut 1- | 0.783 |  |  | 694 | -29.77 | Carbon | 4 | 47.027 | 18.628 |  |  | 18.341 |  | 0 | 1.1519 |  | 0.010847 |  | 1.0731 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.5 | 2021/05/18 00:21:27 | 95
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.18 | 36-UWSIF-UT Glut 1- | 0.783 |  |  | 5661 | -43.8 |  | 5 | 97.279 | 105.925 |  |  | 104.316 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.6 | 2021/05/18 00:21:27 | 95
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.18 | 36-UWSIF-UT Glut 1- | 0.783 |  |  | 5667 | -44.007 |  | 6 | 96.9871 | 106.207 |  |  | 104.594 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.0 | 2021/05/18 00:21:27 | 95
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.19 | 36-UWSIF-UT Glut 1- | 0.794 | 6791 | 0.094 |  |  |  | 1 | 21.8841 | 125.061 |  |  |  | 124.162 | 0 |  | 0.72419 |  | 0.0036785 |  | 0.36651 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 1038.5 | 2021/05/18 00:31:12 | 0
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.19 | 36-UWSIF-UT Glut 1- | 0.794 | 6789 | 0 |  |  |  | 2 | 21.894 | 125.117 |  |  |  | 124.218 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.3; BGD 30: 1163.6 | 2021/05/18 00:31:12 | 0
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.19 | 36-UWSIF-UT Glut 1- | 0.794 | 2519 | -2.841 |  |  | Nitrogen | 3 | 10.6443 | 60.995 |  |  |  | 60.558 | 0 |  | 0.72206 |  | 0.0036677 |  | 0.36544 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1133.1 | 2021/05/18 00:31:12 | 0
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.19 | 36-UWSIF-UT Glut 1- | 0.794 |  |  | 703 | -29.887 | Carbon | 4 | 46.9074 | 18.876 |  |  | 18.585 |  | 0 | 1.1518 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.5 | 2021/05/18 00:31:12 | 95
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.19 | 36-UWSIF-UT Glut 1- | 0.794 |  |  | 5669 | -43.8 |  | 5 | 95.8639 | 105.991 |  |  | 104.382 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/18 00:31:12 | 95
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.19 | 36-UWSIF-UT Glut 1- | 0.794 |  |  | 5663 | -44.029 |  | 6 | 95.6586 | 106.192 |  |  | 104.58 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.1 | 2021/05/18 00:31:12 | 95
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 39-UWSIF-UW Glut 2- | 0.735 | 6788 | 0.101 |  |  |  | 1 | 23.6263 | 124.985 |  |  |  | 124.086 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.5; BGD 29: 34.1; BGD 30: 1036.0 | 2021/05/18 00:40:57 | 0
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 39-UWSIF-UW Glut 2- | 0.735 | 6782 | 0 |  |  |  | 2 | 23.6374 | 125.043 |  |  |  | 124.144 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.3; BGD 30: 1161.5 | 2021/05/18 00:40:57 | 0
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 39-UWSIF-UW Glut 2- | 0.735 | 2520 | 29.733 |  |  | Nitrogen | 3 | 11.5411 | 61.22 |  |  |  | 60.766 | 0 |  | 0.74565 |  | 0.0037876 |  | 0.37733 | BGD 28: 44.3; BGD 29: 33.8; BGD 30: 1129.5 | 2021/05/18 00:40:57 | 0
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 39-UWSIF-UW Glut 2- | 0.735 |  |  | 703 | 22.538 | Carbon | 4 | 50.8276 | 18.943 |  |  | 18.64 |  | 0 | 1.2096 |  | 0.011432 |  | 1.1303 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.4 | 2021/05/18 00:40:57 | 95
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 39-UWSIF-UW Glut 2- | 0.735 |  |  | 5665 | -43.8 |  | 5 | 103.6673 | 105.892 |  |  | 104.285 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.6 | 2021/05/18 00:40:57 | 95
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 39-UWSIF-UW Glut 2- | 0.735 |  |  | 5665 | -44.027 |  | 6 | 103.2765 | 106.247 |  |  | 104.634 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.1 | 2021/05/18 00:40:57 | 95
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 39-UWSIF-UW Glut 2- | 0.722 | 6783 | 0.137 |  |  |  | 1 | 24.031 | 124.877 |  |  |  | 123.98 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.5; BGD 29: 34.1; BGD 30: 1034.7 | 2021/05/18 00:50:43 | 0
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 39-UWSIF-UW Glut 2- | 0.722 | 6780 | 0 |  |  |  | 2 | 24.0591 | 125.023 |  |  |  | 124.124 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.3; BGD 30: 1160.2 | 2021/05/18 00:50:43 | 0
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 39-UWSIF-UW Glut 2- | 0.722 | 2487 | 29.732 |  |  | Nitrogen | 3 | 11.6043 | 60.468 |  |  |  | 60.02 | 0 |  | 0.74563 |  | 0.0037876 |  | 0.37733 | BGD 28: 44.3; BGD 29: 33.8; BGD 30: 1128.5 | 2021/05/18 00:50:43 | 0
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 39-UWSIF-UW Glut 2- | 0.722 |  |  | 694 | 22.268 | Carbon | 4 | 51.1809 | 18.705 |  |  | 18.406 |  | 0 | 1.2093 |  | 0.011429 |  | 1.13 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.5 | 2021/05/18 00:50:43 | 95
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 39-UWSIF-UW Glut 2- | 0.722 |  |  | 5662 | -43.8 |  | 5 | 105.535 | 105.891 |  |  | 104.284 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.6 | 2021/05/18 00:50:43 | 95
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 39-UWSIF-UW Glut 2- | 0.722 |  |  | 5663 | -44.034 |  | 6 | 105.3013 | 106.1 |  |  | 104.489 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.0 | 2021/05/18 00:50:43 | 95
05/18/21 | 68 | 01-UWSIF-Liver-20210048.15 | 01-UWSIF-Liver- | 0.728 | 6773 | 0.121 |  |  |  | 1 | 23.8245 | 124.834 |  |  |  | 123.936 | 0 |  | 0.7242 |  | 0.0036786 |  | 0.36652 | BGD 28: 44.4; BGD 29: 34.1; BGD 30: 1034.1 | 2021/05/18 01:00:28 | 0
05/18/21 | 68 | 01-UWSIF-Liver-20210048.15 | 01-UWSIF-Liver- | 0.728 | 6781 | 0 |  |  |  | 2 | 23.8452 | 124.941 |  |  |  | 124.043 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.3; BGD 30: 1159.4 | 2021/05/18 01:00:28 | 0
05/18/21 | 68 | 01-UWSIF-Liver-20210048.15 | 01-UWSIF-Liver- | 0.728 | 3215 | 8.523 |  |  | Nitrogen | 3 | 14.7519 | 77.452 |  |  |  | 76.891 | 0 |  | 0.73029 |  | 0.0037095 |  | 0.36958 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1128.3 | 2021/05/18 01:00:28 | 0
05/18/21 | 68 | 01-UWSIF-Liver-20210048.15 | 01-UWSIF-Liver- | 0.728 |  |  | 803 | -19.338 | Carbon | 4 | 57.5315 | 21.669 |  |  | 21.332 |  | 0 | 1.1634 |  | 0.010964 |  | 1.0845 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.5 | 2021/05/18 01:00:28 | 95
05/18/21 | 68 | 01-UWSIF-Liver-20210048.15 | 01-UWSIF-Liver- | 0.728 |  |  | 5666 | -43.8 |  | 5 | 104.5869 | 105.962 |  |  | 104.353 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/18 01:00:28 | 95
05/18/21 | 68 | 01-UWSIF-Liver-20210048.15 | 01-UWSIF-Liver- | 0.728 |  |  | 5659 | -44.02 |  | 6 | 104.3431 | 106.181 |  |  | 104.569 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.2 | 2021/05/18 01:00:28 | 95
05/18/21 | 69 | 01-UWSIF-Liver-20210048.16 | 01-UWSIF-Liver- | 0.738 | 6779 | 0.109 |  |  |  | 1 | 23.5047 | 124.85 |  |  |  | 123.952 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 1037.6 | 2021/05/18 01:10:13 | 0
05/18/21 | 69 | 01-UWSIF-Liver-20210048.16 | 01-UWSIF-Liver- | 0.738 | 6775 | 0 |  |  |  | 2 | 23.5134 | 124.896 |  |  |  | 123.998 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.4; BGD 30: 1162.3 | 2021/05/18 01:10:13 | 0
05/18/21 | 69 | 01-UWSIF-Liver-20210048.16 | 01-UWSIF-Liver- | 0.738 | 3269 | 8.516 |  |  | Nitrogen | 3 | 14.7713 | 78.616 |  |  |  | 78.046 | 0 |  | 0.73028 |  | 0.0037095 |  | 0.36958 | BGD 28: 44.5; BGD 29: 33.9; BGD 30: 1130.4 | 2021/05/18 01:10:13 | 0
05/18/21 | 69 | 01-UWSIF-Liver-20210048.16 | 01-UWSIF-Liver- | 0.738 |  |  | 816 | -19.393 | Carbon | 4 | 57.5376 | 22.028 |  |  | 21.685 |  | 0 | 1.1633 |  | 0.010963 |  | 1.0844 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.7 | 2021/05/18 01:10:13 | 95
05/18/21 | 69 | 01-UWSIF-Liver-20210048.16 | 01-UWSIF-Liver- | 0.738 |  |  | 5659 | -43.8 |  | 5 | 103.2464 | 105.892 |  |  | 104.284 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/18 01:10:13 | 95
05/18/21 | 69 | 01-UWSIF-Liver-20210048.16 | 01-UWSIF-Liver- | 0.738 |  |  | 5658 | -44.018 |  | 6 | 103.0449 | 106.076 |  |  | 104.466 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.3 | 2021/05/18 01:10:13 | 95
05/18/21 | 70 | 20210048.046 | 593 | 0.822 | 6779 | 0.102 |  |  |  | 1 | 21.1037 | 124.855 |  |  |  | 123.958 | 0 |  | 0.7242 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 1038.5 | 2021/05/18 01:19:58 | 0
05/18/21 | 70 | 20210048.046 | 593 | 0.822 | 6777 | 0 |  |  |  | 2 | 21.1071 | 124.875 |  |  |  | 123.978 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.3; BGD 30: 1163.8 | 2021/05/18 01:19:58 | 0
05/18/21 | 70 | 20210048.046 | 593 | 0.822 | 4152 | 12.73 |  |  | Nitrogen | 3 | 16.9024 | 100.105 |  |  |  | 99.376 | 0 |  | 0.73334 |  | 0.003725 |  | 0.37112 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1135.8 | 2021/05/18 01:19:58 | 0
05/18/21 | 70 | 20210048.046 | 593 | 0.822 |  |  | 872 | -31.499 | Carbon | 4 | 54.8791 | 23.697 |  |  | 23.331 |  | 0 | 1.15 |  | 0.010828 |  | 1.0712 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.7 | 2021/05/18 01:19:58 | 95
05/18/21 | 70 | 20210048.046 | 593 | 0.822 |  |  | 5657 | -43.8 |  | 5 | 92.6963 | 105.891 |  |  | 104.283 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 6.9 | 2021/05/18 01:19:58 | 95
05/18/21 | 70 | 20210048.046 | 593 | 0.822 |  |  | 5659 | -44.028 |  | 6 | 92.486 | 106.105 |  |  | 104.494 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.3 | 2021/05/18 01:19:58 | 95
05/18/21 | 71 | 20210048.047 | 595 | 0.826 | 6778 | 0.095 |  |  |  | 1 | 21.0044 | 124.873 |  |  |  | 123.975 | 0 |  | 0.72418 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.7; BGD 29: 34.3; BGD 30: 1043.5 | 2021/05/18 01:29:43 | 0
05/18/21 | 71 | 20210048.047 | 595 | 0.826 | 6781 | 0 |  |  |  | 2 | 21.0141 | 124.93 |  |  |  | 124.032 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.4; BGD 30: 1168.7 | 2021/05/18 01:29:43 | 0
05/18/21 | 71 | 20210048.047 | 595 | 0.826 | 3983 | 9.002 |  |  | Nitrogen | 3 | 16.0732 | 95.675 |  |  |  | 94.981 | 0 |  | 0.73063 |  | 0.0037113 |  | 0.36976 | BGD 28: 44.5; BGD 29: 34.0; BGD 30: 1138.5 | 2021/05/18 01:29:43 | 0
05/18/21 | 71 | 20210048.047 | 595 | 0.826 |  |  | 894 | -33.458 | Carbon | 4 | 55.7236 | 24.287 |  |  | 23.913 |  | 0 | 1.1478 |  | 0.010806 |  | 1.0691 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.7 | 2021/05/18 01:29:43 | 95
05/18/21 | 71 | 20210048.047 | 595 | 0.826 |  |  | 5656 | -43.8 |  | 5 | 92.3044 | 105.833 |  |  | 104.226 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 6.9 | 2021/05/18 01:29:43 | 95
05/18/21 | 71 | 20210048.047 | 595 | 0.826 |  |  | 5657 | -44.001 |  | 6 | 92.0382 | 106.105 |  |  | 104.494 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.4 | 2021/05/18 01:29:43 | 95
05/18/21 | 72 | 20210048.048 | 597 | 0.849 | 6773 | 0.115 |  |  |  | 1 | 20.3931 | 124.615 |  |  |  | 123.719 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.7; BGD 29: 34.3; BGD 30: 1045.3 | 2021/05/18 01:39:28 | 0
05/18/21 | 72 | 20210048.048 | 597 | 0.849 | 6767 | 0 |  |  |  | 2 | 20.4017 | 124.668 |  |  |  | 123.771 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.4; BGD 30: 1170.0 | 2021/05/18 01:39:28 | 0
05/18/21 | 72 | 20210048.048 | 597 | 0.849 | 4276 | 9.143 |  |  | Nitrogen | 3 | 16.8633 | 103.14 |  |  |  | 102.392 | 0 |  | 0.73073 |  | 0.0037118 |  | 0.36981 | BGD 28: 44.5; BGD 29: 34.0; BGD 30: 1140.7 | 2021/05/18 01:39:28 | 0
05/18/21 | 72 | 20210048.048 | 597 | 0.849 |  |  | 886 | -31.891 | Carbon | 4 | 53.9118 | 24.121 |  |  | 23.749 |  | 0 | 1.1495 |  | 0.010824 |  | 1.0708 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/18 01:39:28 | 95
05/18/21 | 72 | 20210048.048 | 597 | 0.849 |  |  | 5661 | -43.8 |  | 5 | 89.914 | 105.716 |  |  | 104.111 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.0 | 2021/05/18 01:39:28 | 95
05/18/21 | 72 | 20210048.048 | 597 | 0.849 |  |  | 5651 | -44.011 |  | 6 | 89.5942 | 106.053 |  |  | 104.443 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.4 | 2021/05/18 01:39:28 | 95
05/18/21 | 73 | 20210048.049 | 599 | 0.807 | 6766 | 0.116 |  |  |  | 1 | 21.4666 | 124.685 |  |  |  | 123.789 | 0 |  | 0.7242 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.7; BGD 29: 34.3; BGD 30: 1048.4 | 2021/05/18 01:49:13 | 0
05/18/21 | 73 | 20210048.049 | 599 | 0.807 | 6776 | 0 |  |  |  | 2 | 21.4795 | 124.76 |  |  |  | 123.863 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.5; BGD 30: 1173.2 | 2021/05/18 01:49:13 | 0
05/18/21 | 73 | 20210048.049 | 599 | 0.807 | 3983 | 10.379 |  |  | Nitrogen | 3 | 16.4817 | 95.849 |  |  |  | 95.153 | 0 |  | 0.73163 |  | 0.0037164 |  | 0.37026 | BGD 28: 44.5; BGD 29: 34.0; BGD 30: 1142.9 | 2021/05/18 01:49:13 | 0
05/18/21 | 73 | 20210048.049 | 599 | 0.807 |  |  | 853 | -32.602 | Carbon | 4 | 54.792 | 23.128 |  |  | 22.771 |  | 0 | 1.1487 |  | 0.010816 |  | 1.07 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/18 01:49:13 | 95
05/18/21 | 73 | 20210048.049 | 599 | 0.807 |  |  | 5656 | -43.8 |  | 5 | 94.5359 | 105.774 |  |  | 104.168 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.0 | 2021/05/18 01:49:13 | 95
05/18/21 | 73 | 20210048.049 | 599 | 0.807 |  |  | 5649 | -44.007 |  | 6 | 94.3431 | 105.968 |  |  | 104.359 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.4 | 2021/05/18 01:49:13 | 95
05/18/21 | 74 | 20210048.050 | 601 | 0.833 | 6767 | 0.109 |  |  |  | 1 | 20.7726 | 124.543 |  |  |  | 123.647 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.6; BGD 29: 34.3; BGD 30: 1049.1 | 2021/05/18 01:58:59 | 0
05/18/21 | 74 | 20210048.050 | 601 | 0.833 | 6772 | 0 |  |  |  | 2 | 20.8061 | 124.743 |  |  |  | 123.846 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.4; BGD 30: 1174.6 | 2021/05/18 01:58:59 | 0
05/18/21 | 74 | 20210048.050 | 601 | 0.833 | 4005 | 12.882 |  |  | Nitrogen | 3 | 16.0639 | 96.427 |  |  |  | 95.725 | 0 |  | 0.73344 |  | 0.0037256 |  | 0.37117 | BGD 28: 44.5; BGD 29: 34.0; BGD 30: 1144.1 | 2021/05/18 01:58:59 | 0
05/18/21 | 74 | 20210048.050 | 601 | 0.833 |  |  | 890 | -32.978 | Carbon | 4 | 55.066 | 24.185 |  |  | 23.812 |  | 0 | 1.1483 |  | 0.010811 |  | 1.0696 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/18 01:58:59 | 95
05/18/21 | 74 | 20210048.050 | 601 | 0.833 |  |  | 5650 | -43.8 |  | 5 | 91.7479 | 105.605 |  |  | 104.001 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.0 | 2021/05/18 01:58:59 | 95
05/18/21 | 74 | 20210048.050 | 601 | 0.833 |  |  | 5648 | -44.007 |  | 6 | 91.4256 | 105.94 |  |  | 104.331 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.5 | 2021/05/18 01:58:59 | 95
05/18/21 | 75 | 20210048.051 | 603 | 0.822 | 6757 | 0.124 |  |  |  | 1 | 21.0348 | 124.45 |  |  |  | 123.555 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.7; BGD 29: 34.3; BGD 30: 1049.6 | 2021/05/18 02:08:44 | 0
05/18/21 | 75 | 20210048.051 | 603 | 0.822 | 6765 | 0 |  |  |  | 2 | 21.0636 | 124.619 |  |  |  | 123.723 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.4; BGD 30: 1174.2 | 2021/05/18 02:08:44 | 0
05/18/21 | 75 | 20210048.051 | 603 | 0.822 | 3921 | 10.957 |  |  | Nitrogen | 3 | 15.9353 | 94.4 |  |  |  | 93.714 | 0 |  | 0.73205 |  | 0.0037185 |  | 0.37047 | BGD 28: 44.5; BGD 29: 34.0; BGD 30: 1145.2 | 2021/05/18 02:08:44 | 0
05/18/21 | 75 | 20210048.051 | 603 | 0.822 |  |  | 880 | -32.799 | Carbon | 4 | 55.2327 | 23.883 |  |  | 23.515 |  | 0 | 1.1485 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/18 02:08:44 | 95
05/18/21 | 75 | 20210048.051 | 603 | 0.822 |  |  | 5650 | -43.8 |  | 5 | 92.9734 | 105.607 |  |  | 104.004 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.0 | 2021/05/18 02:08:44 | 95
05/18/21 | 75 | 20210048.051 | 603 | 0.822 |  |  | 5647 | -44 |  | 6 | 92.6435 | 105.945 |  |  | 104.337 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.4 | 2021/05/18 02:08:44 | 95
05/18/21 | 76 | 20210048.052 | 605 | 0.85 | 6763 | 0.11 |  |  |  | 1 | 20.3525 | 124.514 |  |  |  | 123.619 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 1050.1 | 2021/05/18 02:18:29 | 0
05/18/21 | 76 | 20210048.052 | 605 | 0.85 | 6765 | 0 |  |  |  | 2 | 20.3713 | 124.629 |  |  |  | 123.733 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.4; BGD 30: 1175.2 | 2021/05/18 02:18:29 | 0
05/18/21 | 76 | 20210048.052 | 605 | 0.85 | 4231 | 10.06 |  |  | Nitrogen | 3 | 16.5969 | 101.637 |  |  |  | 100.899 | 0 |  | 0.7314 |  | 0.0037152 |  | 0.37015 | BGD 28: 44.4; BGD 29: 34.0; BGD 30: 1146.2 | 2021/05/18 02:18:29 | 0
05/18/21 | 76 | 20210048.052 | 605 | 0.85 |  |  | 893 | -32.559 | Carbon | 4 | 54.118 | 24.269 |  |  | 23.895 |  | 0 | 1.1487 |  | 0.010816 |  | 1.07 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/18 02:18:29 | 95
05/18/21 | 76 | 20210048.052 | 605 | 0.85 |  |  | 5645 | -43.8 |  | 5 | 89.935 | 105.581 |  |  | 103.978 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.1 | 2021/05/18 02:18:29 | 95
05/18/21 | 76 | 20210048.052 | 605 | 0.85 |  |  | 5641 | -44.01 |  | 6 | 89.6263 | 105.909 |  |  | 104.301 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.5 | 2021/05/18 02:18:29 | 95
05/18/21 | 77 | 20210048.053 | 607 | 0.818 | 6763 | 0.103 |  |  |  | 1 | 21.1391 | 124.458 |  |  |  | 123.563 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 1051.5 | 2021/05/18 02:28:14 | 0
05/18/21 | 77 | 20210048.053 | 607 | 0.818 | 6755 | 0 |  |  |  | 2 | 21.1406 | 124.467 |  |  |  | 123.572 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.4; BGD 30: 1176.7 | 2021/05/18 02:28:14 | 0
05/18/21 | 77 | 20210048.053 | 607 | 0.818 | 4068 | 12.396 |  |  | Nitrogen | 3 | 16.6146 | 97.931 |  |  |  | 97.218 | 0 |  | 0.7331 |  | 0.0037238 |  | 0.371 | BGD 28: 44.5; BGD 29: 34.0; BGD 30: 1147.2 | 2021/05/18 02:28:14 | 0
05/18/21 | 77 | 20210048.053 | 607 | 0.818 |  |  | 848 | -31.584 | Carbon | 4 | 53.7966 | 22.994 |  |  | 22.639 |  | 0 | 1.1498 |  | 0.010827 |  | 1.0711 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/18 02:28:14 | 95
05/18/21 | 77 | 20210048.053 | 607 | 0.818 |  |  | 5644 | -43.8 |  | 5 | 93.4158 | 105.62 |  |  | 104.016 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.0 | 2021/05/18 02:28:14 | 95
05/18/21 | 77 | 20210048.053 | 607 | 0.818 |  |  | 5642 | -44.007 |  | 6 | 93.2822 | 105.756 |  |  | 104.151 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.5 | 2021/05/18 02:28:14 | 95
05/18/21 | 78 | 20210048.054 | 609 | 0.833 | 6754 | 0.126 |  |  |  | 1 | 20.7707 | 124.531 |  |  |  | 123.636 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.6; BGD 29: 34.1; BGD 30: 1052.0 | 2021/05/18 02:37:59 | 0
05/18/21 | 78 | 20210048.054 | 609 | 0.833 | 6750 | 0 |  |  |  | 2 | 20.7524 | 124.422 |  |  |  | 123.527 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.4; BGD 30: 1177.0 | 2021/05/18 02:37:59 | 0
05/18/21 | 78 | 20210048.054 | 609 | 0.833 | 3642 | 12.405 |  |  | Nitrogen | 3 | 14.5547 | 87.401 |  |  |  | 86.765 | 0 |  | 0.73309 |  | 0.0037238 |  | 0.371 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1145.3 | 2021/05/18 02:37:59 | 0
05/18/21 | 78 | 20210048.054 | 609 | 0.833 |  |  | 965 | -33.279 | Carbon | 4 | 59.0177 | 26.355 |  |  | 25.949 |  | 0 | 1.1479 |  | 0.010808 |  | 1.0693 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/18 02:37:59 | 95
05/18/21 | 78 | 20210048.054 | 609 | 0.833 |  |  | 5638 | -43.8 |  | 5 | 91.8981 | 105.448 |  |  | 103.847 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.1 | 2021/05/18 02:37:59 | 95
05/18/21 | 78 | 20210048.054 | 609 | 0.833 |  |  | 5643 | -43.995 |  | 6 | 91.5775 | 105.782 |  |  | 104.176 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.5 | 2021/05/18 02:37:59 | 95
05/18/21 | 79 | 20210048.055 | 611 | 0.826 | 6751 | 0.109 |  |  |  | 1 | 20.9128 | 124.331 |  |  |  | 123.437 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.6; BGD 29: 34.3; BGD 30: 1051.5 | 2021/05/18 02:47:44 | 0
05/18/21 | 79 | 20210048.055 | 611 | 0.826 | 6751 | 0 |  |  |  | 2 | 20.9144 | 124.34 |  |  |  | 123.446 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.6; BGD 29: 33.4; BGD 30: 1176.6 | 2021/05/18 02:47:44 | 0
05/18/21 | 79 | 20210048.055 | 611 | 0.826 | 3246 | 12.698 |  |  | Nitrogen | 3 | 13.0751 | 77.888 |  |  |  | 77.321 | 0 |  | 0.7333 |  | 0.0037249 |  | 0.37111 | BGD 28: 44.5; BGD 29: 34.0; BGD 30: 1144.9 | 2021/05/18 02:47:44 | 0
05/18/21 | 79 | 20210048.055 | 611 | 0.826 |  |  | 960 | -33.18 | Carbon | 4 | 59.1847 | 26.17 |  |  | 25.767 |  | 0 | 1.1481 |  | 0.010809 |  | 1.0694 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/18 02:47:44 | 95
05/18/21 | 79 | 20210048.055 | 611 | 0.826 |  |  | 5645 | -43.8 |  | 5 | 92.5519 | 105.578 |  |  | 103.974 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 7.1 | 2021/05/18 02:47:44 | 95
05/18/21 | 79 | 20210048.055 | 611 | 0.826 |  |  | 5639 | -44.028 |  | 6 | 92.3601 | 105.776 |  |  | 104.17 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.5 | 2021/05/18 02:47:44 | 95
05/18/21 | 80 | 20210048.056 | 613 | 0.801 | 6753 | 0.134 |  |  |  | 1 | 21.5732 | 124.375 |  |  |  | 123.481 | 0 |  | 0.72421 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 1050.0 | 2021/05/18 02:57:30 | 0
05/18/21 | 80 | 20210048.056 | 613 | 0.801 | 6753 | 0 |  |  |  | 2 | 21.583 | 124.431 |  |  |  | 123.536 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.4; BGD 30: 1175.2 | 2021/05/18 02:57:30 | 0
05/18/21 | 80 | 20210048.056 | 613 | 0.801 | 4030 | 9.36 |  |  | Nitrogen | 3 | 16.8121 | 97.039 |  |  |  | 96.335 | 0 |  | 0.73089 |  | 0.0037126 |  | 0.36989 | BGD 28: 44.3; BGD 29: 33.9; BGD 30: 1145.9 | 2021/05/18 02:57:30 | 0
05/18/21 | 80 | 20210048.056 | 613 | 0.801 |  |  | 838 | -32.673 | Carbon | 4 | 54.319 | 22.682 |  |  | 22.332 |  | 0 | 1.1487 |  | 0.010815 |  | 1.0699 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/18 02:57:30 | 95
05/18/21 | 80 | 20210048.056 | 613 | 0.801 |  |  | 5650 | -43.8 |  | 5 | 95.4095 | 105.609 |  |  | 104.005 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/18 02:57:30 | 95
05/18/21 | 80 | 20210048.056 | 613 | 0.801 |  |  | 5642 | -44.022 |  | 6 | 95.1526 | 105.865 |  |  | 104.258 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.5 | 2021/05/18 02:57:30 | 95
05/18/21 | 81 | 20210048.057 | 615 | 0.801 | 6749 | 0.097 |  |  |  | 1 | 21.5544 | 124.267 |  |  |  | 123.373 | 0 |  | 0.7242 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.5; BGD 29: 34.2; BGD 30: 1051.7 | 2021/05/18 03:07:15 | 0
05/18/21 | 81 | 20210048.057 | 615 | 0.801 | 6754 | 0 |  |  |  | 2 | 21.5814 | 124.422 |  |  |  | 123.527 | 1 |  | 0.72413 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.3; BGD 30: 1176.4 | 2021/05/18 03:07:15 | 0
05/18/21 | 81 | 20210048.057 | 615 | 0.801 | 3439 | 9.342 |  |  | Nitrogen | 3 | 14.2878 | 82.52 |  |  |  | 81.921 | 0 |  | 0.73089 |  | 0.0037126 |  | 0.36988 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1145.8 | 2021/05/18 03:07:15 | 0
05/18/21 | 81 | 20210048.057 | 615 | 0.801 |  |  | 890 | -34.063 | Carbon | 4 | 57.233 | 24.168 |  |  | 23.796 |  | 0 | 1.1471 |  | 0.010799 |  | 1.0684 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/18 03:07:15 | 95
05/18/21 | 81 | 20210048.057 | 615 | 0.801 |  |  | 5650 | -43.8 |  | 5 | 95.3649 | 105.653 |  |  | 104.049 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.1 | 2021/05/18 03:07:15 | 95
05/18/21 | 81 | 20210048.057 | 615 | 0.801 |  |  | 5642 | -44.006 |  | 6 | 95.1825 | 105.836 |  |  | 104.229 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.6 | 2021/05/18 03:07:15 | 95
05/18/21 | 82 | 20210048.058 | 617 | 0.848 | 6754 | 0.105 |  |  |  | 1 | 20.3709 | 124.335 |  |  |  | 123.441 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.5; BGD 29: 34.2; BGD 30: 1049.0 | 2021/05/18 03:17:00 | 0
05/18/21 | 82 | 20210048.058 | 617 | 0.848 | 6752 | 0 |  |  |  | 2 | 20.3823 | 124.404 |  |  |  | 123.509 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.4; BGD 30: 1175.5 | 2021/05/18 03:17:00 | 0
05/18/21 | 82 | 20210048.058 | 617 | 0.848 | 3887 | 10.602 |  |  | Nitrogen | 3 | 15.2653 | 93.296 |  |  |  | 92.618 | 0 |  | 0.7318 |  | 0.0037172 |  | 0.37034 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1144.9 | 2021/05/18 03:17:00 | 0
05/18/21 | 82 | 20210048.058 | 617 | 0.848 |  |  | 920 | -33.13 | Carbon | 4 | 55.6371 | 25.038 |  |  | 24.652 |  | 0 | 1.1481 |  | 0.01081 |  | 1.0694 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/18 03:17:00 | 95
05/18/21 | 82 | 20210048.058 | 617 | 0.848 |  |  | 5647 | -43.8 |  | 5 | 90.1701 | 105.557 |  |  | 103.954 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.2 | 2021/05/18 03:17:00 | 95
05/18/21 | 82 | 20210048.058 | 617 | 0.848 |  |  | 5647 | -44.018 |  | 6 | 89.8776 | 105.867 |  |  | 104.26 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.6 | 2021/05/18 03:17:00 | 95
05/18/21 | 83 | 20210048.059 | 619 | 0.839 | 6749 | 0.099 |  |  |  | 1 | 20.574 | 124.242 |  |  |  | 123.349 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.6; BGD 29: 34.2; BGD 30: 1051.6 | 2021/05/18 03:26:45 | 0
05/18/21 | 83 | 20210048.059 | 619 | 0.839 | 6751 | 0 |  |  |  | 2 | 20.5977 | 124.384 |  |  |  | 123.49 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.4; BGD 30: 1176.6 | 2021/05/18 03:26:45 | 0
05/18/21 | 83 | 20210048.059 | 619 | 0.839 | 4057 | 13.709 |  |  | Nitrogen | 3 | 16.1315 | 97.526 |  |  |  | 96.815 | 0 |  | 0.73405 |  | 0.0037286 |  | 0.37148 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1147.4 | 2021/05/18 03:26:45 | 0
05/18/21 | 83 | 20210048.059 | 619 | 0.839 |  |  | 863 | -32.012 | Carbon | 4 | 53.2527 | 23.421 |  |  | 23.06 |  | 0 | 1.1494 |  | 0.010822 |  | 1.0706 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/18 03:26:45 | 95
05/18/21 | 83 | 20210048.059 | 619 | 0.839 |  |  | 5640 | -43.8 |  | 5 | 91.203 | 105.488 |  |  | 103.886 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.2 | 2021/05/18 03:26:45 | 95
05/18/21 | 83 | 20210048.059 | 619 | 0.839 |  |  | 5634 | -44.016 |  | 6 | 90.9565 | 105.747 |  |  | 104.142 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.5 | 2021/05/18 03:26:45 | 95
05/18/21 | 84 | 20210048.060 | 623 | 0.817 | 6743 | 0.123 |  |  |  | 1 | 21.1396 | 124.31 |  |  |  | 123.416 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.5; BGD 29: 34.2; BGD 30: 1052.5 | 2021/05/18 03:36:30 | 0
05/18/21 | 84 | 20210048.060 | 623 | 0.817 | 6749 | 0 |  |  |  | 2 | 21.1354 | 124.285 |  |  |  | 123.392 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.4; BGD 30: 1177.7 | 2021/05/18 03:36:30 | 0
05/18/21 | 84 | 20210048.060 | 623 | 0.817 | 4097 | 12.631 |  |  | Nitrogen | 3 | 16.6989 | 98.305 |  |  |  | 97.59 | 0 |  | 0.73325 |  | 0.0037247 |  | 0.37108 | BGD 28: 44.3; BGD 29: 33.9; BGD 30: 1146.7 | 2021/05/18 03:36:30 | 0
05/18/21 | 84 | 20210048.060 | 623 | 0.817 |  |  | 865 | -32.421 | Carbon | 4 | 54.7073 | 23.432 |  |  | 23.071 |  | 0 | 1.1489 |  | 0.010818 |  | 1.0702 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/18 03:36:30 | 95
05/18/21 | 84 | 20210048.060 | 623 | 0.817 |  |  | 5636 | -43.8 |  | 5 | 93.6479 | 105.499 |  |  | 103.897 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.1 | 2021/05/18 03:36:30 | 95
05/18/21 | 84 | 20210048.060 | 623 | 0.817 |  |  | 5648 | -44.004 |  | 6 | 93.3799 | 105.773 |  |  | 104.167 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/18 03:36:30 | 95
05/18/21 | 85 | 20210048.061 | 625 | 0.813 | 6742 | 0.106 |  |  |  | 1 | 21.223 | 124.19 |  |  |  | 123.297 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.5; BGD 29: 34.2; BGD 30: 1053.3 | 2021/05/18 03:46:15 | 0
05/18/21 | 85 | 20210048.061 | 625 | 0.813 | 6745 | 0 |  |  |  | 2 | 21.2358 | 124.264 |  |  |  | 123.371 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.4; BGD 30: 1178.5 | 2021/05/18 03:46:15 | 0
05/18/21 | 85 | 20210048.061 | 625 | 0.813 | 4005 | 12.791 |  |  | Nitrogen | 3 | 16.4888 | 96.6 |  |  |  | 95.897 | 0 |  | 0.73338 |  | 0.0037252 |  | 0.37114 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1148.6 | 2021/05/18 03:46:15 | 0
05/18/21 | 85 | 20210048.061 | 625 | 0.813 |  |  | 855 | -32.506 | Carbon | 4 | 54.5092 | 23.191 |  |  | 22.833 |  | 0 | 1.1489 |  | 0.010817 |  | 1.0701 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/18 03:46:15 | 95
05/18/21 | 85 | 20210048.061 | 625 | 0.813 |  |  | 5637 | -43.8 |  | 5 | 94.188 | 105.418 |  |  | 103.817 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.2 | 2021/05/18 03:46:15 | 95
05/18/21 | 85 | 20210048.061 | 625 | 0.813 |  |  | 5639 | -44.014 |  | 6 | 93.8942 | 105.718 |  |  | 104.113 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/18 03:46:15 | 95
05/18/21 | 86 | 20210048.062 | 627 | 0.816 | 6741 | 0.123 |  |  |  | 1 | 21.1355 | 124.134 |  |  |  | 123.242 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.5; BGD 29: 34.1; BGD 30: 1054.1 | 2021/05/18 03:56:01 | 0
05/18/21 | 86 | 20210048.062 | 627 | 0.816 | 6739 | 0 |  |  |  | 2 | 21.147 | 124.201 |  |  |  | 123.308 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.3; BGD 30: 1179.4 | 2021/05/18 03:56:01 | 0
05/18/21 | 86 | 20210048.062 | 627 | 0.816 | 3791 | 12.504 |  |  | Nitrogen | 3 | 15.532 | 91.351 |  |  |  | 90.686 | 0 |  | 0.73316 |  | 0.0037242 |  | 0.37104 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1149.0 | 2021/05/18 03:56:01 | 0
05/18/21 | 86 | 20210048.062 | 627 | 0.816 |  |  | 861 | -33.077 | Carbon | 4 | 54.5964 | 23.34 |  |  | 22.98 |  | 0 | 1.1482 |  | 0.01081 |  | 1.0695 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/18 03:56:01 | 95
05/18/21 | 86 | 20210048.062 | 627 | 0.816 |  |  | 5634 | -43.8 |  | 5 | 93.8938 | 105.364 |  |  | 103.764 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.2 | 2021/05/18 03:56:01 | 95
05/18/21 | 86 | 20210048.062 | 627 | 0.816 |  |  | 5632 | -44.013 |  | 6 | 93.6652 | 105.599 |  |  | 103.996 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/18 03:56:01 | 95
05/18/21 | 87 | 20210048.063 | 629 | 0.843 | 6743 | 0.124 |  |  |  | 1 | 20.4505 | 124.085 |  |  |  | 123.193 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.5; BGD 29: 34.2; BGD 30: 1054.5 | 2021/05/18 04:05:46 | 0
05/18/21 | 87 | 20210048.063 | 629 | 0.843 | 6739 | 0 |  |  |  | 2 | 20.4613 | 124.151 |  |  |  | 123.259 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.3; BGD 30: 1179.6 | 2021/05/18 04:05:46 | 0
05/18/21 | 87 | 20210048.063 | 629 | 0.843 | 4210 | 10.196 |  |  | Nitrogen | 3 | 16.6645 | 101.212 |  |  |  | 100.477 | 0 |  | 0.73149 |  | 0.0037157 |  | 0.37019 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1148.5 | 2021/05/18 04:05:46 | 0
05/18/21 | 87 | 20210048.063 | 629 | 0.843 |  |  | 873 | -32.288 | Carbon | 4 | 53.5029 | 23.692 |  |  | 23.327 |  | 0 | 1.1491 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/18 04:05:46 | 95
05/18/21 | 87 | 20210048.063 | 629 | 0.843 |  |  | 5636 | -43.8 |  | 5 | 90.8849 | 105.366 |  |  | 103.766 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.3 | 2021/05/18 04:05:46 | 95
05/18/21 | 87 | 20210048.063 | 629 | 0.843 |  |  | 5627 | -44.025 |  | 6 | 90.7208 | 105.54 |  |  | 103.938 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.6 | 2021/05/18 04:05:46 | 95
05/18/21 | 88 | 20210048.064 | 631 | 0.825 | 6738 | 0.128 |  |  |  | 1 | 20.8966 | 124.085 |  |  |  | 123.193 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.5; BGD 29: 34.2; BGD 30: 1055.1 | 2021/05/18 04:15:31 | 0
05/18/21 | 88 | 20210048.064 | 631 | 0.825 | 6736 | 0 |  |  |  | 2 | 20.9063 | 124.142 |  |  |  | 123.25 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.3; BGD 30: 1180.8 | 2021/05/18 04:15:31 | 0
05/18/21 | 88 | 20210048.064 | 631 | 0.825 | 3648 | 12.644 |  |  | Nitrogen | 3 | 14.7453 | 87.694 |  |  |  | 87.056 | 0 |  | 0.73326 |  | 0.0037247 |  | 0.37109 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1149.9 | 2021/05/18 04:15:31 | 0
05/18/21 | 88 | 20210048.064 | 631 | 0.825 |  |  | 901 | -32.41 | Carbon | 4 | 56.179 | 24.495 |  |  | 24.117 |  | 0 | 1.1489 |  | 0.010818 |  | 1.0702 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/18 04:15:31 | 95
05/18/21 | 88 | 20210048.064 | 631 | 0.825 |  |  | 5631 | -43.8 |  | 5 | 92.8925 | 105.34 |  |  | 103.741 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.3 | 2021/05/18 04:15:31 | 95
05/18/21 | 88 | 20210048.064 | 631 | 0.825 |  |  | 5635 | -44.003 |  | 6 | 92.655 | 105.587 |  |  | 103.984 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.7 | 2021/05/18 04:15:31 | 95
05/18/21 | 89 | 20210048.065 | 635 | 0.821 | 6736 | 0.116 |  |  |  | 1 | 20.9606 | 123.863 |  |  |  | 122.972 | 0 |  | 0.7242 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.5; BGD 29: 34.1; BGD 30: 1054.4 | 2021/05/18 04:25:16 | 0
05/18/21 | 89 | 20210048.065 | 635 | 0.821 | 6738 | 0 |  |  |  | 2 | 21.0043 | 124.12 |  |  |  | 123.228 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.4; BGD 30: 1180.5 | 2021/05/18 04:25:16 | 0
05/18/21 | 89 | 20210048.065 | 635 | 0.821 | 4069 | 11.307 |  |  | Nitrogen | 3 | 16.5324 | 97.804 |  |  |  | 97.093 | 0 |  | 0.7323 |  | 0.0037198 |  | 0.3706 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1150.7 | 2021/05/18 04:25:16 | 0
05/18/21 | 89 | 20210048.065 | 635 | 0.821 |  |  | 878 | -32.128 | Carbon | 4 | 55.2405 | 23.852 |  |  | 23.484 |  | 0 | 1.1493 |  | 0.010821 |  | 1.0705 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.2 | 2021/05/18 04:25:16 | 95
05/18/21 | 89 | 20210048.065 | 635 | 0.821 |  |  | 5640 | -43.8 |  | 5 | 93.2876 | 105.4 |  |  | 103.799 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.3 | 2021/05/18 04:25:16 | 95
05/18/21 | 89 | 20210048.065 | 635 | 0.821 |  |  | 5625 | -44.023 |  | 6 | 93.137 | 105.555 |  |  | 103.953 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 8.8 | 2021/05/18 04:25:16 | 95
05/18/21 | 90 | 20210048.066 | 637 | 0.836 | 6739 | 0.102 |  |  |  | 1 | 20.6235 | 124.097 |  |  |  | 123.204 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.5; BGD 29: 34.1; BGD 30: 1056.7 | 2021/05/18 04:35:01 | 0
05/18/21 | 90 | 20210048.066 | 637 | 0.836 | 6738 | 0 |  |  |  | 2 | 20.6264 | 124.114 |  |  |  | 123.222 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.3; BGD 30: 1181.7 | 2021/05/18 04:35:01 | 0
05/18/21 | 90 | 20210048.066 | 637 | 0.836 | 4090 | 12.918 |  |  | Nitrogen | 3 | 16.3607 | 98.554 |  |  |  | 97.836 | 0 |  | 0.73347 |  | 0.0037257 |  | 0.37119 | BGD 28: 44.3; BGD 29: 33.9; BGD 30: 1152.1 | 2021/05/18 04:35:01 | 0
05/18/21 | 90 | 20210048.066 | 637 | 0.836 |  |  | 888 | -32.717 | Carbon | 4 | 54.7951 | 24.146 |  |  | 23.773 |  | 0 | 1.1486 |  | 0.010814 |  | 1.0699 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.2 | 2021/05/18 04:35:01 | 95
05/18/21 | 90 | 20210048.066 | 637 | 0.836 |  |  | 5634 | -43.8 |  | 5 | 91.6033 | 105.411 |  |  | 103.81 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.3 | 2021/05/18 04:35:01 | 95
05/18/21 | 90 | 20210048.066 | 637 | 0.836 |  |  | 5626 | -44.004 |  | 6 | 91.4416 | 105.581 |  |  | 103.978 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.8 | 2021/05/18 04:35:01 | 95
05/18/21 | 91 | 20210048.067 | 639 | 0.83 | 6737 | 0.122 |  |  |  | 1 | 20.7701 | 124.082 |  |  |  | 123.19 | 0 |  | 0.7242 |  | 0.0036786 |  | 0.36652 | BGD 28: 44.5; BGD 29: 34.1; BGD 30: 1057.9 | 2021/05/18 04:44:47 | 0
05/18/21 | 91 | 20210048.067 | 639 | 0.83 | 6732 | 0 |  |  |  | 2 | 20.767 | 124.063 |  |  |  | 123.171 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.5; BGD 29: 33.3; BGD 30: 1183.3 | 2021/05/18 04:44:47 | 0
05/18/21 | 91 | 20210048.067 | 639 | 0.83 | 3947 | 13.182 |  |  | Nitrogen | 3 | 15.8872 | 95.028 |  |  |  | 94.336 | 0 |  | 0.73365 |  | 0.0037267 |  | 0.37128 | BGD 28: 44.4; BGD 29: 33.9; BGD 30: 1153.0 | 2021/05/18 04:44:47 | 0
05/18/21 | 91 | 20210048.067 | 639 | 0.83 |  |  | 875 | -32.597 | Carbon | 4 | 54.5387 | 23.797 |  |  | 23.43 |  | 0 | 1.1488 |  | 0.010816 |  | 1.07 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/18 04:44:47 | 95
05/18/21 | 91 | 20210048.067 | 639 | 0.83 |  |  | 5637 | -43.8 |  | 5 | 92.3616 | 105.31 |  |  | 103.711 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.3 | 2021/05/18 04:44:47 | 95
05/18/21 | 91 | 20210048.067 | 639 | 0.83 |  |  | 5637 | -44.024 |  | 6 | 92.0121 | 105.675 |  |  | 104.071 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.8 | 2021/05/18 04:44:47 | 95
05/18/21 | 92 | 20210048.068 | 641 | 0.848 | 6733 | 0.117 |  |  |  | 1 | 20.2941 | 123.868 |  |  |  | 122.977 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.4; BGD 29: 34.1; BGD 30: 1057.7 | 2021/05/18 04:54:32 | 0
05/18/21 | 92 | 20210048.068 | 641 | 0.848 | 6742 | 0 |  |  |  | 2 | 20.3343 | 124.113 |  |  |  | 123.22 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.3; BGD 30: 1182.5 | 2021/05/18 04:54:32 | 0
05/18/21 | 92 | 20210048.068 | 641 | 0.848 | 4252 | 10.997 |  |  | Nitrogen | 3 | 16.7774 | 102.497 |  |  |  | 101.752 | 0 |  | 0.73207 |  | 0.0037186 |  | 0.37049 | BGD 28: 44.3; BGD 29: 33.8; BGD 30: 1152.8 | 2021/05/18 04:54:32 | 0
05/18/21 | 92 | 20210048.068 | 641 | 0.848 |  |  | 870 | -32.064 | Carbon | 4 | 53.0893 | 23.639 |  |  | 23.274 |  | 0 | 1.1493 |  | 0.010822 |  | 1.0706 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.2 | 2021/05/18 04:54:32 | 95
05/18/21 | 92 | 20210048.068 | 641 | 0.848 |  |  | 5634 | -43.8 |  | 5 | 90.4069 | 105.304 |  |  | 103.705 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.4 | 2021/05/18 04:54:32 | 95
05/18/21 | 92 | 20210048.068 | 641 | 0.848 |  |  | 5622 | -44.006 |  | 6 | 90.1567 | 105.571 |  |  | 103.969 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.8 | 2021/05/18 04:54:32 | 95
05/18/21 | 93 | 20210048.068 | Replicate H8 | 0.828 | 6731 | 0.139 |  |  |  | 1 | 20.7919 | 123.913 |  |  |  | 123.022 | 0 |  | 0.72419 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.4; BGD 29: 34.1; BGD 30: 1058.8 | 2021/05/18 05:04:17 | 0
05/18/21 | 93 | 20210048.068 | Replicate H8 | 0.828 | 6735 | 0 |  |  |  | 2 | 20.7988 | 123.954 |  |  |  | 123.063 | 1 |  | 0.72409 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.3; BGD 30: 1183.9 | 2021/05/18 05:04:17 | 0
05/18/21 | 93 | 20210048.068 | Replicate H8 | 0.828 | 4218 | 11.025 |  |  | Nitrogen | 3 | 16.9864 | 101.332 |  |  |  | 100.595 | 0 |  | 0.73208 |  | 0.0037188 |  | 0.3705 | BGD 28: 44.3; BGD 29: 33.8; BGD 30: 1153.8 | 2021/05/18 05:04:17 | 0
05/18/21 | 93 | 20210048.068 | Replicate H8 | 0.828 |  |  | 860 | -32.029 | Carbon | 4 | 53.753 | 23.312 |  |  | 22.953 |  | 0 | 1.1494 |  | 0.010822 |  | 1.0706 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/18 05:04:17 | 95
05/18/21 | 93 | 20210048.068 | Replicate H8 | 0.828 |  |  | 5627 | -43.8 |  | 5 | 92.639 | 105.253 |  |  | 103.655 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.3 | 2021/05/18 05:04:17 | 95
05/18/21 | 93 | 20210048.068 | Replicate H8 | 0.828 |  |  | 5619 | -44.017 |  | 6 | 92.3973 | 105.506 |  |  | 103.904 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 8.8 | 2021/05/18 05:04:17 | 95
05/18/21 | 94 | 36-UWSIF-UT Glut 1-20210048.111 | 36-UWSIF-UT Glut 1- | 0.754 | 6721 | 0.117 |  |  |  | 1 | 22.8245 | 123.87 |  |  |  | 122.98 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.4; BGD 29: 34.1; BGD 30: 1059.1 | 2021/05/18 05:14:03 | 0
05/18/21 | 94 | 36-UWSIF-UT Glut 1-20210048.111 | 36-UWSIF-UT Glut 1- | 0.754 | 6731 | 0 |  |  |  | 2 | 22.8464 | 123.989 |  |  |  | 123.097 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.4; BGD 29: 33.2; BGD 30: 1184.6 | 2021/05/18 05:14:03 | 0
05/18/21 | 94 | 36-UWSIF-UT Glut 1-20210048.111 | 36-UWSIF-UT Glut 1- | 0.754 | 2364 | -2.858 |  |  | Nitrogen | 3 | 10.5422 | 57.376 |  |  |  | 56.965 | 0 |  | 0.72203 |  | 0.0036677 |  | 0.36543 | BGD 28: 44.3; BGD 29: 33.8; BGD 30: 1152.1 | 2021/05/18 05:14:03 | 0
05/18/21 | 94 | 36-UWSIF-UT Glut 1-20210048.111 | 36-UWSIF-UT Glut 1- | 0.754 |  |  | 664 | -29.871 | Carbon | 4 | 46.8165 | 17.744 |  |  | 17.47 |  | 0 | 1.1518 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 9.0 | 2021/05/18 05:14:03 | 95
05/18/21 | 94 | 36-UWSIF-UT Glut 1-20210048.111 | 36-UWSIF-UT Glut 1- | 0.754 |  |  | 5623 | -43.8 |  | 5 | 101.7173 | 105.266 |  |  | 103.667 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/18 05:14:03 | 95
05/18/21 | 94 | 36-UWSIF-UT Glut 1-20210048.111 | 36-UWSIF-UT Glut 1- | 0.754 |  |  | 5630 | -44.036 |  | 6 | 101.4495 | 105.521 |  |  | 103.919 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/18 05:14:03 | 95
05/18/21 | 95 | 36-UWSIF-UT Glut 1-20210048.112 | 36-UWSIF-UT Glut 1- | 0.787 | 6729 | 0.117 |  |  |  | 1 | 21.8635 | 123.848 |  |  |  | 122.957 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.2; BGD 29: 33.8; BGD 30: 1050.2 | 2021/05/18 05:23:48 | 0
05/18/21 | 95 | 36-UWSIF-UT Glut 1-20210048.112 | 36-UWSIF-UT Glut 1- | 0.787 | 6727 | 0 |  |  |  | 2 | 21.8703 | 123.886 |  |  |  | 122.996 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.2; BGD 29: 33.0; BGD 30: 1176.9 | 2021/05/18 05:23:48 | 0
05/18/21 | 95 | 36-UWSIF-UT Glut 1-20210048.112 | 36-UWSIF-UT Glut 1- | 0.787 | 2474 | -2.84 |  |  | Nitrogen | 3 | 10.5453 | 59.898 |  |  |  | 59.469 | 0 |  | 0.72205 |  | 0.0036678 |  | 0.36544 | BGD 28: 44.1; BGD 29: 33.6; BGD 30: 1143.7 | 2021/05/18 05:23:48 | 0
05/18/21 | 95 | 36-UWSIF-UT Glut 1-20210048.112 | 36-UWSIF-UT Glut 1- | 0.787 |  |  | 692 | -29.818 | Carbon | 4 | 46.6696 | 18.574 |  |  | 18.287 |  | 0 | 1.1519 |  | 0.010847 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/18 05:23:48 | 95
05/18/21 | 95 | 36-UWSIF-UT Glut 1-20210048.112 | 36-UWSIF-UT Glut 1- | 0.787 |  |  | 5631 | -43.8 |  | 5 | 97.3849 | 105.333 |  |  | 103.733 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/18 05:23:48 | 95
05/18/21 | 95 | 36-UWSIF-UT Glut 1-20210048.112 | 36-UWSIF-UT Glut 1- | 0.787 |  |  | 5627 | -44.018 |  | 6 | 97.2603 | 105.457 |  |  | 103.856 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/18 05:23:48 | 95
05/18/21 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 39-UWSIF-UW Glut 2- | 0.703 | 6716 | 0.11 |  |  |  | 1 | 24.4741 | 123.839 |  |  |  | 122.948 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.1; BGD 29: 33.8; BGD 30: 1047.9 | 2021/05/18 05:33:33 | 0
05/18/21 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 39-UWSIF-UW Glut 2- | 0.703 | 6724 | 0 |  |  |  | 2 | 24.4856 | 123.897 |  |  |  | 123.006 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1173.5 | 2021/05/18 05:33:33 | 0
05/18/21 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 39-UWSIF-UW Glut 2- | 0.703 | 2394 | 29.779 |  |  | Nitrogen | 3 | 11.4688 | 58.195 |  |  |  | 57.764 | 0 |  | 0.74567 |  | 0.0037877 |  | 0.37734 | BGD 28: 44.1; BGD 29: 33.6; BGD 30: 1140.8 | 2021/05/18 05:33:33 | 0
05/18/21 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 39-UWSIF-UW Glut 2- | 0.703 |  |  | 670 | 22.266 | Carbon | 4 | 50.7657 | 17.968 |  |  | 17.681 |  | 0 | 1.2093 |  | 0.011429 |  | 1.13 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/18 05:33:33 | 95
05/18/21 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 39-UWSIF-UW Glut 2- | 0.703 |  |  | 5625 | -43.8 |  | 5 | 109.1849 | 105.187 |  |  | 103.59 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/18 05:33:33 | 95
05/18/21 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 39-UWSIF-UW Glut 2- | 0.703 |  |  | 5618 | -44.025 |  | 6 | 108.8802 | 105.458 |  |  | 103.857 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/18 05:33:33 | 95
05/18/21 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 39-UWSIF-UW Glut 2- | 0.83 | 6727 | 0.107 |  |  |  | 1 | 20.7247 | 123.812 |  |  |  | 122.922 | 0 |  | 0.72418 |  | 0.0036786 |  | 0.36651 | BGD 28: 44.1; BGD 29: 33.8; BGD 30: 1044.7 | 2021/05/18 05:43:18 | 0
05/18/21 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 39-UWSIF-UW Glut 2- | 0.83 | 6725 | 0 |  |  |  | 2 | 20.7186 | 123.776 |  |  |  | 122.886 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.0; BGD 29: 33.0; BGD 30: 1170.5 | 2021/05/18 05:43:18 | 0
05/18/21 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 39-UWSIF-UW Glut 2- | 0.83 | 2829 | 29.692 |  |  | Nitrogen | 3 | 11.4373 | 68.489 |  |  |  | 67.982 | 0 |  | 0.74561 |  | 0.0037874 |  | 0.37731 | BGD 28: 44.0; BGD 29: 33.6; BGD 30: 1139.0 | 2021/05/18 05:43:18 | 0
05/18/21 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 39-UWSIF-UW Glut 2- | 0.83 |  |  | 787 | 22.35 | Carbon | 4 | 49.7711 | 21.316 |  |  | 20.975 |  | 0 | 1.2094 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/18 05:43:18 | 95
05/18/21 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 39-UWSIF-UW Glut 2- | 0.83 |  |  | 5622 | -43.8 |  | 5 | 92.5502 | 105.111 |  |  | 103.515 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/18 05:43:18 | 95
05/18/21 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 39-UWSIF-UW Glut 2- | 0.83 |  |  | 5625 | -44.017 |  | 6 | 92.3076 | 105.367 |  |  | 103.767 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/18 05:43:18 | 95
05/18/21 | 98 | 01-UWSIF-Liver-20210048.17 | 01-UWSIF-Liver- | 0.753 | 6721 | 0.132 |  |  |  | 1 | 22.8166 | 123.664 |  |  |  | 122.775 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.1; BGD 29: 33.8; BGD 30: 1044.3 | 2021/05/18 05:53:04 | 0
05/18/21 | 98 | 01-UWSIF-Liver-20210048.17 | 01-UWSIF-Liver- | 0.753 | 6711 | 0 |  |  |  | 2 | 22.8463 | 123.824 |  |  |  | 122.934 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1170.7 | 2021/05/18 05:53:04 | 0
05/18/21 | 98 | 01-UWSIF-Liver-20210048.17 | 01-UWSIF-Liver- | 0.753 | 3279 | 8.488 |  |  | Nitrogen | 3 | 14.5628 | 79.08 |  |  |  | 78.506 | 0 |  | 0.73025 |  | 0.0037094 |  | 0.36957 | BGD 28: 44.0; BGD 29: 33.6; BGD 30: 1138.6 | 2021/05/18 05:53:04 | 0
05/18/21 | 98 | 01-UWSIF-Liver-20210048.17 | 01-UWSIF-Liver- | 0.753 |  |  | 818 | -19.536 | Carbon | 4 | 56.6475 | 22.148 |  |  | 21.803 |  | 0 | 1.1632 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/18 05:53:04 | 95
05/18/21 | 98 | 01-UWSIF-Liver-20210048.17 | 01-UWSIF-Liver- | 0.753 |  |  | 5626 | -43.8 |  | 5 | 101.9817 | 105.142 |  |  | 103.545 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.2 | 2021/05/18 05:53:04 | 95
05/18/21 | 98 | 01-UWSIF-Liver-20210048.17 | 01-UWSIF-Liver- | 0.753 |  |  | 5624 | -44.017 |  | 6 | 101.6411 | 105.467 |  |  | 103.866 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/18 05:53:04 | 95
05/18/21 | 99 | 01-UWSIF-Liver-20210048.18 | 01-UWSIF-Liver- | 0.735 | 6722 | 0.123 |  |  |  | 1 | 23.4052 | 123.821 |  |  |  | 122.931 | 0 |  | 0.7242 |  | 0.0036787 |  | 0.36652 | BGD 28: 44.2; BGD 29: 33.9; BGD 30: 1047.8 | 2021/05/18 06:02:50 | 0
05/18/21 | 99 | 01-UWSIF-Liver-20210048.18 | 01-UWSIF-Liver- | 0.735 | 6721 | 0 |  |  |  | 2 | 23.4074 | 123.833 |  |  |  | 122.942 | 1 |  | 0.72411 |  | 0.0036782 |  | 0.36647 | BGD 28: 43.2; BGD 29: 33.1; BGD 30: 1172.6 | 2021/05/18 06:02:50 | 0
05/18/21 | 99 | 01-UWSIF-Liver-20210048.18 | 01-UWSIF-Liver- | 0.735 | 3199 | 8.539 |  |  | Nitrogen | 3 | 14.5707 | 77.237 |  |  |  | 76.677 | 0 |  | 0.73029 |  | 0.0037096 |  | 0.36959 | BGD 28: 44.1; BGD 29: 33.6; BGD 30: 1141.2 | 2021/05/18 06:02:50 | 0
05/18/21 | 99 | 01-UWSIF-Liver-20210048.18 | 01-UWSIF-Liver- | 0.735 |  |  | 801 | -19.438 | Carbon | 4 | 56.9301 | 21.645 |  |  | 21.308 |  | 0 | 1.1633 |  | 0.010963 |  | 1.0844 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/18 06:02:50 | 95
05/18/21 | 99 | 01-UWSIF-Liver-20210048.18 | 01-UWSIF-Liver- | 0.735 |  |  | 5620 | -43.8 |  | 5 | 104.4355 | 105.183 |  |  | 103.586 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.3 | 2021/05/18 06:02:50 | 95
05/18/21 | 99 | 01-UWSIF-Liver-20210048.18 | 01-UWSIF-Liver- | 0.735 |  |  | 5625 | -44.034 |  | 6 | 104.119 | 105.477 |  |  | 103.876 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.7 | 2021/05/18 06:02:50 | 95

## Original 2
Date | Row | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Component | Peak Nr | Amt% | Area All | Comment | Preparation | Area 44 | Area 28 | Is Ref _ | rR 45CO2/44CO2 | rR 29N2/28N2 | R 13C/12C | R 15N/14N | AT% 13C/12C | AT% 15N/14N | BGD 28 | BGD 29 | BGD 30 | BGD 44 | BGD 45 | BGD 46 | Time Code | Sample Dilution
05/18/21 | 1 | 36-UWSIF-UT Glut 1-20210048.21 | 36-UWSIF-UT Glut 1- | 0.769 | 6697 | 0.129 |  |  |  | 1 | 22.2866 | 123.36 | JOB: 2021-0048 Tray 2 Nelson |  |  | 122.473 | 0 |  | 0.72429 |  | 0.0036787 |  | 0.36652 | BGD 28: 40.6; BGD 29: 30.5; BGD 30: 852.5 | 2021/05/18 10:51:49 | 0
05/18/21 | 1 | 36-UWSIF-UT Glut 1-20210048.21 | 36-UWSIF-UT Glut 1- | 0.769 | 6697 | 0 |  |  |  | 2 | 22.2821 | 123.335 | JOB: 2021-0048 Tray 2 Nelson |  |  | 122.448 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.8; BGD 29: 29.9; BGD 30: 975.7 | 2021/05/18 10:51:49 | 0
05/18/21 | 1 | 36-UWSIF-UT Glut 1-20210048.21 | 36-UWSIF-UT Glut 1- | 0.769 | 2461 | -2.774 |  |  | Nitrogen | 3 | 10.5236 | 58.412 | JOB: 2021-0048 Tray 2 Nelson |  |  | 57.993 | 0 |  | 0.72219 |  | 0.003668 |  | 0.36546 | BGD 28: 41.0; BGD 29: 30.8; BGD 30: 941.8 | 2021/05/18 10:51:49 | 0
05/18/21 | 1 | 36-UWSIF-UT Glut 1-20210048.21 | 36-UWSIF-UT Glut 1- | 0.769 |  |  | 685 | -29.792 | Carbon | 4 | 46.903 | 18.189 | JOB: 2021-0048 Tray 2 Nelson |  | 17.908 |  | 0 | 1.1516 |  | 0.010847 |  | 1.0731 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 8.0 | 2021/05/18 10:51:49 | 95
05/18/21 | 1 | 36-UWSIF-UT Glut 1-20210048.21 | 36-UWSIF-UT Glut 1- | 0.769 |  |  | 5599 | -43.8 |  | 5 | 100.4629 | 104.544 | JOB: 2021-0048 Tray 2 Nelson |  | 102.957 |  | 1 | 1.1349 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 6.0 | 2021/05/18 10:51:49 | 95
05/18/21 | 1 | 36-UWSIF-UT Glut 1-20210048.21 | 36-UWSIF-UT Glut 1- | 0.769 |  |  | 5595 | -44.116 |  | 6 | 100.0962 | 104.909 | JOB: 2021-0048 Tray 2 Nelson |  | 103.317 |  | 0 | 1.1346 |  | 0.010687 |  | 1.0574 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 7.4 | 2021/05/18 10:51:49 | 95
05/18/21 | 2 | 36-UWSIF-UT Glut 1-20210048.22 | 36-UWSIF-UT Glut 1- | 0.824 | 6679 | 0.137 |  |  |  | 1 | 20.7313 | 122.96 | ANALYST: sme |  |  | 122.076 | 0 |  | 0.72425 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.4; BGD 29: 32.0; BGD 30: 867.5 | 2021/05/18 11:01:35 | 0
05/18/21 | 2 | 36-UWSIF-UT Glut 1-20210048.22 | 36-UWSIF-UT Glut 1- | 0.824 | 6675 | 0 |  |  |  | 2 | 20.7427 | 123.028 | ANALYST: sme |  |  | 122.143 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.4; BGD 29: 31.3; BGD 30: 990.1 | 2021/05/18 11:01:35 | 0
05/18/21 | 2 | 36-UWSIF-UT Glut 1-20210048.22 | 36-UWSIF-UT Glut 1- | 0.824 | 2532 | -2.754 |  |  | Nitrogen | 3 | 10.4061 | 61.881 | ANALYST: sme |  |  | 61.437 | 0 |  | 0.72216 |  | 0.0036681 |  | 0.36547 | BGD 28: 42.4; BGD 29: 32.0; BGD 30: 955.5 | 2021/05/18 11:01:35 | 0
05/18/21 | 2 | 36-UWSIF-UT Glut 1-20210048.22 | 36-UWSIF-UT Glut 1- | 0.824 |  |  | 710 | -29.974 | Carbon | 4 | 45.9465 | 19.239 | ANALYST: sme |  | 18.942 |  | 0 | 1.1516 |  | 0.010845 |  | 1.0729 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.6 | 2021/05/18 11:01:35 | 95
05/18/21 | 2 | 36-UWSIF-UT Glut 1-20210048.22 | 36-UWSIF-UT Glut 1- | 0.824 |  |  | 5582 | -43.8 |  | 5 | 93.9073 | 104.382 | ANALYST: sme |  | 102.797 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 5.9 | 2021/05/18 11:01:35 | 95
05/18/21 | 2 | 36-UWSIF-UT Glut 1-20210048.22 | 36-UWSIF-UT Glut 1- | 0.824 |  |  | 5580 | -44.047 |  | 6 | 93.6662 | 104.642 | ANALYST: sme |  | 103.053 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 7.2 | 2021/05/18 11:01:35 | 95
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 | 6664 | 0.141 |  |  |  | 1 | 9.52 | 122.79 | SAMPLE COUNT: 0108 |  |  | 121.907 | 0 |  | 0.72426 |  | 0.0036787 |  | 0.36652 | BGD 28: 43.0; BGD 29: 32.6; BGD 30: 876.8 | 2021/05/18 11:11:20 | 0
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 | 6666 | 0 |  |  |  | 2 | 9.52 | 122.773 | SAMPLE COUNT: 0108 |  |  | 121.89 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.0; BGD 29: 31.9; BGD 30: 998.3 | 2021/05/18 11:11:20 | 0
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 | 1441 | -2.877 |  |  | Nitrogen | 3 | 9.52 | 35.554 | SAMPLE COUNT: 0108 |  |  | 35.3 | 0 |  | 0.72207 |  | 0.0036676 |  | 0.36542 | BGD 28: 43.0; BGD 29: 32.5; BGD 30: 962.4 | 2021/05/18 11:11:20 | 0
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 |  |  | 407 | -29.845 | Carbon | 4 | 40.81 | 10.822 | SAMPLE COUNT: 0108 |  | 10.655 |  | 0 | 1.1519 |  | 0.010847 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.5 | 2021/05/18 11:11:20 | 95
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 |  |  | 5574 | -43.8 |  | 5 | 40.81 | 104.303 | SAMPLE COUNT: 0108 |  | 102.719 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.5; BGD 45: 0.6; BGD 46: 5.6 | 2021/05/18 11:11:20 | 95
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 |  |  | 5571 | -44.047 |  | 6 | 40.81 | 104.529 | SAMPLE COUNT: 0108 |  | 102.942 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 7.0 | 2021/05/18 11:11:20 | 95
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 | 6645 | 0.128 |  |  |  | 1 | 9.52 | 122.414 | HE PSI: 1900 |  |  | 121.533 | 0 |  | 0.72425 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.8; BGD 29: 32.4; BGD 30: 880.7 | 2021/05/18 11:21:04 | 0
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 | 6650 | 0 |  |  |  | 2 | 9.52 | 122.599 | HE PSI: 1900 |  |  | 121.718 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 31.7; BGD 30: 1002.6 | 2021/05/18 11:21:04 | 0
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 | 2250 | -2.812 |  |  | Nitrogen | 3 | 9.52 | 55.169 | HE PSI: 1900 |  |  | 54.774 | 0 |  | 0.72212 |  | 0.0036679 |  | 0.36545 | BGD 28: 42.8; BGD 29: 32.3; BGD 30: 969.5 | 2021/05/18 11:21:04 | 0
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 |  |  | 634 | -29.899 | Carbon | 4 | 40.81 | 17.1 | HE PSI: 1900 |  | 16.836 |  | 0 | 1.1518 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/18 11:21:04 | 95
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 |  |  | 5573 | -43.8 |  | 5 | 40.81 | 104.189 | HE PSI: 1900 |  | 102.607 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 5.8 | 2021/05/18 11:21:04 | 95
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 |  |  | 5567 | -44.048 |  | 6 | 40.81 | 104.526 | HE PSI: 1900 |  | 102.939 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 7.2 | 2021/05/18 11:21:04 | 95
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 | 6637 | 0.089 |  |  |  | 1 | 9.52 | 122.231 | O2 PSI: 500 |  |  | 121.352 | 0 |  | 0.72424 |  | 0.0036785 |  | 0.36651 | BGD 28: 43.0; BGD 29: 32.7; BGD 30: 892.5 | 2021/05/18 11:30:49 | 0
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 | 6645 | 0 |  |  |  | 2 | 9.52 | 122.38 | O2 PSI: 500 |  |  | 121.5 | 1 |  | 0.72418 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.0; BGD 29: 31.9; BGD 30: 1013.9 | 2021/05/18 11:30:49 | 0
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 | 4838 | -2.865 |  |  | Nitrogen | 3 | 9.52 | 115.77 | O2 PSI: 500 |  |  | 114.94 | 0 |  | 0.7221 |  | 0.0036677 |  | 0.36543 | BGD 28: 43.0; BGD 29: 32.5; BGD 30: 986.5 | 2021/05/18 11:30:49 | 0
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 |  |  | 1304 | -30.108 | Carbon | 4 | 40.81 | 36.572 | O2 PSI: 500 |  | 36.008 |  | 0 | 1.1515 |  | 0.010844 |  | 1.0727 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.9 | 2021/05/18 11:30:49 | 95
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 |  |  | 5567 | -43.8 |  | 5 | 40.81 | 104.099 | O2 PSI: 500 |  | 102.518 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 6.4 | 2021/05/18 11:30:49 | 95
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 |  |  | 5557 | -44.016 |  | 6 | 40.81 | 104.262 | O2 PSI: 500 |  | 102.679 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 7.7 | 2021/05/18 11:30:49 | 95
05/18/21 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 39-UWSIF-UW Glut 2- | 0.742 | 6639 | 0.132 |  |  |  | 1 | 15.5935 | 122.105 | MASS 28: |  |  | 121.227 | 0 |  | 0.72424 |  | 0.0036787 |  | 0.36652 | BGD 28: 43.7; BGD 29: 33.4; BGD 30: 916.0 | 2021/05/18 11:40:34 | 0
05/18/21 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 39-UWSIF-UW Glut 2- | 0.742 | 6632 | 0 |  |  |  | 2 | 15.6008 | 122.229 | MASS 28: |  |  | 121.35 | 1 |  | 0.72414 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.5; BGD 29: 32.5; BGD 30: 1037.0 | 2021/05/18 11:40:34 | 0
05/18/21 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 39-UWSIF-UW Glut 2- | 0.742 | 2482 | 29.77 |  |  | Nitrogen | 3 | 9.8299 | 60.675 | MASS 28: |  |  | 60.225 | 0 |  | 0.7457 |  | 0.0037877 |  | 0.37734 | BGD 28: 43.4; BGD 29: 33.0; BGD 30: 1003.7 | 2021/05/18 11:40:34 | 0
05/18/21 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 39-UWSIF-UW Glut 2- | 0.742 |  |  | 698 | 22.086 | Carbon | 4 | 46.7862 | 18.924 | MASS 28: |  | 18.621 |  | 0 | 1.2094 |  | 0.011427 |  | 1.1298 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.8 | 2021/05/18 11:40:34 | 95
05/18/21 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 39-UWSIF-UW Glut 2- | 0.742 |  |  | 5560 | -43.8 |  | 5 | 56.615 | 103.975 | MASS 28: |  | 102.396 |  | 1 | 1.1355 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.1 | 2021/05/18 11:40:34 | 95
05/18/21 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 39-UWSIF-UW Glut 2- | 0.742 |  |  | 5557 | -44.038 |  | 6 | 56.2459 | 104.178 | MASS 28: |  | 102.596 |  | 0 | 1.1353 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 7.5 | 2021/05/18 11:40:34 | 95
05/18/21 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 39-UWSIF-UW Glut 2- | 0.736 | 6634 | 0.124 |  |  |  | 1 | 15.7156 | 122.021 | MASS 29: |  |  | 121.144 | 0 |  | 0.72424 |  | 0.0036787 |  | 0.36652 | BGD 28: 43.2; BGD 29: 33.0; BGD 30: 920.6 | 2021/05/18 11:50:19 | 0
05/18/21 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 39-UWSIF-UW Glut 2- | 0.736 | 6632 | 0 |  |  |  | 2 | 15.7181 | 122.063 | MASS 29: |  |  | 121.185 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.2; BGD 29: 32.1; BGD 30: 1042.4 | 2021/05/18 11:50:19 | 0
05/18/21 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 39-UWSIF-UW Glut 2- | 0.736 | 2330 | 29.795 |  |  | Nitrogen | 3 | 9.4387 | 57.082 | MASS 29: |  |  | 56.659 | 0 |  | 0.74573 |  | 0.0037878 |  | 0.37735 | BGD 28: 43.1; BGD 29: 32.6; BGD 30: 1011.8 | 2021/05/18 11:50:19 | 0
05/18/21 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 39-UWSIF-UW Glut 2- | 0.736 |  |  | 657 | 22.324 | Carbon | 4 | 44.759 | 17.769 | MASS 29: |  | 17.485 |  | 0 | 1.2095 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.9 | 2021/05/18 11:50:19 | 95
05/18/21 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 39-UWSIF-UW Glut 2- | 0.736 |  |  | 5557 | -43.8 |  | 5 | 57.2184 | 103.897 | MASS 29: |  | 102.319 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.1 | 2021/05/18 11:50:19 | 95
05/18/21 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 39-UWSIF-UW Glut 2- | 0.736 |  |  | 5553 | -44.034 |  | 6 | 56.7771 | 104.138 | MASS 29: |  | 102.557 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 7.5 | 2021/05/18 11:50:19 | 95
05/18/21 | 8 | 01-UWSIF-Liver-20210048.21 | 01-UWSIF-Liver- | 0.723 | 6622 | 0.095 |  |  |  | 1 | 15.9956 | 121.979 | MASS 30: |  |  | 121.102 | 0 |  | 0.72423 |  | 0.0036785 |  | 0.36651 | BGD 28: 43.2; BGD 29: 32.9; BGD 30: 928.7 | 2021/05/18 12:00:04 | 0
05/18/21 | 8 | 01-UWSIF-Liver-20210048.21 | 01-UWSIF-Liver- | 0.723 | 6625 | 0 |  |  |  | 2 | 15.9971 | 122.004 | MASS 30: |  |  | 121.127 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.2; BGD 29: 32.1; BGD 30: 1049.5 | 2021/05/18 12:00:04 | 0
05/18/21 | 8 | 01-UWSIF-Liver-20210048.21 | 01-UWSIF-Liver- | 0.723 | 3081 | 8.603 |  |  | Nitrogen | 3 | 11.8466 | 74.933 | MASS 30: |  |  | 74.389 | 0 |  | 0.73039 |  | 0.0037098 |  | 0.36961 | BGD 28: 43.1; BGD 29: 32.7; BGD 30: 1018.6 | 2021/05/18 12:00:04 | 0
05/18/21 | 8 | 01-UWSIF-Liver-20210048.21 | 01-UWSIF-Liver- | 0.723 |  |  | 775 | -19.377 | Carbon | 4 | 52.3943 | 21.062 | MASS 30: |  | 20.734 |  | 0 | 1.1633 |  | 0.010964 |  | 1.0845 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.0 | 2021/05/18 12:00:04 | 95
05/18/21 | 8 | 01-UWSIF-Liver-20210048.21 | 01-UWSIF-Liver- | 0.723 |  |  | 5554 | -43.8 |  | 5 | 58.3461 | 103.844 | MASS 30: |  | 102.267 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.3 | 2021/05/18 12:00:04 | 95
05/18/21 | 8 | 01-UWSIF-Liver-20210048.21 | 01-UWSIF-Liver- | 0.723 |  |  | 5547 | -44.013 |  | 6 | 57.8883 | 104.09 | MASS 30: |  | 102.51 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 7.7 | 2021/05/18 12:00:04 | 95
05/18/21 | 9 | 01-UWSIF-Liver-20210048.22 | 01-UWSIF-Liver- | 0.717 | 6624 | 0.095 |  |  |  | 1 | 16.1321 | 122.022 |  |  |  | 121.145 | 0 |  | 0.72432 |  | 0.0036785 |  | 0.36651 | BGD 28: 43.3; BGD 29: 33.0; BGD 30: 938.1 | 2021/05/18 12:09:49 | 0
05/18/21 | 9 | 01-UWSIF-Liver-20210048.22 | 01-UWSIF-Liver- | 0.717 | 6620 | 0 |  |  |  | 2 | 16.1341 | 122.054 |  |  |  | 121.177 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.3; BGD 29: 32.2; BGD 30: 1060.3 | 2021/05/18 12:09:49 | 0
05/18/21 | 9 | 01-UWSIF-Liver-20210048.22 | 01-UWSIF-Liver- | 0.717 | 3071 | 8.584 |  |  | Nitrogen | 3 | 11.8859 | 74.418 |  |  |  | 73.879 | 0 |  | 0.73047 |  | 0.0037098 |  | 0.36961 | BGD 28: 43.1; BGD 29: 32.7; BGD 30: 1029.4 | 2021/05/18 12:09:49 | 0
05/18/21 | 9 | 01-UWSIF-Liver-20210048.22 | 01-UWSIF-Liver- | 0.717 |  |  | 773 | -19.431 | Carbon | 4 | 52.5544 | 20.924 |  |  | 20.599 |  | 0 | 1.1632 |  | 0.010963 |  | 1.0844 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/18 12:09:49 | 95
05/18/21 | 9 | 01-UWSIF-Liver-20210048.22 | 01-UWSIF-Liver- | 0.717 |  |  | 5544 | -43.8 |  | 5 | 59.0475 | 103.73 |  |  | 102.155 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.3 | 2021/05/18 12:09:49 | 95
05/18/21 | 9 | 01-UWSIF-Liver-20210048.22 | 01-UWSIF-Liver- | 0.717 |  |  | 5540 | -44.028 |  | 6 | 58.6837 | 103.924 |  |  | 102.347 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 7.7 | 2021/05/18 12:09:49 | 95
05/18/21 | 10 | 20210048.069 | 643 | 0.806 | 6622 | 0.114 |  |  |  | 1 | 14.3421 | 121.864 |  |  |  | 120.987 | 0 |  | 0.72441 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.3; BGD 29: 33.1; BGD 30: 947.5 | 2021/05/18 12:19:34 | 0
05/18/21 | 10 | 20210048.069 | 643 | 0.806 | 6617 | 0 |  |  |  | 2 | 14.3505 | 122.017 |  |  |  | 121.139 | 1 |  | 0.72432 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.3; BGD 29: 32.3; BGD 30: 1069.4 | 2021/05/18 12:19:34 | 0
05/18/21 | 10 | 20210048.069 | 643 | 0.806 | 3894 | 12.548 |  |  | Nitrogen | 3 | 12.4552 | 94.551 |  |  |  | 93.862 | 0 |  | 0.73341 |  | 0.0037244 |  | 0.37105 | BGD 28: 43.2; BGD 29: 32.8; BGD 30: 1040.6 | 2021/05/18 12:19:34 | 0
05/18/21 | 10 | 20210048.069 | 643 | 0.806 |  |  | 844 | -31.649 | Carbon | 4 | 50.426 | 23.019 |  |  | 22.664 |  | 0 | 1.1498 |  | 0.010826 |  | 1.071 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.2 | 2021/05/18 12:19:34 | 95
05/18/21 | 10 | 20210048.069 | 643 | 0.806 |  |  | 5540 | -43.8 |  | 5 | 52.6966 | 103.627 |  |  | 102.054 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.5 | 2021/05/18 12:19:34 | 95
05/18/21 | 10 | 20210048.069 | 643 | 0.806 |  |  | 5545 | -43.998 |  | 6 | 52.3083 | 103.861 |  |  | 102.285 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 7.8 | 2021/05/18 12:19:34 | 95
05/18/21 | 11 | 20210048.070 | 645 | 0.814 | 6615 | 0.093 |  |  |  | 1 | 14.1953 | 121.757 |  |  |  | 120.882 | 0 |  | 0.72432 |  | 0.0036785 |  | 0.36651 | BGD 28: 43.3; BGD 29: 33.1; BGD 30: 958.3 | 2021/05/18 12:29:19 | 0
05/18/21 | 11 | 20210048.070 | 645 | 0.814 | 6619 | 0 |  |  |  | 2 | 14.2086 | 122.001 |  |  |  | 121.124 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.4; BGD 29: 32.3; BGD 30: 1080.4 | 2021/05/18 12:29:19 | 0
05/18/21 | 11 | 20210048.070 | 645 | 0.814 | 4036 | 11.909 |  |  | Nitrogen | 3 | 12.6053 | 97.932 |  |  |  | 97.22 | 0 |  | 0.73288 |  | 0.003722 |  | 0.37082 | BGD 28: 43.4; BGD 29: 33.0; BGD 30: 1051.4 | 2021/05/18 12:29:19 | 0
05/18/21 | 11 | 20210048.070 | 645 | 0.814 |  |  | 846 | -31.644 | Carbon | 4 | 50.0601 | 23.096 |  |  | 22.74 |  | 0 | 1.1498 |  | 0.010826 |  | 1.071 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.4 | 2021/05/18 12:29:19 | 95
05/18/21 | 11 | 20210048.070 | 645 | 0.814 |  |  | 5543 | -43.8 |  | 5 | 52.0759 | 103.69 |  |  | 102.116 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.6 | 2021/05/18 12:29:19 | 95
05/18/21 | 11 | 20210048.070 | 645 | 0.814 |  |  | 5536 | -44.03 |  | 6 | 51.6822 | 103.929 |  |  | 102.352 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.9 | 2021/05/18 12:29:19 | 95
05/18/21 | 12 | 20210048.070 | Replicate A11 | 0.84 | 6615 | 0.122 |  |  |  | 1 | 13.7674 | 121.975 |  |  |  | 121.098 | 0 |  | 0.72443 |  | 0.0036786 |  | 0.36652 | BGD 28: 43.4; BGD 29: 33.1; BGD 30: 969.1 | 2021/05/18 12:39:04 | 0
05/18/21 | 12 | 20210048.070 | Replicate A11 | 0.84 | 6608 | 0 |  |  |  | 2 | 13.7632 | 121.895 |  |  |  | 121.019 | 1 |  | 0.72434 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.3; BGD 29: 32.3; BGD 30: 1090.7 | 2021/05/18 12:39:04 | 0
05/18/21 | 12 | 20210048.070 | Replicate A11 | 0.84 | 4158 | 11.929 |  |  | Nitrogen | 3 | 12.4336 | 100.845 |  |  |  | 100.111 | 0 |  | 0.73298 |  | 0.0037221 |  | 0.37083 | BGD 28: 43.3; BGD 29: 32.9; BGD 30: 1060.9 | 2021/05/18 12:39:04 | 0
05/18/21 | 12 | 20210048.070 | Replicate A11 | 0.84 |  |  | 872 | -31.616 | Carbon | 4 | 49.708 | 23.834 |  |  | 23.466 |  | 0 | 1.1499 |  | 0.010827 |  | 1.0711 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.5 | 2021/05/18 12:39:04 | 95
05/18/21 | 12 | 20210048.070 | Replicate A11 | 0.84 |  |  | 5539 | -43.8 |  | 5 | 50.6104 | 103.598 |  |  | 102.025 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/18 12:39:04 | 95
05/18/21 | 12 | 20210048.070 | Replicate A11 | 0.84 |  |  | 5541 | -44.015 |  | 6 | 50.0991 | 103.919 |  |  | 102.341 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.0 | 2021/05/18 12:39:04 | 95
05/18/21 | 13 | 20210048.071 | 651 | 0.824 | 6601 | 0.093 |  |  |  | 1 | 14.0222 | 121.742 |  |  |  | 120.867 | 0 |  | 0.72423 |  | 0.0036785 |  | 0.36651 | BGD 28: 43.4; BGD 29: 33.2; BGD 30: 976.1 | 2021/05/18 12:48:49 | 0
05/18/21 | 13 | 20210048.071 | 651 | 0.824 | 6607 | 0 |  |  |  | 2 | 14.0197 | 121.695 |  |  |  | 120.821 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.3; BGD 29: 32.3; BGD 30: 1098.7 | 2021/05/18 12:48:49 | 0
05/18/21 | 13 | 20210048.071 | 651 | 0.824 | 4064 | 8.812 |  |  | Nitrogen | 3 | 12.5082 | 98.653 |  |  |  | 97.937 | 0 |  | 0.73054 |  | 0.0037106 |  | 0.36969 | BGD 28: 43.2; BGD 29: 32.9; BGD 30: 1069.6 | 2021/05/18 12:48:49 | 0
05/18/21 | 13 | 20210048.071 | 651 | 0.824 |  |  | 838 | -31.711 | Carbon | 4 | 49.0444 | 22.852 |  |  | 22.5 |  | 0 | 1.1498 |  | 0.010826 |  | 1.071 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.5 | 2021/05/18 12:48:49 | 95
05/18/21 | 13 | 20210048.071 | 651 | 0.824 |  |  | 5542 | -43.8 |  | 5 | 51.3837 | 103.727 |  |  | 102.152 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/18 12:48:49 | 95
05/18/21 | 13 | 20210048.071 | 651 | 0.824 |  |  | 5542 | -44.04 |  | 6 | 51.0786 | 103.915 |  |  | 102.337 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.1 | 2021/05/18 12:48:49 | 95
05/18/21 | 14 | 20210048.072 | 653 | 0.845 | 6603 | 0.103 |  |  |  | 1 | 13.6774 | 121.812 |  |  |  | 120.936 | 0 |  | 0.72424 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.4; BGD 29: 33.2; BGD 30: 983.7 | 2021/05/18 12:58:34 | 0
05/18/21 | 14 | 20210048.072 | 653 | 0.845 | 6622 | 0 |  |  |  | 2 | 13.6912 | 122.075 |  |  |  | 121.197 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.3; BGD 29: 32.3; BGD 30: 1104.6 | 2021/05/18 12:58:34 | 0
05/18/21 | 14 | 20210048.072 | 653 | 0.845 | 4071 | 8.639 |  |  | Nitrogen | 3 | 12.1698 | 98.287 |  |  |  | 97.574 | 0 |  | 0.73042 |  | 0.00371 |  | 0.36963 | BGD 28: 43.2; BGD 29: 32.8; BGD 30: 1076.4 | 2021/05/18 12:58:34 | 0
05/18/21 | 14 | 20210048.072 | 653 | 0.845 |  |  | 881 | -32.292 | Carbon | 4 | 49.6513 | 23.982 |  |  | 23.612 |  | 0 | 1.1492 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.6 | 2021/05/18 12:58:34 | 95
05/18/21 | 14 | 20210048.072 | 653 | 0.845 |  |  | 5552 | -43.8 |  | 5 | 50.017 | 103.784 |  |  | 102.208 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/18 12:58:34 | 95
05/18/21 | 14 | 20210048.072 | 653 | 0.845 |  |  | 5545 | -44.015 |  | 6 | 49.6533 | 104.013 |  |  | 102.434 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.2 | 2021/05/18 12:58:34 | 95
05/18/21 | 15 | 20210048.073 | 655 | 0.84 | 6614 | 0.108 |  |  |  | 1 | 13.7683 | 121.991 |  |  |  | 121.114 | 0 |  | 0.72425 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.5; BGD 29: 33.2; BGD 30: 988.5 | 2021/05/18 13:08:19 | 0
05/18/21 | 15 | 20210048.073 | 655 | 0.84 | 6627 | 0 |  |  |  | 2 | 13.7674 | 121.975 |  |  |  | 121.098 | 1 |  | 0.72417 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.4; BGD 29: 32.4; BGD 30: 1111.1 | 2021/05/18 13:08:19 | 0
05/18/21 | 15 | 20210048.073 | 655 | 0.84 | 4065 | 12.569 |  |  | Nitrogen | 3 | 12.221 | 98.009 |  |  |  | 97.296 | 0 |  | 0.73327 |  | 0.0037244 |  | 0.37106 | BGD 28: 43.3; BGD 29: 33.0; BGD 30: 1081.8 | 2021/05/18 13:08:19 | 0
05/18/21 | 15 | 20210048.073 | 655 | 0.84 |  |  | 888 | -31.397 | Carbon | 4 | 50.26 | 24.178 |  |  | 23.805 |  | 0 | 1.1502 |  | 0.010829 |  | 1.0713 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.6 | 2021/05/18 13:08:19 | 95
05/18/21 | 15 | 20210048.073 | 655 | 0.84 |  |  | 5545 | -43.8 |  | 5 | 50.405 | 103.727 |  |  | 102.152 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/18 13:08:19 | 95
05/18/21 | 15 | 20210048.073 | 655 | 0.84 |  |  | 5541 | -44.034 |  | 6 | 50.1479 | 103.889 |  |  | 102.311 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.3 | 2021/05/18 13:08:19 | 95
05/18/21 | 16 | 20210048.074 | 657 | 0.837 | 6621 | 0.115 |  |  |  | 1 | 13.812 | 121.885 |  |  |  | 121.009 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.5; BGD 29: 33.3; BGD 30: 993.4 | 2021/05/18 13:18:04 | 0
05/18/21 | 16 | 20210048.074 | 657 | 0.837 | 6618 | 0 |  |  |  | 2 | 13.819 | 122.017 |  |  |  | 121.14 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.4; BGD 29: 32.4; BGD 30: 1115.9 | 2021/05/18 13:18:04 | 0
05/18/21 | 16 | 20210048.074 | 657 | 0.837 | 4012 | 8.67 |  |  | Nitrogen | 3 | 12.199 | 97.155 |  |  |  | 96.45 | 0 |  | 0.73054 |  | 0.0037101 |  | 0.36964 | BGD 28: 43.4; BGD 29: 33.0; BGD 30: 1085.9 | 2021/05/18 13:18:04 | 0
05/18/21 | 16 | 20210048.074 | 657 | 0.837 |  |  | 873 | -31.821 | Carbon | 4 | 49.8764 | 23.828 |  |  | 23.46 |  | 0 | 1.1497 |  | 0.010824 |  | 1.0709 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.7 | 2021/05/18 13:18:04 | 95
05/18/21 | 16 | 20210048.074 | 657 | 0.837 |  |  | 5550 | -43.8 |  | 5 | 50.4579 | 103.807 |  |  | 102.231 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/18 13:18:04 | 95
05/18/21 | 16 | 20210048.074 | 657 | 0.837 |  |  | 5541 | -44.028 |  | 6 | 50.3471 | 103.876 |  |  | 102.299 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.3 | 2021/05/18 13:18:04 | 95
05/18/21 | 17 | 20210048.075 | 659 | 0.838 | 6611 | 0.086 |  |  |  | 1 | 13.7831 | 121.651 |  |  |  | 120.776 | 0 |  | 0.72424 |  | 0.0036785 |  | 0.3665 | BGD 28: 43.7; BGD 29: 33.4; BGD 30: 997.0 | 2021/05/18 13:27:50 | 0
05/18/21 | 17 | 20210048.075 | 659 | 0.838 | 6614 | 0 |  |  |  | 2 | 13.7967 | 121.907 |  |  |  | 121.031 | 1 |  | 0.72417 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.6; BGD 29: 32.6; BGD 30: 1119.1 | 2021/05/18 13:27:50 | 0
05/18/21 | 17 | 20210048.075 | 659 | 0.838 | 3833 | 11.905 |  |  | Nitrogen | 3 | 11.8123 | 92.484 |  |  |  | 91.811 | 0 |  | 0.73279 |  | 0.003722 |  | 0.37082 | BGD 28: 43.4; BGD 29: 33.1; BGD 30: 1090.1 | 2021/05/18 13:27:50 | 0
05/18/21 | 17 | 20210048.075 | 659 | 0.838 |  |  | 903 | -29.306 | Carbon | 4 | 51.2029 | 24.696 |  |  | 24.315 |  | 0 | 1.1524 |  | 0.010853 |  | 1.0736 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.6 | 2021/05/18 13:27:50 | 95
05/18/21 | 17 | 20210048.075 | 659 | 0.838 |  |  | 5552 | -43.8 |  | 5 | 50.3479 | 103.838 |  |  | 102.262 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/18 13:27:50 | 95
05/18/21 | 17 | 20210048.075 | 659 | 0.838 |  |  | 5550 | -43.999 |  | 6 | 49.8874 | 104.125 |  |  | 102.544 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.3 | 2021/05/18 13:27:50 | 95
05/18/21 | 18 | 20210048.076 | 661 | 0.846 | 6628 | 0.119 |  |  |  | 1 | 13.676 | 122.094 |  |  |  | 121.216 | 0 |  | 0.72425 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.5; BGD 29: 33.3; BGD 30: 998.8 | 2021/05/18 13:37:35 | 0
05/18/21 | 18 | 20210048.076 | 661 | 0.846 | 6627 | 0 |  |  |  | 2 | 13.6741 | 122.058 |  |  |  | 121.181 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.4; BGD 29: 32.4; BGD 30: 1119.8 | 2021/05/18 13:37:35 | 0
05/18/21 | 18 | 20210048.076 | 661 | 0.846 | 4009 | 12.152 |  |  | Nitrogen | 3 | 12.0544 | 96.962 |  |  |  | 96.256 | 0 |  | 0.73296 |  | 0.0037229 |  | 0.37091 | BGD 28: 43.3; BGD 29: 33.0; BGD 30: 1089.3 | 2021/05/18 13:37:35 | 0
05/18/21 | 18 | 20210048.076 | 661 | 0.846 |  |  | 862 | -30.974 | Carbon | 4 | 48.8429 | 23.514 |  |  | 23.151 |  | 0 | 1.1506 |  | 0.010834 |  | 1.0718 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.7 | 2021/05/18 13:37:35 | 95
05/18/21 | 18 | 20210048.076 | 661 | 0.846 |  |  | 5564 | -43.8 |  | 5 | 49.4818 | 104.084 |  |  | 102.503 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/18 13:37:35 | 95
05/18/21 | 18 | 20210048.076 | 661 | 0.846 |  |  | 5564 | -44.02 |  | 6 | 49.1066 | 104.318 |  |  | 102.735 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.3 | 2021/05/18 13:37:35 | 95
05/18/21 | 19 | 20210048.077 | 663 | 0.818 | 6637 | 0.112 |  |  |  | 1 | 14.1543 | 122.284 |  |  |  | 121.404 | 0 |  | 0.72426 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.4; BGD 29: 33.2; BGD 30: 998.8 | 2021/05/18 13:47:20 | 0
05/18/21 | 19 | 20210048.077 | 663 | 0.818 | 6641 | 0 |  |  |  | 2 | 14.1517 | 122.236 |  |  |  | 121.357 | 1 |  | 0.72417 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.2; BGD 29: 32.3; BGD 30: 1119.3 | 2021/05/18 13:47:20 | 0
05/18/21 | 19 | 20210048.077 | 663 | 0.818 | 4046 | 12.894 |  |  | Nitrogen | 3 | 12.5449 | 97.948 |  |  |  | 97.234 | 0 |  | 0.73351 |  | 0.0037256 |  | 0.37118 | BGD 28: 43.1; BGD 29: 32.8; BGD 30: 1089.6 | 2021/05/18 13:47:20 | 0
05/18/21 | 19 | 20210048.077 | 663 | 0.818 |  |  | 839 | -32.338 | Carbon | 4 | 49.4436 | 22.875 |  |  | 22.523 |  | 0 | 1.1491 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.8 | 2021/05/18 13:47:20 | 95
05/18/21 | 19 | 20210048.077 | 663 | 0.818 |  |  | 5570 | -43.8 |  | 5 | 50.9798 | 104.202 |  |  | 102.62 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/18 13:47:20 | 95
05/18/21 | 19 | 20210048.077 | 663 | 0.818 |  |  | 5573 | -44.013 |  | 6 | 50.626 | 104.416 |  |  | 102.831 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.2 | 2021/05/18 13:47:20 | 95
05/18/21 | 20 | 20210048.078 | 665 | 0.824 | 6652 | 0.123 |  |  |  | 1 | 14.0538 | 122.331 |  |  |  | 121.451 | 0 |  | 0.72425 |  | 0.0036787 |  | 0.36652 | BGD 28: 43.2; BGD 29: 33.0; BGD 30: 997.1 | 2021/05/18 13:57:05 | 0
05/18/21 | 20 | 20210048.078 | 665 | 0.824 | 6646 | 0 |  |  |  | 2 | 14.0565 | 122.382 |  |  |  | 121.502 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.2; BGD 30: 1117.8 | 2021/05/18 13:57:05 | 0
05/18/21 | 20 | 20210048.078 | 665 | 0.824 | 3053 | 12.812 |  |  | Nitrogen | 3 | 10.3084 | 74.083 |  |  |  | 73.544 | 0 |  | 0.73344 |  | 0.0037253 |  | 0.37115 | BGD 28: 43.0; BGD 29: 32.7; BGD 30: 1084.9 | 2021/05/18 13:57:05 | 0
05/18/21 | 20 | 20210048.078 | 665 | 0.824 |  |  | 878 | -34.457 | Carbon | 4 | 50.9657 | 24.012 |  |  | 23.642 |  | 0 | 1.1467 |  | 0.010795 |  | 1.068 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.6 | 2021/05/18 13:57:05 | 95
05/18/21 | 20 | 20210048.078 | 665 | 0.824 |  |  | 5582 | -43.8 |  | 5 | 50.3375 | 104.367 |  |  | 102.782 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/18 13:57:05 | 95
05/18/21 | 20 | 20210048.078 | 665 | 0.824 |  |  | 5584 | -44.022 |  | 6 | 49.8906 | 104.637 |  |  | 103.049 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.2 | 2021/05/18 13:57:05 | 95
05/18/21 | 21 | 20210048.079 | 667 | 0.802 | 6646 | 0.122 |  |  |  | 1 | 14.4415 | 122.37 |  |  |  | 121.49 | 0 |  | 0.72425 |  | 0.0036786 |  | 0.36652 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 991.1 | 2021/05/18 14:06:50 | 0
05/18/21 | 21 | 20210048.079 | 667 | 0.802 | 6649 | 0 |  |  |  | 2 | 14.4406 | 122.354 |  |  |  | 121.475 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.0; BGD 29: 32.1; BGD 30: 1112.4 | 2021/05/18 14:06:50 | 0
05/18/21 | 21 | 20210048.079 | 667 | 0.802 | 3890 | 10.027 |  |  | Nitrogen | 3 | 12.4776 | 94.077 |  |  |  | 93.394 | 0 |  | 0.73142 |  | 0.0037151 |  | 0.37013 | BGD 28: 42.9; BGD 29: 32.7; BGD 30: 1083.0 | 2021/05/18 14:06:50 | 0
05/18/21 | 21 | 20210048.079 | 667 | 0.802 |  |  | 819 | -32.76 | Carbon | 4 | 49.2419 | 22.192 |  |  | 21.85 |  | 0 | 1.1486 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.6 | 2021/05/18 14:06:50 | 95
05/18/21 | 21 | 20210048.079 | 667 | 0.802 |  |  | 5589 | -43.8 |  | 5 | 51.5578 | 104.462 |  |  | 102.876 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/18 14:06:50 | 95
05/18/21 | 21 | 20210048.079 | 667 | 0.802 |  |  | 5583 | -44.045 |  | 6 | 51.2273 | 104.656 |  |  | 103.067 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.2 | 2021/05/18 14:06:50 | 95
05/18/21 | 22 | 20210048.080 | 669 | 0.836 | 6648 | 0.098 |  |  |  | 1 | 13.8569 | 122.422 |  |  |  | 121.542 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 988.6 | 2021/05/18 14:16:35 | 0
05/18/21 | 22 | 20210048.080 | 669 | 0.836 | 6660 | 0 |  |  |  | 2 | 13.8595 | 122.473 |  |  |  | 121.592 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.0; BGD 30: 1110.2 | 2021/05/18 14:16:35 | 0
05/18/21 | 22 | 20210048.080 | 669 | 0.836 | 4143 | 12.39 |  |  | Nitrogen | 3 | 12.4306 | 100.005 |  |  |  | 99.277 | 0 |  | 0.73322 |  | 0.0037238 |  | 0.371 | BGD 28: 42.8; BGD 29: 32.6; BGD 30: 1080.7 | 2021/05/18 14:16:35 | 0
05/18/21 | 22 | 20210048.080 | 669 | 0.836 |  |  | 856 | -31.62 | Carbon | 4 | 48.9727 | 23.236 |  |  | 22.877 |  | 0 | 1.1498 |  | 0.010827 |  | 1.0711 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.6 | 2021/05/18 14:16:35 | 95
05/18/21 | 22 | 20210048.080 | 669 | 0.836 |  |  | 5589 | -43.8 |  | 5 | 49.4139 | 104.491 |  |  | 102.904 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/18 14:16:35 | 95
05/18/21 | 22 | 20210048.080 | 669 | 0.836 |  |  | 5589 | -44.011 |  | 6 | 48.8535 | 104.833 |  |  | 103.242 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.2 | 2021/05/18 14:16:35 | 95
05/18/21 | 23 | 20210048.081 | 671 | 0.828 | 6644 | 0.12 |  |  |  | 1 | 13.9871 | 122.353 |  |  |  | 121.473 | 0 |  | 0.72421 |  | 0.0036786 |  | 0.36652 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 988.4 | 2021/05/18 14:26:20 | 0
05/18/21 | 23 | 20210048.081 | 671 | 0.828 | 6653 | 0 |  |  |  | 2 | 13.9946 | 122.494 |  |  |  | 121.613 | 1 |  | 0.72413 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1109.5 | 2021/05/18 14:26:20 | 0
05/18/21 | 23 | 20210048.081 | 671 | 0.828 | 4104 | 9.875 |  |  | Nitrogen | 3 | 12.5 | 99.335 |  |  |  | 98.614 | 0 |  | 0.73128 |  | 0.0037145 |  | 0.37008 | BGD 28: 42.8; BGD 29: 32.7; BGD 30: 1081.4 | 2021/05/18 14:26:20 | 0
05/18/21 | 23 | 20210048.081 | 671 | 0.828 |  |  | 848 | -32.656 | Carbon | 4 | 49.036 | 22.989 |  |  | 22.634 |  | 0 | 1.1487 |  | 0.010815 |  | 1.0699 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.6 | 2021/05/18 14:26:20 | 95
05/18/21 | 23 | 20210048.081 | 671 | 0.828 |  |  | 5583 | -43.8 |  | 5 | 49.9898 | 104.431 |  |  | 102.845 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/18 14:26:20 | 95
05/18/21 | 23 | 20210048.081 | 671 | 0.828 |  |  | 5581 | -44.013 |  | 6 | 49.5638 | 104.689 |  |  | 103.1 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.1 | 2021/05/18 14:26:20 | 95
05/18/21 | 24 | 20210048.082 | 673 | 0.816 | 6655 | 0.114 |  |  |  | 1 | 14.2007 | 122.5 |  |  |  | 121.619 | 0 |  | 0.72419 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 988.6 | 2021/05/18 14:36:05 | 0
05/18/21 | 24 | 20210048.082 | 673 | 0.816 | 6657 | 0 |  |  |  | 2 | 14.196 | 122.413 |  |  |  | 121.533 | 1 |  | 0.7241 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1108.3 | 2021/05/18 14:36:05 | 0
05/18/21 | 24 | 20210048.082 | 673 | 0.816 | 3965 | 12.267 |  |  | Nitrogen | 3 | 12.4326 | 96.149 |  |  |  | 95.45 | 0 |  | 0.73299 |  | 0.0037233 |  | 0.37095 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1080.6 | 2021/05/18 14:36:05 | 0
05/18/21 | 24 | 20210048.082 | 673 | 0.816 |  |  | 828 | -32.448 | Carbon | 4 | 48.8977 | 22.484 |  |  | 22.137 |  | 0 | 1.1488 |  | 0.010817 |  | 1.0702 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.6 | 2021/05/18 14:36:05 | 95
05/18/21 | 24 | 20210048.082 | 673 | 0.816 |  |  | 5588 | -43.8 |  | 5 | 50.7468 | 104.418 |  |  | 102.832 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/18 14:36:05 | 95
05/18/21 | 24 | 20210048.082 | 673 | 0.816 |  |  | 5587 | -44.051 |  | 6 | 50.4472 | 104.597 |  |  | 103.009 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.1 | 2021/05/18 14:36:05 | 95
05/18/21 | 25 | 20210048.083 | 703 | 0.832 | 6654 | 0.111 |  |  |  | 1 | 13.9272 | 122.492 |  |  |  | 121.611 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 987.9 | 2021/05/18 14:45:50 | 0
05/18/21 | 25 | 20210048.083 | 703 | 0.832 | 6652 | 0 |  |  |  | 2 | 13.9273 | 122.494 |  |  |  | 121.613 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1108.6 | 2021/05/18 14:45:50 | 0
05/18/21 | 25 | 20210048.083 | 703 | 0.832 | 3511 | 10.044 |  |  | Nitrogen | 3 | 11.2207 | 84.658 |  |  |  | 84.043 | 0 |  | 0.73154 |  | 0.0037151 |  | 0.37014 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1078.7 | 2021/05/18 14:45:50 | 0
05/18/21 | 25 | 20210048.083 | 703 | 0.832 |  |  | 901 | -33.758 | Carbon | 4 | 51.3672 | 24.568 |  |  | 24.189 |  | 0 | 1.1472 |  | 0.010803 |  | 1.0687 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.5 | 2021/05/18 14:45:50 | 95
05/18/21 | 25 | 20210048.083 | 703 | 0.832 |  |  | 5588 | -43.8 |  | 5 | 49.8237 | 104.385 |  |  | 102.801 |  | 1 | 1.1349 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/18 14:45:50 | 95
05/18/21 | 25 | 20210048.083 | 703 | 0.832 |  |  | 5581 | -44.026 |  | 6 | 49.485 | 104.592 |  |  | 103.005 |  | 0 | 1.1347 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.2 | 2021/05/18 14:45:50 | 95
05/18/21 | 26 | 20210048.084 | 705 | 0.801 | 6645 | 0.108 |  |  |  | 1 | 14.4564 | 122.314 |  |  |  | 121.435 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 987.6 | 2021/05/18 14:55:35 | 0
05/18/21 | 26 | 20210048.084 | 705 | 0.801 | 6634 | 0 |  |  |  | 2 | 14.451 | 122.216 |  |  |  | 121.337 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1107.7 | 2021/05/18 14:55:35 | 0
05/18/21 | 26 | 20210048.084 | 705 | 0.801 | 3498 | 13.419 |  |  | Nitrogen | 3 | 11.6794 | 84.919 |  |  |  | 84.3 | 0 |  | 0.73396 |  | 0.0037276 |  | 0.37137 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1079.4 | 2021/05/18 14:55:35 | 0
05/18/21 | 26 | 20210048.084 | 705 | 0.801 |  |  | 856 | -32.737 | Carbon | 4 | 51.2535 | 23.318 |  |  | 22.959 |  | 0 | 1.1482 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.5 | 2021/05/18 14:55:35 | 95
05/18/21 | 26 | 20210048.084 | 705 | 0.801 |  |  | 5575 | -43.8 |  | 5 | 52.0535 | 104.207 |  |  | 102.625 |  | 1 | 1.1349 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/18 14:55:35 | 95
05/18/21 | 26 | 20210048.084 | 705 | 0.801 |  |  | 5571 | -44.016 |  | 6 | 51.6826 | 104.426 |  |  | 102.841 |  | 0 | 1.1346 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.2 | 2021/05/18 14:55:35 | 95
05/18/21 | 27 | 20210048.085 | 709 | 0.835 | 6621 | 0.102 |  |  |  | 1 | 13.855 | 122.071 |  |  |  | 121.194 | 0 |  | 0.72422 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 988.7 | 2021/05/18 15:05:20 | 0
05/18/21 | 27 | 20210048.085 | 709 | 0.835 | 6625 | 0 |  |  |  | 2 | 13.8512 | 122 |  |  |  | 121.123 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1109.2 | 2021/05/18 15:05:20 | 0
05/18/21 | 27 | 20210048.085 | 709 | 0.835 | 4087 | 8.927 |  |  | Nitrogen | 3 | 12.3571 | 98.832 |  |  |  | 98.115 | 0 |  | 0.73061 |  | 0.003711 |  | 0.36973 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1082.2 | 2021/05/18 15:05:20 | 0
05/18/21 | 27 | 20210048.085 | 709 | 0.835 |  |  | 858 | -32.089 | Carbon | 4 | 49.1758 | 23.324 |  |  | 22.964 |  | 0 | 1.1489 |  | 0.010821 |  | 1.0706 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.6 | 2021/05/18 15:05:20 | 95
05/18/21 | 27 | 20210048.085 | 709 | 0.835 |  |  | 5570 | -43.8 |  | 5 | 50.123 | 104.09 |  |  | 102.51 |  | 1 | 1.1349 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/18 15:05:20 | 95
05/18/21 | 27 | 20210048.085 | 709 | 0.835 |  |  | 5570 | -44.035 |  | 6 | 49.6674 | 104.371 |  |  | 102.787 |  | 0 | 1.1346 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.2 | 2021/05/18 15:05:20 | 95
05/18/21 | 28 | 20210048.086 | 713 | 0.829 | 6619 | 0.136 |  |  |  | 1 | 13.947 | 121.916 |  |  |  | 121.04 | 0 |  | 0.72431 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 993.7 | 2021/05/18 15:15:05 | 0
05/18/21 | 28 | 20210048.086 | 713 | 0.829 | 6628 | 0 |  |  |  | 2 | 13.9512 | 121.995 |  |  |  | 121.118 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.1; BGD 30: 1113.6 | 2021/05/18 15:15:05 | 0
05/18/21 | 28 | 20210048.086 | 713 | 0.829 | 3576 | 12.342 |  |  | Nitrogen | 3 | 11.4289 | 86.526 |  |  |  | 85.896 | 0 |  | 0.73315 |  | 0.0037236 |  | 0.37098 | BGD 28: 42.6; BGD 29: 32.6; BGD 30: 1084.2 | 2021/05/18 15:15:05 | 0
05/18/21 | 28 | 20210048.086 | 713 | 0.829 |  |  | 885 | -31.08 | Carbon | 4 | 50.8654 | 24.14 |  |  | 23.768 |  | 0 | 1.15 |  | 0.010833 |  | 1.0717 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.6 | 2021/05/18 15:15:05 | 95
05/18/21 | 28 | 20210048.086 | 713 | 0.829 |  |  | 5561 | -43.8 |  | 5 | 50.6804 | 103.971 |  |  | 102.392 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/18 15:15:05 | 95
05/18/21 | 28 | 20210048.086 | 713 | 0.829 |  |  | 5564 | -43.997 |  | 6 | 50.1836 | 104.276 |  |  | 102.693 |  | 0 | 1.1346 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.3 | 2021/05/18 15:15:05 | 95
05/18/21 | 29 | 20210048.087 | 715 | 0.81 | 6619 | 0.123 |  |  |  | 1 | 14.2707 | 121.854 |  |  |  | 120.978 | 0 |  | 0.72428 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 994.8 | 2021/05/18 15:24:50 | 0
05/18/21 | 29 | 20210048.087 | 715 | 0.81 | 6627 | 0 |  |  |  | 2 | 14.2799 | 122.022 |  |  |  | 121.145 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.1; BGD 30: 1116.0 | 2021/05/18 15:24:50 | 0
05/18/21 | 29 | 20210048.087 | 715 | 0.81 | 3136 | 10.601 |  |  | Nitrogen | 3 | 10.6252 | 75.43 |  |  |  | 74.882 | 0 |  | 0.73187 |  | 0.0037172 |  | 0.37034 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1086.7 | 2021/05/18 15:24:50 | 0
05/18/21 | 29 | 20210048.087 | 715 | 0.81 |  |  | 935 | -36.106 | Carbon | 4 | 54.2765 | 25.503 |  |  | 25.111 |  | 0 | 1.1444 |  | 0.010776 |  | 1.0662 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.6 | 2021/05/18 15:24:50 | 95
05/18/21 | 29 | 20210048.087 | 715 | 0.81 |  |  | 5562 | -43.8 |  | 5 | 51.8061 | 104.008 |  |  | 102.43 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/18 15:24:50 | 95
05/18/21 | 29 | 20210048.087 | 715 | 0.81 |  |  | 5556 | -44.034 |  | 6 | 51.4792 | 104.205 |  |  | 102.623 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.3 | 2021/05/18 15:24:50 | 95
05/18/21 | 30 | 20210048.088 | 719 | 0.829 | 6635 | 0.111 |  |  |  | 1 | 13.9523 | 122.016 |  |  |  | 121.139 | 0 |  | 0.72428 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 995.8 | 2021/05/18 15:34:35 | 0
05/18/21 | 30 | 20210048.088 | 719 | 0.829 | 6616 | 0 |  |  |  | 2 | 13.9473 | 121.922 |  |  |  | 121.045 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1117.7 | 2021/05/18 15:34:35 | 0
05/18/21 | 30 | 20210048.088 | 719 | 0.829 | 3940 | 12.448 |  |  | Nitrogen | 3 | 12.1186 | 94.661 |  |  |  | 93.972 | 0 |  | 0.73322 |  | 0.003724 |  | 0.37102 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1090.0 | 2021/05/18 15:34:35 | 0
05/18/21 | 30 | 20210048.088 | 719 | 0.829 |  |  | 859 | -32.646 | Carbon | 4 | 49.3983 | 23.243 |  |  | 22.885 |  | 0 | 1.1482 |  | 0.010815 |  | 1.0699 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/18 15:34:35 | 95
05/18/21 | 30 | 20210048.088 | 719 | 0.829 |  |  | 5565 | -43.8 |  | 5 | 50.6454 | 103.992 |  |  | 102.414 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/18 15:34:35 | 95
05/18/21 | 30 | 20210048.088 | 719 | 0.829 |  |  | 5554 | -44.029 |  | 6 | 50.3875 | 104.151 |  |  | 102.57 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.3 | 2021/05/18 15:34:35 | 95
05/18/21 | 31 | 20210048.089 | 721 | 0.814 | 6616 | 0.108 |  |  |  | 1 | 14.1947 | 121.747 |  |  |  | 120.872 | 0 |  | 0.72436 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1000.5 | 2021/05/18 15:44:20 | 0
05/18/21 | 31 | 20210048.089 | 721 | 0.814 | 6617 | 0 |  |  |  | 2 | 14.1984 | 121.814 |  |  |  | 120.938 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 32.0; BGD 30: 1122.1 | 2021/05/18 15:44:20 | 0
05/18/21 | 31 | 20210048.089 | 721 | 0.814 | 3828 | 12.965 |  |  | Nitrogen | 3 | 12.1696 | 92.591 |  |  |  | 91.917 | 0 |  | 0.73367 |  | 0.0037259 |  | 0.37121 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1094.8 | 2021/05/18 15:44:20 | 0
05/18/21 | 31 | 20210048.089 | 721 | 0.814 |  |  | 845 | -32.968 | Carbon | 4 | 49.8275 | 22.959 |  |  | 22.605 |  | 0 | 1.1479 |  | 0.010812 |  | 1.0696 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/18 15:44:20 | 95
05/18/21 | 31 | 20210048.089 | 721 | 0.814 |  |  | 5560 | -43.8 |  | 5 | 51.7251 | 103.903 |  |  | 102.326 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/18 15:44:20 | 95
05/18/21 | 31 | 20210048.089 | 721 | 0.814 |  |  | 5560 | -44.021 |  | 6 | 51.1689 | 104.239 |  |  | 102.657 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.3 | 2021/05/18 15:44:20 | 95
05/18/21 | 32 | 20210048.090 | 723 | 0.825 | 6607 | 0.089 |  |  |  | 1 | 14.0066 | 121.769 |  |  |  | 120.893 | 0 |  | 0.72425 |  | 0.0036785 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.9; BGD 30: 1003.5 | 2021/05/18 15:54:05 | 0
05/18/21 | 32 | 20210048.090 | 723 | 0.825 | 6607 | 0 |  |  |  | 2 | 14.0002 | 121.649 |  |  |  | 120.775 | 1 |  | 0.72418 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 32.0; BGD 30: 1125.3 | 2021/05/18 15:54:05 | 0
05/18/21 | 32 | 20210048.090 | 723 | 0.825 | 4061 | 9.013 |  |  | Nitrogen | 3 | 12.4181 | 97.687 |  |  |  | 96.978 | 0 |  | 0.73071 |  | 0.0037114 |  | 0.36976 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1097.6 | 2021/05/18 15:54:05 | 0
05/18/21 | 32 | 20210048.090 | 723 | 0.825 |  |  | 841 | -32.878 | Carbon | 4 | 48.8639 | 22.78 |  |  | 22.429 |  | 0 | 1.1479 |  | 0.010813 |  | 1.0697 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/18 15:54:05 | 95
05/18/21 | 32 | 20210048.090 | 723 | 0.825 |  |  | 5553 | -43.8 |  | 5 | 51.1551 | 103.83 |  |  | 102.254 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.0 | 2021/05/18 15:54:05 | 95
05/18/21 | 32 | 20210048.090 | 723 | 0.825 |  |  | 5553 | -44.014 |  | 6 | 50.7666 | 104.068 |  |  | 102.489 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/18 15:54:05 | 95
05/18/21 | 33 | 20210048.090 | Replicate C8 | 0.81 | 6604 | 0.11 |  |  |  | 1 | 14.2551 | 121.569 |  |  |  | 120.695 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 1008.0 | 2021/05/18 16:03:51 | 0
05/18/21 | 33 | 20210048.090 | Replicate C8 | 0.81 | 6606 | 0 |  |  |  | 2 | 14.2561 | 121.588 |  |  |  | 120.713 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1129.6 | 2021/05/18 16:03:51 | 0
05/18/21 | 33 | 20210048.090 | Replicate C8 | 0.81 | 4024 | 9.098 |  |  | Nitrogen | 3 | 12.6013 | 97.101 |  |  |  | 96.396 | 0 |  | 0.73084 |  | 0.0037117 |  | 0.36979 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1102.3 | 2021/05/18 16:03:51 | 0
05/18/21 | 33 | 20210048.090 | Replicate C8 | 0.81 |  |  | 831 | -32.931 | Carbon | 4 | 49.2824 | 22.497 |  |  | 22.15 |  | 0 | 1.1479 |  | 0.010812 |  | 1.0696 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/18 16:03:51 | 95
05/18/21 | 33 | 20210048.090 | Replicate C8 | 0.81 |  |  | 5547 | -43.8 |  | 5 | 52.3447 | 103.683 |  |  | 102.109 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.0 | 2021/05/18 16:03:51 | 95
05/18/21 | 33 | 20210048.090 | Replicate C8 | 0.81 |  |  | 5545 | -44.012 |  | 6 | 51.9089 | 103.947 |  |  | 102.369 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.5 | 2021/05/18 16:03:51 | 95
05/18/21 | 34 | 36-UWSIF-UT Glut 1-20210048.26 | 36-UWSIF-UT Glut 1- | 0.731 | 6595 | 0.129 |  |  |  | 1 | 15.7852 | 121.399 |  |  |  | 120.526 | 0 |  | 0.72424 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 1011.3 | 2021/05/18 16:13:36 | 0
05/18/21 | 34 | 36-UWSIF-UT Glut 1-20210048.26 | 36-UWSIF-UT Glut 1- | 0.731 | 6598 | 0 |  |  |  | 2 | 15.796 | 121.576 |  |  |  | 120.701 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1133.5 | 2021/05/18 16:13:36 | 0
05/18/21 | 34 | 36-UWSIF-UT Glut 1-20210048.26 | 36-UWSIF-UT Glut 1- | 0.731 | 2255 | -2.832 |  |  | Nitrogen | 3 | 9.1972 | 54.821 |  |  |  | 54.428 | 0 |  | 0.7221 |  | 0.0036678 |  | 0.36544 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1101.6 | 2021/05/18 16:13:36 | 0
05/18/21 | 34 | 36-UWSIF-UT Glut 1-20210048.26 | 36-UWSIF-UT Glut 1- | 0.731 |  |  | 627 | -29.752 | Carbon | 4 | 42.8828 | 16.754 |  |  | 16.495 |  | 0 | 1.1515 |  | 0.010848 |  | 1.0731 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/18 16:13:36 | 95
05/18/21 | 34 | 36-UWSIF-UT Glut 1-20210048.26 | 36-UWSIF-UT Glut 1- | 0.731 |  |  | 5539 | -43.8 |  | 5 | 58.1836 | 103.583 |  |  | 102.011 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.9 | 2021/05/18 16:13:36 | 95
05/18/21 | 34 | 36-UWSIF-UT Glut 1-20210048.26 | 36-UWSIF-UT Glut 1- | 0.731 |  |  | 5535 | -44.032 |  | 6 | 57.694 | 103.851 |  |  | 102.275 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.3 | 2021/05/18 16:13:36 | 95
05/18/21 | 35 | 36-UWSIF-UT Glut 1-20210048.27 | 36-UWSIF-UT Glut 1- | 0.829 | 6588 | 0.122 |  |  |  | 1 | 13.9147 | 121.316 |  |  |  | 120.444 | 0 |  | 0.72425 |  | 0.0036786 |  | 0.36652 | BGD 28: 42.6; BGD 29: 32.6; BGD 30: 1007.5 | 2021/05/18 16:23:21 | 0
05/18/21 | 35 | 36-UWSIF-UT Glut 1-20210048.27 | 36-UWSIF-UT Glut 1- | 0.829 | 6581 | 0 |  |  |  | 2 | 13.9152 | 121.326 |  |  |  | 120.454 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.6; BGD 29: 31.9; BGD 30: 1130.2 | 2021/05/18 16:23:21 | 0
05/18/21 | 35 | 36-UWSIF-UT Glut 1-20210048.27 | 36-UWSIF-UT Glut 1- | 0.829 | 2572 | -2.801 |  |  | Nitrogen | 3 | 8.9616 | 62.107 |  |  |  | 61.662 | 0 |  | 0.72213 |  | 0.0036679 |  | 0.36545 | BGD 28: 42.5; BGD 29: 32.4; BGD 30: 1099.8 | 2021/05/18 16:23:21 | 0
05/18/21 | 35 | 36-UWSIF-UT Glut 1-20210048.27 | 36-UWSIF-UT Glut 1- | 0.829 |  |  | 714 | -29.84 | Carbon | 4 | 42.255 | 19.131 |  |  | 18.835 |  | 0 | 1.1513 |  | 0.010847 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.7 | 2021/05/18 16:23:21 | 95
05/18/21 | 35 | 36-UWSIF-UT Glut 1-20210048.27 | 36-UWSIF-UT Glut 1- | 0.829 |  |  | 5535 | -43.8 |  | 5 | 51.4664 | 103.483 |  |  | 101.912 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.0 | 2021/05/18 16:23:21 | 95
05/18/21 | 35 | 36-UWSIF-UT Glut 1-20210048.27 | 36-UWSIF-UT Glut 1- | 0.829 |  |  | 5529 | -44.019 |  | 6 | 51.1612 | 103.673 |  |  | 102.1 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/18 16:23:21 | 95
05/18/21 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 39-UWSIF-UW Glut 2- | 0.769 | 6573 | 0.113 |  |  |  | 1 | 14.9862 | 121.073 |  |  |  | 120.203 | 0 |  | 0.72424 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.7; BGD 29: 32.7; BGD 30: 1009.7 | 2021/05/18 16:33:06 | 0
05/18/21 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 39-UWSIF-UW Glut 2- | 0.769 | 6579 | 0 |  |  |  | 2 | 14.9957 | 121.237 |  |  |  | 120.365 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.6; BGD 29: 31.9; BGD 30: 1133.1 | 2021/05/18 16:33:06 | 0
05/18/21 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 39-UWSIF-UW Glut 2- | 0.769 | 2594 | 29.79 |  |  | Nitrogen | 3 | 9.7486 | 62.828 |  |  |  | 62.363 | 0 |  | 0.74573 |  | 0.0037878 |  | 0.37735 | BGD 28: 42.4; BGD 29: 32.3; BGD 30: 1103.4 | 2021/05/18 16:33:06 | 0
05/18/21 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 39-UWSIF-UW Glut 2- | 0.769 |  |  | 720 | 22.208 | Carbon | 4 | 46.0377 | 19.379 |  |  | 19.069 |  | 0 | 1.2088 |  | 0.011428 |  | 1.1299 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/18 16:33:06 | 95
05/18/21 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 39-UWSIF-UW Glut 2- | 0.769 |  |  | 5526 | -43.8 |  | 5 | 55.8046 | 103.296 |  |  | 101.728 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.0 | 2021/05/18 16:33:06 | 95
05/18/21 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 39-UWSIF-UW Glut 2- | 0.769 |  |  | 5524 | -44.004 |  | 6 | 55.2268 | 103.63 |  |  | 102.058 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.5 | 2021/05/18 16:33:06 | 95
05/18/21 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 39-UWSIF-UW Glut 2- | 0.841 | 6570 | 0.103 |  |  |  | 1 | 13.7005 | 121.023 |  |  |  | 120.153 | 0 |  | 0.72415 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.7; BGD 29: 32.7; BGD 30: 1012.4 | 2021/05/18 16:42:51 | 0
05/18/21 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 39-UWSIF-UW Glut 2- | 0.841 | 6568 | 0 |  |  |  | 2 | 13.705 | 121.108 |  |  |  | 120.237 | 1 |  | 0.72408 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.6; BGD 29: 31.9; BGD 30: 1136.0 | 2021/05/18 16:42:51 | 0
05/18/21 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 39-UWSIF-UW Glut 2- | 0.841 | 2808 | 29.667 |  |  | Nitrogen | 3 | 9.4809 | 68.061 |  |  |  | 67.557 | 0 |  | 0.74556 |  | 0.0037873 |  | 0.3773 | BGD 28: 42.5; BGD 29: 32.5; BGD 30: 1106.5 | 2021/05/18 16:42:51 | 0
05/18/21 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 39-UWSIF-UW Glut 2- | 0.841 |  |  | 779 | 22.32 | Carbon | 4 | 45.1159 | 21.105 |  |  | 20.767 |  | 0 | 1.2089 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/18 16:42:51 | 95
05/18/21 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 39-UWSIF-UW Glut 2- | 0.841 |  |  | 5523 | -43.8 |  | 5 | 51.1346 | 103.227 |  |  | 101.66 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/18 16:42:51 | 95
05/18/21 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 39-UWSIF-UW Glut 2- | 0.841 |  |  | 5515 | -44.024 |  | 6 | 50.7472 | 103.473 |  |  | 101.903 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.7 | 2021/05/18 16:42:51 | 95
05/18/21 | 38 | 01-UWSIF-Liver-20210048.23 | 01-UWSIF-Liver- | 0.736 | 6561 | 0.14 |  |  |  | 1 | 15.639 | 120.761 |  |  |  | 119.893 | 0 |  | 0.72425 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.7; BGD 29: 32.8; BGD 30: 1017.0 | 2021/05/18 16:52:36 | 0
05/18/21 | 38 | 01-UWSIF-Liver-20210048.23 | 01-UWSIF-Liver- | 0.736 | 6576 | 0 |  |  |  | 2 | 15.6557 | 121.033 |  |  |  | 120.163 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 32.0; BGD 30: 1140.6 | 2021/05/18 16:52:36 | 0
05/18/21 | 38 | 01-UWSIF-Liver-20210048.23 | 01-UWSIF-Liver- | 0.736 | 3138 | 8.547 |  |  | Nitrogen | 3 | 11.721 | 75.675 |  |  |  | 75.126 | 0 |  | 0.73034 |  | 0.0037096 |  | 0.36959 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1110.2 | 2021/05/18 16:52:36 | 0
05/18/21 | 38 | 01-UWSIF-Liver-20210048.23 | 01-UWSIF-Liver- | 0.736 |  |  | 785 | -19.604 | Carbon | 4 | 51.7494 | 21.205 |  |  | 20.875 |  | 0 | 1.1627 |  | 0.010961 |  | 1.0842 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/18 16:52:36 | 95
05/18/21 | 38 | 01-UWSIF-Liver-20210048.23 | 01-UWSIF-Liver- | 0.736 |  |  | 5521 | -43.8 |  | 5 | 58.4273 | 103.228 |  |  | 101.662 |  | 1 | 1.1349 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/18 16:52:36 | 95
05/18/21 | 38 | 01-UWSIF-Liver-20210048.23 | 01-UWSIF-Liver- | 0.736 |  |  | 5517 | -43.999 |  | 6 | 58.0649 | 103.43 |  |  | 101.86 |  | 0 | 1.1346 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/18 16:52:36 | 95
05/18/21 | 39 | 01-UWSIF-Liver-20210048.24 | 01-UWSIF-Liver- | 0.731 | 6560 | 0.109 |  |  |  | 1 | 15.7518 | 120.855 |  |  |  | 119.986 | 0 |  | 0.72426 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 1023.7 | 2021/05/18 17:02:21 | 0
05/18/21 | 39 | 01-UWSIF-Liver-20210048.24 | 01-UWSIF-Liver- | 0.731 | 6571 | 0 |  |  |  | 2 | 15.7574 | 120.947 |  |  |  | 120.077 | 1 |  | 0.72418 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1147.4 | 2021/05/18 17:02:21 | 0
05/18/21 | 39 | 01-UWSIF-Liver-20210048.24 | 01-UWSIF-Liver- | 0.731 | 3123 | 8.548 |  |  | Nitrogen | 3 | 11.7612 | 75.322 |  |  |  | 74.776 | 0 |  | 0.73037 |  | 0.0037096 |  | 0.36959 | BGD 28: 42.7; BGD 29: 32.5; BGD 30: 1115.8 | 2021/05/18 17:02:21 | 0
05/18/21 | 39 | 01-UWSIF-Liver-20210048.24 | 01-UWSIF-Liver- | 0.731 |  |  | 781 | -19.591 | Carbon | 4 | 51.8187 | 21.061 |  |  | 20.733 |  | 0 | 1.1628 |  | 0.010961 |  | 1.0842 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.2 | 2021/05/18 17:02:21 | 95
05/18/21 | 39 | 01-UWSIF-Liver-20210048.24 | 01-UWSIF-Liver- | 0.731 |  |  | 5512 | -43.8 |  | 5 | 59.1295 | 103.061 |  |  | 101.496 |  | 1 | 1.1349 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/18 17:02:21 | 95
05/18/21 | 39 | 01-UWSIF-Liver-20210048.24 | 01-UWSIF-Liver- | 0.731 |  |  | 5508 | -44.023 |  | 6 | 58.6915 | 103.303 |  |  | 101.736 |  | 0 | 1.1347 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.7 | 2021/05/18 17:02:21 | 95
05/18/21 | 40 | 20210048.091 | 725 | 0.827 | 6550 | 0.106 |  |  |  | 1 | 13.9121 | 120.652 |  |  |  | 119.785 | 0 |  | 0.72416 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1027.5 | 2021/05/18 17:12:06 | 0
05/18/21 | 40 | 20210048.091 | 725 | 0.827 | 6552 | 0 |  |  |  | 2 | 13.9194 | 120.785 |  |  |  | 119.916 | 1 |  | 0.72408 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1151.7 | 2021/05/18 17:12:06 | 0
05/18/21 | 40 | 20210048.091 | 725 | 0.827 | 3783 | 12.448 |  |  | Nitrogen | 3 | 11.8627 | 91.21 |  |  |  | 90.546 | 0 |  | 0.73309 |  | 0.003724 |  | 0.37102 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1123.8 | 2021/05/18 17:12:06 | 0
05/18/21 | 40 | 20210048.091 | 725 | 0.827 |  |  | 881 | -33.413 | Carbon | 4 | 50.7121 | 23.97 |  |  | 23.601 |  | 0 | 1.1476 |  | 0.010807 |  | 1.0691 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.4 | 2021/05/18 17:12:06 | 95
05/18/21 | 40 | 20210048.091 | 725 | 0.827 |  |  | 5508 | -43.8 |  | 5 | 52.2728 | 103.056 |  |  | 101.492 |  | 1 | 1.135 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.4 | 2021/05/18 17:12:06 | 95
05/18/21 | 40 | 20210048.091 | 725 | 0.827 |  |  | 5504 | -44.01 |  | 6 | 51.9944 | 103.231 |  |  | 101.664 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.8 | 2021/05/18 17:12:06 | 95
05/18/21 | 41 | 20210048.092 | 727 | 0.807 | 6557 | 0.107 |  |  |  | 1 | 14.2759 | 120.99 |  |  |  | 120.12 | 0 |  | 0.72427 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.0; BGD 29: 33.0; BGD 30: 1036.1 | 2021/05/18 17:21:52 | 0
05/18/21 | 41 | 20210048.092 | 727 | 0.807 | 6565 | 0 |  |  |  | 2 | 14.2735 | 120.947 |  |  |  | 120.078 | 1 |  | 0.72419 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.0; BGD 29: 32.2; BGD 30: 1160.9 | 2021/05/18 17:21:52 | 0
05/18/21 | 41 | 20210048.092 | 727 | 0.807 | 3914 | 10.218 |  |  | Nitrogen | 3 | 12.3737 | 93.758 |  |  |  | 93.077 | 0 |  | 0.73159 |  | 0.0037158 |  | 0.3702 | BGD 28: 42.8; BGD 29: 32.7; BGD 30: 1132.4 | 2021/05/18 17:21:52 | 0
05/18/21 | 41 | 20210048.092 | 727 | 0.807 |  |  | 837 | -32.867 | Carbon | 4 | 49.7025 | 22.634 |  |  | 22.285 |  | 0 | 1.1482 |  | 0.010813 |  | 1.0697 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.5 | 2021/05/18 17:21:52 | 95
05/18/21 | 41 | 20210048.092 | 727 | 0.807 |  |  | 5508 | -43.8 |  | 5 | 53.6529 | 103.004 |  |  | 101.44 |  | 1 | 1.135 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.5 | 2021/05/18 17:21:52 | 95
05/18/21 | 41 | 20210048.092 | 727 | 0.807 |  |  | 5504 | -44.002 |  | 6 | 53.3026 | 103.219 |  |  | 101.652 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.9 | 2021/05/18 17:21:52 | 95
05/18/21 | 42 | 20210048.093 | 729 | 0.85 | 6556 | 0.117 |  |  |  | 1 | 13.5387 | 120.708 |  |  |  | 119.84 | 0 |  | 0.72428 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.0; BGD 29: 33.0; BGD 30: 1042.1 | 2021/05/18 17:31:37 | 0
05/18/21 | 42 | 20210048.093 | 729 | 0.85 | 6561 | 0 |  |  |  | 2 | 13.5479 | 120.881 |  |  |  | 120.012 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1166.5 | 2021/05/18 17:31:37 | 0
05/18/21 | 42 | 20210048.093 | 729 | 0.85 | 4022 | 13.663 |  |  | Nitrogen | 3 | 11.9563 | 96.422 |  |  |  | 95.719 | 0 |  | 0.73409 |  | 0.0037285 |  | 0.37146 | BGD 28: 42.8; BGD 29: 32.6; BGD 30: 1138.2 | 2021/05/18 17:31:37 | 0
05/18/21 | 42 | 20210048.093 | 729 | 0.85 |  |  | 894 | -32.187 | Carbon | 4 | 49.7634 | 24.238 |  |  | 23.864 |  | 0 | 1.149 |  | 0.01082 |  | 1.0705 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/18 17:31:37 | 95
05/18/21 | 42 | 20210048.093 | 729 | 0.85 |  |  | 5502 | -43.8 |  | 5 | 51.0394 | 102.939 |  |  | 101.376 |  | 1 | 1.135 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.6 | 2021/05/18 17:31:37 | 95
05/18/21 | 42 | 20210048.093 | 729 | 0.85 |  |  | 5503 | -44.04 |  | 6 | 50.6897 | 103.165 |  |  | 101.599 |  | 0 | 1.1347 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.0 | 2021/05/18 17:31:37 | 95
05/18/21 | 43 | 20210048.094 | 731 | 0.823 | 6551 | 0.094 |  |  |  | 1 | 13.9787 | 120.633 |  |  |  | 119.765 | 0 |  | 0.72419 |  | 0.0036785 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1047.4 | 2021/05/18 17:41:22 | 0
05/18/21 | 43 | 20210048.094 | 731 | 0.823 | 6557 | 0 |  |  |  | 2 | 13.9814 | 120.681 |  |  |  | 119.814 | 1 |  | 0.72412 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1172.6 | 2021/05/18 17:41:22 | 0
05/18/21 | 43 | 20210048.094 | 731 | 0.823 | 3563 | 12.792 |  |  | Nitrogen | 3 | 11.411 | 85.401 |  |  |  | 84.779 | 0 |  | 0.73338 |  | 0.0037253 |  | 0.37114 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 1141.2 | 2021/05/18 17:41:22 | 0
05/18/21 | 43 | 20210048.094 | 731 | 0.823 |  |  | 908 | -33.942 | Carbon | 4 | 52.1044 | 24.677 |  |  | 24.296 |  | 0 | 1.1469 |  | 0.010801 |  | 1.0685 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.5 | 2021/05/18 17:41:22 | 95
05/18/21 | 43 | 20210048.094 | 731 | 0.823 |  |  | 5503 | -43.8 |  | 5 | 52.6083 | 103.005 |  |  | 101.441 |  | 1 | 1.1349 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 7.7 | 2021/05/18 17:41:22 | 95
05/18/21 | 43 | 20210048.094 | 731 | 0.823 |  |  | 5499 | -44.013 |  | 6 | 52.4176 | 103.124 |  |  | 101.559 |  | 0 | 1.1346 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.0 | 2021/05/18 17:41:22 | 95
05/18/21 | 44 | 20210048.095 | 733 | 0.837 | 6551 | 0.123 |  |  |  | 1 | 13.7442 | 120.62 |  |  |  | 119.752 | 0 |  | 0.72438 |  | 0.0036787 |  | 0.36652 | BGD 28: 43.2; BGD 29: 33.1; BGD 30: 1051.5 | 2021/05/18 17:51:07 | 0
05/18/21 | 44 | 20210048.095 | 733 | 0.837 | 6557 | 0 |  |  |  | 2 | 13.7488 | 120.704 |  |  |  | 119.836 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1176.5 | 2021/05/18 17:51:07 | 0
05/18/21 | 44 | 20210048.095 | 733 | 0.837 | 4184 | 12.792 |  |  | Nitrogen | 3 | 12.4318 | 100.22 |  |  |  | 99.49 | 0 |  | 0.73355 |  | 0.0037253 |  | 0.37114 | BGD 28: 43.0; BGD 29: 32.9; BGD 30: 1147.4 | 2021/05/18 17:51:07 | 0
05/18/21 | 44 | 20210048.095 | 733 | 0.837 |  |  | 854 | -31.889 | Carbon | 4 | 48.7209 | 23.118 |  |  | 22.761 |  | 0 | 1.1493 |  | 0.010824 |  | 1.0708 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.7 | 2021/05/18 17:51:07 | 95
05/18/21 | 44 | 20210048.095 | 733 | 0.837 |  |  | 5501 | -43.8 |  | 5 | 51.9547 | 102.86 |  |  | 101.299 |  | 1 | 1.135 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.7 | 2021/05/18 17:51:07 | 95
05/18/21 | 44 | 20210048.095 | 733 | 0.837 |  |  | 5497 | -44.011 |  | 6 | 51.4628 | 103.174 |  |  | 101.608 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.1 | 2021/05/18 17:51:07 | 95
05/18/21 | 45 | 20210048.096 | 735 | 0.85 | 6555 | 0.115 |  |  |  | 1 | 13.5338 | 120.616 |  |  |  | 119.749 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.1; BGD 30: 1055.9 | 2021/05/18 18:00:52 | 0
05/18/21 | 45 | 20210048.096 | 735 | 0.85 | 6537 | 0 |  |  |  | 2 | 13.5323 | 120.588 |  |  |  | 119.721 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1181.5 | 2021/05/18 18:00:52 | 0
05/18/21 | 45 | 20210048.096 | 735 | 0.85 | 3998 | 9.97 |  |  | Nitrogen | 3 | 11.9007 | 95.703 |  |  |  | 95.008 | 0 |  | 0.73144 |  | 0.0037149 |  | 0.37011 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1152.2 | 2021/05/18 18:00:52 | 0
05/18/21 | 45 | 20210048.096 | 735 | 0.85 |  |  | 893 | -33.027 | Carbon | 4 | 49.7519 | 24.231 |  |  | 23.857 |  | 0 | 1.1481 |  | 0.010811 |  | 1.0695 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.7 | 2021/05/18 18:00:52 | 95
05/18/21 | 45 | 20210048.096 | 735 | 0.85 |  |  | 5497 | -43.8 |  | 5 | 51.1962 | 102.837 |  |  | 101.276 |  | 1 | 1.135 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 7.8 | 2021/05/18 18:00:52 | 95
05/18/21 | 45 | 20210048.096 | 735 | 0.85 |  |  | 5489 | -44.012 |  | 6 | 50.8632 | 103.053 |  |  | 101.489 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.2 | 2021/05/18 18:00:52 | 95
05/18/21 | 46 | 20210048.097 | 739 | 0.826 | 6549 | 0.107 |  |  |  | 1 | 13.9311 | 120.69 |  |  |  | 119.822 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.1; BGD 30: 1059.8 | 2021/05/18 18:10:37 | 0
05/18/21 | 46 | 20210048.097 | 739 | 0.826 | 6559 | 0 |  |  |  | 2 | 13.9356 | 120.773 |  |  |  | 119.904 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1185.5 | 2021/05/18 18:10:37 | 0
05/18/21 | 46 | 20210048.097 | 739 | 0.826 | 3972 | 11.513 |  |  | Nitrogen | 3 | 12.1996 | 95.119 |  |  |  | 94.427 | 0 |  | 0.73256 |  | 0.0037205 |  | 0.37068 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1156.3 | 2021/05/18 18:10:37 | 0
05/18/21 | 46 | 20210048.097 | 739 | 0.826 |  |  | 835 | -32.072 | Carbon | 4 | 48.4429 | 22.565 |  |  | 22.217 |  | 0 | 1.1492 |  | 0.010822 |  | 1.0706 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.8 | 2021/05/18 18:10:37 | 95
05/18/21 | 46 | 20210048.097 | 739 | 0.826 |  |  | 5493 | -43.8 |  | 5 | 52.7604 | 102.788 |  |  | 101.228 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.8 | 2021/05/18 18:10:37 | 95
05/18/21 | 46 | 20210048.097 | 739 | 0.826 |  |  | 5488 | -44.035 |  | 6 | 52.3965 | 103.018 |  |  | 101.454 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.3 | 2021/05/18 18:10:37 | 95
05/18/21 | 47 | 20210048.098 | 741 | 0.825 | 6532 | 0.118 |  |  |  | 1 | 13.9345 | 120.446 |  |  |  | 119.579 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.2; BGD 29: 33.0; BGD 30: 1062.8 | 2021/05/18 18:20:23 | 0
05/18/21 | 47 | 20210048.098 | 741 | 0.825 | 6552 | 0 |  |  |  | 2 | 13.9421 | 120.583 |  |  |  | 119.716 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1188.7 | 2021/05/18 18:20:23 | 0
05/18/21 | 47 | 20210048.098 | 741 | 0.825 | 3878 | 13.087 |  |  | Nitrogen | 3 | 12.0721 | 93.373 |  |  |  | 92.693 | 0 |  | 0.73372 |  | 0.0037263 |  | 0.37125 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1159.1 | 2021/05/18 18:20:23 | 0
05/18/21 | 47 | 20210048.098 | 741 | 0.825 |  |  | 851 | -31.878 | Carbon | 4 | 49.3832 | 23.09 |  |  | 22.734 |  | 0 | 1.1494 |  | 0.010824 |  | 1.0708 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.8 | 2021/05/18 18:20:23 | 95
05/18/21 | 47 | 20210048.098 | 741 | 0.825 |  |  | 5497 | -43.8 |  | 5 | 52.8403 | 102.778 |  |  | 101.218 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.9 | 2021/05/18 18:20:23 | 95
05/18/21 | 47 | 20210048.098 | 741 | 0.825 |  |  | 5487 | -44.005 |  | 6 | 52.5593 | 102.956 |  |  | 101.393 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.3 | 2021/05/18 18:20:23 | 95
05/18/21 | 48 | 20210048.099 | 747 | 0.849 | 6542 | 0.111 |  |  |  | 1 | 13.5358 | 120.357 |  |  |  | 119.492 | 0 |  | 0.72423 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1064.9 | 2021/05/18 18:30:08 | 0
05/18/21 | 48 | 20210048.099 | 747 | 0.849 | 6539 | 0 |  |  |  | 2 | 13.5431 | 120.493 |  |  |  | 119.627 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.2; BGD 30: 1191.0 | 2021/05/18 18:30:08 | 0
05/18/21 | 48 | 20210048.099 | 747 | 0.849 | 4041 | 10.437 |  |  | Nitrogen | 3 | 11.9994 | 96.799 |  |  |  | 96.096 | 0 |  | 0.73171 |  | 0.0037166 |  | 0.37028 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1161.8 | 2021/05/18 18:30:08 | 0
05/18/21 | 48 | 20210048.099 | 747 | 0.849 |  |  | 892 | -32.791 | Carbon | 4 | 49.8074 | 24.229 |  |  | 23.855 |  | 0 | 1.1484 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.8 | 2021/05/18 18:30:08 | 95
05/18/21 | 48 | 20210048.099 | 747 | 0.849 |  |  | 5489 | -43.8 |  | 5 | 51.4068 | 102.739 |  |  | 101.179 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 7.9 | 2021/05/18 18:30:08 | 95
05/18/21 | 48 | 20210048.099 | 747 | 0.849 |  |  | 5489 | -43.984 |  | 6 | 50.9692 | 103.023 |  |  | 101.459 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.3 | 2021/05/18 18:30:08 | 95
05/18/21 | 49 | 20210048.100 | 847 | 0.815 | 6539 | 0.135 |  |  |  | 1 | 14.1017 | 120.379 |  |  |  | 119.513 | 0 |  | 0.72426 |  | 0.0036787 |  | 0.36652 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1068.3 | 2021/05/18 18:39:53 | 0
05/18/21 | 49 | 20210048.100 | 847 | 0.815 | 6536 | 0 |  |  |  | 2 | 14.1056 | 120.448 |  |  |  | 119.582 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1194.5 | 2021/05/18 18:39:53 | 0
05/18/21 | 49 | 20210048.100 | 847 | 0.815 | 2397 | 11.908 |  |  | Nitrogen | 3 | 8.5628 | 57.407 |  |  |  | 56.989 | 0 |  | 0.73278 |  | 0.003722 |  | 0.37082 | BGD 28: 43.0; BGD 29: 32.9; BGD 30: 1161.3 | 2021/05/18 18:39:53 | 0
05/18/21 | 49 | 20210048.100 | 847 | 0.815 |  |  | 964 | -35.331 | Carbon | 4 | 55.2931 | 26.362 |  |  | 25.957 |  | 0 | 1.1456 |  | 0.010785 |  | 1.067 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/18 18:39:53 | 95
05/18/21 | 49 | 20210048.100 | 847 | 0.815 |  |  | 5488 | -43.8 |  | 5 | 53.6684 | 102.666 |  |  | 101.107 |  | 1 | 1.135 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 8.0 | 2021/05/18 18:39:53 | 95
05/18/21 | 49 | 20210048.100 | 847 | 0.815 |  |  | 5483 | -44.001 |  | 6 | 53.2149 | 102.949 |  |  | 101.386 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.3 | 2021/05/18 18:39:53 | 95
05/18/21 | 50 | 20210048.101 | 849 | 0.833 | 6544 | 0.104 |  |  |  | 1 | 13.796 | 120.361 |  |  |  | 119.495 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.1; BGD 30: 1064.6 | 2021/05/18 18:49:38 | 0
05/18/21 | 50 | 20210048.101 | 849 | 0.833 | 6545 | 0 |  |  |  | 2 | 13.8004 | 120.442 |  |  |  | 119.576 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1190.9 | 2021/05/18 18:49:38 | 0
05/18/21 | 50 | 20210048.101 | 849 | 0.833 | 3440 | 12.408 |  |  | Nitrogen | 3 | 11.0021 | 82.414 |  |  |  | 81.814 | 0 |  | 0.73326 |  | 0.0037238 |  | 0.371 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1160.9 | 2021/05/18 18:49:38 | 0
05/18/21 | 50 | 20210048.101 | 849 | 0.833 |  |  | 922 | -34.271 | Carbon | 4 | 52.1536 | 25.104 |  |  | 24.717 |  | 0 | 1.1467 |  | 0.010797 |  | 1.0682 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.9 | 2021/05/18 18:49:38 | 95
05/18/21 | 50 | 20210048.101 | 849 | 0.833 |  |  | 5483 | -43.8 |  | 5 | 52.5429 | 102.644 |  |  | 101.086 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/18 18:49:38 | 95
05/18/21 | 50 | 20210048.101 | 849 | 0.833 |  |  | 5479 | -43.977 |  | 6 | 52.2083 | 102.858 |  |  | 101.296 |  | 0 | 1.1349 |  | 0.010689 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 9.4 | 2021/05/18 18:49:38 | 95
05/18/21 | 51 | 20210048.102 | 851 | 0.82 | 6538 | 0.119 |  |  |  | 1 | 14.0137 | 120.343 |  |  |  | 119.477 | 0 |  | 0.72436 |  | 0.0036786 |  | 0.36652 | BGD 28: 43.2; BGD 29: 33.1; BGD 30: 1067.4 | 2021/05/18 18:59:23 | 0
05/18/21 | 51 | 20210048.102 | 851 | 0.82 | 6530 | 0 |  |  |  | 2 | 14.0185 | 120.429 |  |  |  | 119.563 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1194.2 | 2021/05/18 18:59:23 | 0
05/18/21 | 51 | 20210048.102 | 851 | 0.82 | 3904 | 10.429 |  |  | Nitrogen | 3 | 12.1672 | 93.634 |  |  |  | 92.953 | 0 |  | 0.73183 |  | 0.0037166 |  | 0.37028 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1164.4 | 2021/05/18 18:59:23 | 0
05/18/21 | 51 | 20210048.102 | 851 | 0.82 |  |  | 818 | -30.963 | Carbon | 4 | 48.0222 | 22.111 |  |  | 21.77 |  | 0 | 1.1505 |  | 0.010834 |  | 1.0718 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.9 | 2021/05/18 18:59:23 | 95
05/18/21 | 51 | 20210048.102 | 851 | 0.82 |  |  | 5483 | -43.8 |  | 5 | 53.5222 | 102.551 |  |  | 100.994 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.9 | 2021/05/18 18:59:23 | 95
05/18/21 | 51 | 20210048.102 | 851 | 0.82 |  |  | 5475 | -44.001 |  | 6 | 53.2252 | 102.739 |  |  | 101.179 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/18 18:59:23 | 95
05/18/21 | 52 | 20210048.103 | 853 | 0.833 | 6532 | 0.119 |  |  |  | 1 | 13.8 | 120.433 |  |  |  | 119.567 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.0; BGD 29: 33.0; BGD 30: 1069.8 | 2021/05/18 19:09:08 | 0
05/18/21 | 52 | 20210048.103 | 853 | 0.833 | 6535 | 0 |  |  |  | 2 | 13.8023 | 120.475 |  |  |  | 119.609 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.0; BGD 29: 32.2; BGD 30: 1196.9 | 2021/05/18 19:09:08 | 0
05/18/21 | 52 | 20210048.103 | 853 | 0.833 | 3722 | 11.684 |  |  | Nitrogen | 3 | 11.5885 | 88.986 |  |  |  | 88.338 | 0 |  | 0.73266 |  | 0.0037212 |  | 0.37074 | BGD 28: 42.8; BGD 29: 32.7; BGD 30: 1166.0 | 2021/05/18 19:09:08 | 0
05/18/21 | 52 | 20210048.103 | 853 | 0.833 |  |  | 887 | -32.471 | Carbon | 4 | 50.4639 | 24.043 |  |  | 23.672 |  | 0 | 1.1488 |  | 0.010817 |  | 1.0701 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.9 | 2021/05/18 19:09:08 | 95
05/18/21 | 52 | 20210048.103 | 853 | 0.833 |  |  | 5480 | -43.8 |  | 5 | 52.6319 | 102.587 |  |  | 101.029 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/18 19:09:08 | 95
05/18/21 | 52 | 20210048.103 | 853 | 0.833 |  |  | 5472 | -44.015 |  | 6 | 52.378 | 102.749 |  |  | 101.19 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/18 19:09:08 | 95
05/18/21 | 53 | 20210048.104 | 855 | 0.825 | 6536 | 0.106 |  |  |  | 1 | 13.929 | 120.347 |  |  |  | 119.482 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1072.3 | 2021/05/18 19:18:54 | 0
05/18/21 | 53 | 20210048.104 | 855 | 0.825 | 6534 | 0 |  |  |  | 2 | 13.9224 | 120.226 |  |  |  | 119.362 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.2; BGD 30: 1198.9 | 2021/05/18 19:18:54 | 0
05/18/21 | 53 | 20210048.104 | 855 | 0.825 | 3690 | 12.721 |  |  | Nitrogen | 3 | 11.6639 | 88.561 |  |  |  | 87.916 | 0 |  | 0.73342 |  | 0.003725 |  | 0.37112 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1168.0 | 2021/05/18 19:18:54 | 0
05/18/21 | 53 | 20210048.104 | 855 | 0.825 |  |  | 872 | -31.982 | Carbon | 4 | 50.3137 | 23.652 |  |  | 23.287 |  | 0 | 1.1494 |  | 0.010823 |  | 1.0707 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 9.9 | 2021/05/18 19:18:54 | 95
05/18/21 | 53 | 20210048.104 | 855 | 0.825 |  |  | 5479 | -43.8 |  | 5 | 53.2982 | 102.488 |  |  | 100.931 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/18 19:18:54 | 95
05/18/21 | 53 | 20210048.104 | 855 | 0.825 |  |  | 5476 | -44.009 |  | 6 | 52.9671 | 102.698 |  |  | 101.139 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/18 19:18:54 | 95
05/18/21 | 54 | 20210048.105 | 857 | 0.848 | 6531 | 0.102 |  |  |  | 1 | 13.5466 | 120.261 |  |  |  | 119.396 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1073.0 | 2021/05/18 19:28:39 | 0
05/18/21 | 54 | 20210048.105 | 857 | 0.848 | 6537 | 0 |  |  |  | 2 | 13.5501 | 120.325 |  |  |  | 119.46 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.0; BGD 29: 32.2; BGD 30: 1199.5 | 2021/05/18 19:28:39 | 0
05/18/21 | 54 | 20210048.105 | 857 | 0.848 | 3649 | 12.047 |  |  | Nitrogen | 3 | 11.2623 | 87.561 |  |  |  | 86.924 | 0 |  | 0.73295 |  | 0.0037225 |  | 0.37087 | BGD 28: 42.8; BGD 29: 32.7; BGD 30: 1169.4 | 2021/05/18 19:28:39 | 0
05/18/21 | 54 | 20210048.105 | 857 | 0.848 |  |  | 921 | -32.47 | Carbon | 4 | 51.2648 | 25.126 |  |  | 24.738 |  | 0 | 1.1489 |  | 0.010817 |  | 1.0701 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/18 19:28:39 | 95
05/18/21 | 54 | 20210048.105 | 857 | 0.848 |  |  | 5483 | -43.8 |  | 5 | 51.7849 | 102.532 |  |  | 100.975 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 19:28:39 | 95
05/18/21 | 54 | 20210048.105 | 857 | 0.848 |  |  | 5482 | -44.009 |  | 6 | 51.3679 | 102.804 |  |  | 101.243 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/18 19:28:39 | 95
05/18/21 | 55 | 20210048.106 | 859 | 0.848 | 6523 | 0.104 |  |  |  | 1 | 13.5451 | 120.234 |  |  |  | 119.369 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1074.1 | 2021/05/18 19:38:24 | 0
05/18/21 | 55 | 20210048.106 | 859 | 0.848 | 6526 | 0 |  |  |  | 2 | 13.548 | 120.286 |  |  |  | 119.421 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1201.3 | 2021/05/18 19:38:24 | 0
05/18/21 | 55 | 20210048.106 | 859 | 0.848 | 3771 | 12.024 |  |  | Nitrogen | 3 | 11.5238 | 90.663 |  |  |  | 90.003 | 0 |  | 0.73294 |  | 0.0037224 |  | 0.37086 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 1170.0 | 2021/05/18 19:38:24 | 0
05/18/21 | 55 | 20210048.106 | 859 | 0.848 |  |  | 894 | -32.933 | Carbon | 4 | 50.0344 | 24.336 |  |  | 23.961 |  | 0 | 1.1484 |  | 0.010812 |  | 1.0696 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/18 19:38:24 | 95
05/18/21 | 55 | 20210048.106 | 859 | 0.848 |  |  | 5480 | -43.8 |  | 5 | 51.8557 | 102.486 |  |  | 100.929 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 19:38:24 | 95
05/18/21 | 55 | 20210048.106 | 859 | 0.848 |  |  | 5474 | -44.029 |  | 6 | 51.3888 | 102.79 |  |  | 101.23 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 9.5 | 2021/05/18 19:38:24 | 95
05/18/21 | 56 | 20210048.107 | 861 | 0.806 | 6531 | 0.136 |  |  |  | 1 | 14.2531 | 120.272 |  |  |  | 119.407 | 0 |  | 0.72425 |  | 0.0036787 |  | 0.36652 | BGD 28: 43.0; BGD 29: 32.9; BGD 30: 1075.4 | 2021/05/18 19:48:09 | 0
05/18/21 | 56 | 20210048.107 | 861 | 0.806 | 6539 | 0 |  |  |  | 2 | 14.2595 | 120.384 |  |  |  | 119.518 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.0; BGD 29: 32.2; BGD 30: 1202.8 | 2021/05/18 19:48:09 | 0
05/18/21 | 56 | 20210048.107 | 861 | 0.806 | 3701 | 12.5 |  |  | Nitrogen | 3 | 11.9651 | 88.855 |  |  |  | 88.208 | 0 |  | 0.73321 |  | 0.0037242 |  | 0.37104 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 1171.7 | 2021/05/18 19:48:09 | 0
05/18/21 | 56 | 20210048.107 | 861 | 0.806 |  |  | 847 | -32.174 | Carbon | 4 | 50.3474 | 22.973 |  |  | 22.619 |  | 0 | 1.1493 |  | 0.010821 |  | 1.0705 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/18 19:48:09 | 95
05/18/21 | 56 | 20210048.107 | 861 | 0.806 |  |  | 5484 | -43.8 |  | 5 | 54.5342 | 102.5 |  |  | 100.944 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 19:48:09 | 95
05/18/21 | 56 | 20210048.107 | 861 | 0.806 |  |  | 5478 | -44.002 |  | 6 | 54.1685 | 102.727 |  |  | 101.168 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/18 19:48:09 | 95
05/18/21 | 57 | 20210048.108 | 863 | 0.812 | 6534 | 0.119 |  |  |  | 1 | 14.1578 | 120.449 |  |  |  | 119.583 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36652 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1076.3 | 2021/05/18 19:57:55 | 0
05/18/21 | 57 | 20210048.108 | 863 | 0.812 | 6543 | 0 |  |  |  | 2 | 14.1613 | 120.512 |  |  |  | 119.646 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1204.0 | 2021/05/18 19:57:55 | 0
05/18/21 | 57 | 20210048.108 | 863 | 0.812 | 3878 | 9.885 |  |  | Nitrogen | 3 | 12.2563 | 93.265 |  |  |  | 92.588 | 0 |  | 0.73141 |  | 0.0037146 |  | 0.37008 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1174.0 | 2021/05/18 19:57:55 | 0
05/18/21 | 57 | 20210048.108 | 863 | 0.812 |  |  | 826 | -28.967 | Carbon | 4 | 48.952 | 22.376 |  |  | 22.029 |  | 0 | 1.1529 |  | 0.010856 |  | 1.074 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/18 19:57:55 | 95
05/18/21 | 57 | 20210048.108 | 863 | 0.812 |  |  | 5480 | -43.8 |  | 5 | 54.0942 | 102.523 |  |  | 100.967 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 19:57:55 | 95
05/18/21 | 57 | 20210048.108 | 863 | 0.812 |  |  | 5479 | -44.008 |  | 6 | 53.6816 | 102.781 |  |  | 101.221 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/18 19:57:55 | 95
05/18/21 | 58 | 20210048.109 | 865 | 0.827 | 6530 | 0.1 |  |  |  | 1 | 13.8923 | 120.291 |  |  |  | 119.426 | 0 |  | 0.72426 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1077.3 | 2021/05/18 20:07:40 | 0
05/18/21 | 58 | 20210048.109 | 865 | 0.827 | 6538 | 0 |  |  |  | 2 | 13.9032 | 120.49 |  |  |  | 119.624 | 1 |  | 0.72418 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.2; BGD 30: 1204.9 | 2021/05/18 20:07:40 | 0
05/18/21 | 58 | 20210048.109 | 865 | 0.827 | 3753 | 12.073 |  |  | Nitrogen | 3 | 11.7676 | 90.09 |  |  |  | 89.435 | 0 |  | 0.73293 |  | 0.0037226 |  | 0.37088 | BGD 28: 42.9; BGD 29: 32.7; BGD 30: 1173.7 | 2021/05/18 20:07:40 | 0
05/18/21 | 58 | 20210048.109 | 865 | 0.827 |  |  | 874 | -31.502 | Carbon | 4 | 50.4131 | 23.787 |  |  | 23.42 |  | 0 | 1.1501 |  | 0.010828 |  | 1.0712 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 10.1 | 2021/05/18 20:07:40 | 95
05/18/21 | 58 | 20210048.109 | 865 | 0.827 |  |  | 5487 | -43.8 |  | 5 | 53.0058 | 102.592 |  |  | 101.034 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 20:07:40 | 95
05/18/21 | 58 | 20210048.109 | 865 | 0.827 |  |  | 5477 | -44.005 |  | 6 | 52.7678 | 102.743 |  |  | 101.183 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/18 20:07:40 | 95
05/18/21 | 59 | 20210048.110 | 867 | 0.816 | 6543 | 0.12 |  |  |  | 1 | 14.0833 | 120.359 |  |  |  | 119.493 | 0 |  | 0.72427 |  | 0.0036786 |  | 0.36652 | BGD 28: 43.1; BGD 29: 33.0; BGD 30: 1078.1 | 2021/05/18 20:17:25 | 0
05/18/21 | 59 | 20210048.110 | 867 | 0.816 | 6543 | 0 |  |  |  | 2 | 14.0905 | 120.487 |  |  |  | 119.621 | 1 |  | 0.72418 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1206.2 | 2021/05/18 20:17:25 | 0
05/18/21 | 59 | 20210048.110 | 867 | 0.816 | 2236 | 13.475 |  |  | Nitrogen | 3 | 8.0812 | 53.539 |  |  |  | 53.149 | 0 |  | 0.73394 |  | 0.0037278 |  | 0.37139 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1170.0 | 2021/05/18 20:17:25 | 0
05/18/21 | 59 | 20210048.110 | 867 | 0.816 |  |  | 1032 | -35.787 | Carbon | 4 | 58.2521 | 28.367 |  |  | 27.93 |  | 0 | 1.1453 |  | 0.01078 |  | 1.0665 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.9 | 2021/05/18 20:17:25 | 95
05/18/21 | 59 | 20210048.110 | 867 | 0.816 |  |  | 5491 | -43.8 |  | 5 | 53.5109 | 102.723 |  |  | 101.163 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 8.2 | 2021/05/18 20:17:25 | 95
05/18/21 | 59 | 20210048.110 | 867 | 0.816 |  |  | 5490 | -44.015 |  | 6 | 53.1696 | 102.936 |  |  | 101.374 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 9.6 | 2021/05/18 20:17:25 | 95
05/18/21 | 60 | 20210048.110 | Replicate E11 | 0.845 | 6539 | 0.112 |  |  |  | 1 | 13.6033 | 120.42 |  |  |  | 119.555 | 0 |  | 0.72428 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.1; BGD 29: 33.1; BGD 30: 1073.4 | 2021/05/18 20:27:10 | 0
05/18/21 | 60 | 20210048.110 | Replicate E11 | 0.845 | 6549 | 0 |  |  |  | 2 | 13.6104 | 120.552 |  |  |  | 119.685 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.1; BGD 29: 32.3; BGD 30: 1201.9 | 2021/05/18 20:27:10 | 0
05/18/21 | 60 | 20210048.110 | Replicate E11 | 0.845 | 2727 | 13.506 |  |  | Nitrogen | 3 | 9.1316 | 65.207 |  |  |  | 64.732 | 0 |  | 0.73398 |  | 0.0037279 |  | 0.3714 | BGD 28: 43.0; BGD 29: 32.8; BGD 30: 1168.6 | 2021/05/18 20:27:10 | 0
05/18/21 | 60 | 20210048.110 | Replicate E11 | 0.845 |  |  | 1032 | -34.933 | Carbon | 4 | 56.1672 | 28.306 |  |  | 27.87 |  | 0 | 1.1463 |  | 0.01079 |  | 1.0674 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/18 20:27:10 | 95
05/18/21 | 60 | 20210048.110 | Replicate E11 | 0.845 |  |  | 5491 | -43.8 |  | 5 | 51.6928 | 102.711 |  |  | 101.152 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 8.2 | 2021/05/18 20:27:10 | 95
05/18/21 | 60 | 20210048.110 | Replicate E11 | 0.845 |  |  | 5484 | -44 |  | 6 | 51.3909 | 102.907 |  |  | 101.344 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 9.5 | 2021/05/18 20:27:10 | 95
05/18/21 | 61 | 20210048.111 | 869 | 0.834 | 6548 | 0.099 |  |  |  | 1 | 13.7863 | 120.486 |  |  |  | 119.62 | 0 |  | 0.72428 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.2; BGD 29: 33.2; BGD 30: 1073.3 | 2021/05/18 20:36:55 | 0
05/18/21 | 61 | 20210048.111 | 869 | 0.834 | 6551 | 0 |  |  |  | 2 | 13.7939 | 120.625 |  |  |  | 119.757 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 42.2; BGD 29: 32.4; BGD 30: 1201.4 | 2021/05/18 20:36:55 | 0
05/18/21 | 61 | 20210048.111 | 869 | 0.834 |  |  | 5493 | -43.8 |  | 3 | 52.2032 | 102.821 |  |  | 101.259 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.5; BGD 45: 0.6; BGD 46: 7.3 | 2021/05/18 20:36:55 | 95
05/18/21 | 61 | 20210048.111 | 869 | 0.834 |  |  | 5487 | -44.046 |  | 4 | 51.8325 | 103.057 |  |  | 101.492 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.3; BGD 45: 1.5; BGD 46: 9.0 | 2021/05/18 20:36:55 | 95
05/18/21 | 62 | 20210048.112 | 871 | 0.825 | 6533 | 0.132 |  |  |  | 1 | 13.9449 | 120.635 |  |  |  | 119.767 | 0 |  | 0.7243 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.2; BGD 29: 32.1; BGD 30: 1055.3 | 2021/05/18 20:46:41 | 0
05/18/21 | 62 | 20210048.112 | 871 | 0.825 | 6543 | 0 |  |  |  | 2 | 13.9388 | 120.523 |  |  |  | 119.657 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.3; BGD 29: 31.4; BGD 30: 1183.6 | 2021/05/18 20:46:41 | 0
05/18/21 | 62 | 20210048.112 | 871 | 0.825 | 3917 | 11.976 |  |  | Nitrogen | 3 | 12.1347 | 94.137 |  |  |  | 93.452 | 0 |  | 0.73287 |  | 0.0037222 |  | 0.37084 | BGD 28: 42.2; BGD 29: 32.1; BGD 30: 1154.3 | 2021/05/18 20:46:41 | 0
05/18/21 | 62 | 20210048.112 | 871 | 0.825 |  |  | 843 | -31.703 | Carbon | 4 | 49.0339 | 22.881 |  |  | 22.528 |  | 0 | 1.1499 |  | 0.010826 |  | 1.071 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/18 20:46:41 | 95
05/18/21 | 62 | 20210048.112 | 871 | 0.825 |  |  | 5494 | -43.8 |  | 5 | 52.9278 | 102.723 |  |  | 101.163 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/18 20:46:41 | 95
05/18/21 | 62 | 20210048.112 | 871 | 0.825 |  |  | 5488 | -43.998 |  | 6 | 52.6291 | 102.912 |  |  | 101.349 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/18 20:46:41 | 95
05/18/21 | 63 | 20210048.113 | 873 | 0.839 | 6548 | 0.121 |  |  |  | 1 | 13.7121 | 120.632 |  |  |  | 119.765 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36652 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 1065.5 | 2021/05/18 20:56:26 | 0
05/18/21 | 63 | 20210048.113 | 873 | 0.839 | 6550 | 0 |  |  |  | 2 | 13.7153 | 120.693 |  |  |  | 119.825 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.0; BGD 30: 1193.5 | 2021/05/18 20:56:26 | 0
05/18/21 | 63 | 20210048.113 | 873 | 0.839 | 4048 | 11.291 |  |  | Nitrogen | 3 | 12.1874 | 97.381 |  |  |  | 96.673 | 0 |  | 0.73238 |  | 0.0037197 |  | 0.37059 | BGD 28: 42.8; BGD 29: 32.6; BGD 30: 1164.1 | 2021/05/18 20:56:26 | 0
05/18/21 | 63 | 20210048.113 | 873 | 0.839 |  |  | 850 | -31.78 | Carbon | 4 | 48.5628 | 23.092 |  |  | 22.736 |  | 0 | 1.1499 |  | 0.010825 |  | 1.0709 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/18 20:56:26 | 95
05/18/21 | 63 | 20210048.113 | 873 | 0.839 |  |  | 5499 | -43.8 |  | 5 | 51.8909 | 102.822 |  |  | 101.26 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 20:56:26 | 95
05/18/21 | 63 | 20210048.113 | 873 | 0.839 |  |  | 5504 | -44.005 |  | 6 | 51.613 | 103 |  |  | 101.436 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/18 20:56:26 | 95
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.28 | 36-UWSIF-UT Glut 1- | 0.827 | 6548 | 0.104 |  |  |  | 1 | 13.9123 | 120.654 |  |  |  | 119.787 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 43.0; BGD 29: 32.9; BGD 30: 1070.4 | 2021/05/18 21:06:11 | 0
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.28 | 36-UWSIF-UT Glut 1- | 0.827 | 6548 | 0 |  |  |  | 2 | 13.9138 | 120.682 |  |  |  | 119.815 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1198.0 | 2021/05/18 21:06:11 | 0
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.28 | 36-UWSIF-UT Glut 1- | 0.827 | 2538 | -2.827 |  |  | Nitrogen | 3 | 8.8726 | 61.137 |  |  |  | 60.698 | 0 |  | 0.72217 |  | 0.0036678 |  | 0.36544 | BGD 28: 42.8; BGD 29: 32.6; BGD 30: 1164.4 | 2021/05/18 21:06:11 | 0
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.28 | 36-UWSIF-UT Glut 1- | 0.827 |  |  | 717 | -29.749 | Carbon | 4 | 42.6017 | 19.265 |  |  | 18.967 |  | 0 | 1.1522 |  | 0.010848 |  | 1.0731 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.9 | 2021/05/18 21:06:11 | 95
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.28 | 36-UWSIF-UT Glut 1- | 0.827 |  |  | 5499 | -43.8 |  | 5 | 52.5615 | 102.874 |  |  | 101.312 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.0 | 2021/05/18 21:06:11 | 95
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.28 | 36-UWSIF-UT Glut 1- | 0.827 |  |  | 5494 | -44.018 |  | 6 | 52.3845 | 102.986 |  |  | 101.422 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/18 21:06:11 | 95
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.29 | 36-UWSIF-UT Glut 1- | 0.791 | 6548 | 0.104 |  |  |  | 1 | 14.5511 | 120.754 |  |  |  | 119.886 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.7; BGD 30: 1065.7 | 2021/05/18 21:15:56 | 0
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.29 | 36-UWSIF-UT Glut 1- | 0.791 | 6554 | 0 |  |  |  | 2 | 14.551 | 120.752 |  |  |  | 119.884 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1194.2 | 2021/05/18 21:15:56 | 0
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.29 | 36-UWSIF-UT Glut 1- | 0.791 | 2411 | -2.848 |  |  | Nitrogen | 3 | 8.9603 | 58.525 |  |  |  | 58.106 | 0 |  | 0.72215 |  | 0.0036677 |  | 0.36543 | BGD 28: 42.7; BGD 29: 32.5; BGD 30: 1160.8 | 2021/05/18 21:15:56 | 0
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.29 | 36-UWSIF-UT Glut 1- | 0.791 |  |  | 680 | -29.812 | Carbon | 4 | 42.7513 | 18.334 |  |  | 18.051 |  | 0 | 1.1521 |  | 0.010847 |  | 1.0731 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.9 | 2021/05/18 21:15:56 | 95
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.29 | 36-UWSIF-UT Glut 1- | 0.791 |  |  | 5503 | -43.8 |  | 5 | 54.8439 | 102.94 |  |  | 101.377 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.9 | 2021/05/18 21:15:56 | 95
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.29 | 36-UWSIF-UT Glut 1- | 0.791 |  |  | 5502 | -44.036 |  | 6 | 54.3306 | 103.249 |  |  | 101.681 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/18 21:15:56 | 95
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 39-UWSIF-UW Glut 2- | 0.716 | 6556 | 0.11 |  |  |  | 1 | 16.0765 | 120.771 |  |  |  | 119.903 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.7; BGD 30: 1063.6 | 2021/05/18 21:25:42 | 0
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 39-UWSIF-UW Glut 2- | 0.716 | 6559 | 0 |  |  |  | 2 | 16.0829 | 120.874 |  |  |  | 120.005 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 31.9; BGD 30: 1191.6 | 2021/05/18 21:25:42 | 0
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 39-UWSIF-UW Glut 2- | 0.716 | 2387 | 29.695 |  |  | Nitrogen | 3 | 9.804 | 57.827 |  |  |  | 57.399 | 0 |  | 0.74572 |  | 0.0037874 |  | 0.37731 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1159.7 | 2021/05/18 21:25:42 | 0
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 39-UWSIF-UW Glut 2- | 0.716 |  |  | 675 | 22.277 | Carbon | 4 | 46.8616 | 18.163 |  |  | 17.872 |  | 0 | 1.2096 |  | 0.011429 |  | 1.13 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 9.9 | 2021/05/18 21:25:42 | 95
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 39-UWSIF-UW Glut 2- | 0.716 |  |  | 5518 | -43.8 |  | 5 | 60.4521 | 103.015 |  |  | 101.45 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.9 | 2021/05/18 21:25:42 | 95
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 39-UWSIF-UW Glut 2- | 0.716 |  |  | 5501 | -44.019 |  | 6 | 60.2469 | 103.127 |  |  | 101.561 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.4 | 2021/05/18 21:25:42 | 95
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 39-UWSIF-UW Glut 2- | 0.798 | 6547 | 0.115 |  |  |  | 1 | 14.4173 | 120.645 |  |  |  | 119.778 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.7; BGD 29: 32.7; BGD 30: 1060.9 | 2021/05/18 21:35:27 | 0
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 39-UWSIF-UW Glut 2- | 0.798 | 6559 | 0 |  |  |  | 2 | 14.4263 | 120.803 |  |  |  | 119.935 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 31.9; BGD 30: 1189.2 | 2021/05/18 21:35:27 | 0
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 39-UWSIF-UW Glut 2- | 0.798 | 2665 | 29.723 |  |  | Nitrogen | 3 | 9.5744 | 64.379 |  |  |  | 63.903 | 0 |  | 0.74574 |  | 0.0037875 |  | 0.37732 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1157.3 | 2021/05/18 21:35:27 | 0
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 39-UWSIF-UW Glut 2- | 0.798 |  |  | 753 | 22.299 | Carbon | 4 | 46.1759 | 20.353 |  |  | 20.027 |  | 0 | 1.2096 |  | 0.01143 |  | 1.13 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 9.9 | 2021/05/18 21:35:27 | 95
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 39-UWSIF-UW Glut 2- | 0.798 |  |  | 5508 | -43.8 |  | 5 | 54.2258 | 103.024 |  |  | 101.459 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.9 | 2021/05/18 21:35:27 | 95
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 39-UWSIF-UW Glut 2- | 0.798 |  |  | 5506 | -44.009 |  | 6 | 53.9363 | 103.199 |  |  | 101.632 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/18 21:35:27 | 95
05/18/21 | 68 | 01-UWSIF-Liver-20210048.25 | 01-UWSIF-Liver- | 0.741 | 6556 | 0.103 |  |  |  | 1 | 15.5321 | 120.74 |  |  |  | 119.872 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.7; BGD 29: 32.7; BGD 30: 1060.8 | 2021/05/18 21:45:12 | 0
05/18/21 | 68 | 01-UWSIF-Liver-20210048.25 | 01-UWSIF-Liver- | 0.741 | 6562 | 0 |  |  |  | 2 | 15.5417 | 120.897 |  |  |  | 120.028 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 31.9; BGD 30: 1188.9 | 2021/05/18 21:45:12 | 0
05/18/21 | 68 | 01-UWSIF-Liver-20210048.25 | 01-UWSIF-Liver- | 0.741 | 3140 | 8.538 |  |  | Nitrogen | 3 | 11.6248 | 75.521 |  |  |  | 74.974 | 0 |  | 0.7304 |  | 0.0037096 |  | 0.36959 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1157.2 | 2021/05/18 21:45:12 | 0
05/18/21 | 68 | 01-UWSIF-Liver-20210048.25 | 01-UWSIF-Liver- | 0.741 |  |  | 795 | -19.359 | Carbon | 4 | 51.9879 | 21.509 |  |  | 21.174 |  | 0 | 1.1636 |  | 0.010964 |  | 1.0845 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/18 21:45:12 | 95
05/18/21 | 68 | 01-UWSIF-Liver-20210048.25 | 01-UWSIF-Liver- | 0.741 |  |  | 5515 | -43.8 |  | 5 | 58.3021 | 103.077 |  |  | 101.512 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/18 21:45:12 | 95
05/18/21 | 68 | 01-UWSIF-Liver-20210048.25 | 01-UWSIF-Liver- | 0.741 |  |  | 5501 | -44.007 |  | 6 | 58.066 | 103.21 |  |  | 101.643 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/18 21:45:12 | 95
05/18/21 | 69 | 01-UWSIF-Liver-20210048.26 | 01-UWSIF-Liver- | 0.767 | 6567 | 0.1 |  |  |  | 1 | 15.0132 | 120.869 |  |  |  | 120 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1062.3 | 2021/05/18 21:54:58 | 0
05/18/21 | 69 | 01-UWSIF-Liver-20210048.26 | 01-UWSIF-Liver- | 0.767 | 6566 | 0 |  |  |  | 2 | 15.0131 | 120.866 |  |  |  | 119.997 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1190.4 | 2021/05/18 21:54:58 | 0
05/18/21 | 69 | 01-UWSIF-Liver-20210048.26 | 01-UWSIF-Liver- | 0.767 | 3266 | 8.529 |  |  | Nitrogen | 3 | 11.5586 | 78.608 |  |  |  | 78.038 | 0 |  | 0.7304 |  | 0.0037096 |  | 0.36959 | BGD 28: 42.7; BGD 29: 32.5; BGD 30: 1158.3 | 2021/05/18 21:54:58 | 0
05/18/21 | 69 | 01-UWSIF-Liver-20210048.26 | 01-UWSIF-Liver- | 0.767 |  |  | 823 | -19.484 | Carbon | 4 | 51.7181 | 22.318 |  |  | 21.97 |  | 0 | 1.1635 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/18 21:54:58 | 95
05/18/21 | 69 | 01-UWSIF-Liver-20210048.26 | 01-UWSIF-Liver- | 0.767 |  |  | 5511 | -43.8 |  | 5 | 56.2529 | 103.12 |  |  | 101.554 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 21:54:58 | 95
05/18/21 | 69 | 01-UWSIF-Liver-20210048.26 | 01-UWSIF-Liver- | 0.767 |  |  | 5506 | -44.022 |  | 6 | 56.207 | 103.146 |  |  | 101.58 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/18 21:54:58 | 95
05/18/21 | 70 | 20210048.114 | 875 | 0.826 | 6562 | 0.121 |  |  |  | 1 | 13.9404 | 120.861 |  |  |  | 119.992 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36652 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1063.2 | 2021/05/18 22:04:43 | 0
05/18/21 | 70 | 20210048.114 | 875 | 0.826 | 6570 | 0 |  |  |  | 2 | 13.9443 | 120.932 |  |  |  | 120.062 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1190.9 | 2021/05/18 22:04:43 | 0
05/18/21 | 70 | 20210048.114 | 875 | 0.826 | 3949 | 9.294 |  |  | Nitrogen | 3 | 12.2015 | 95.143 |  |  |  | 94.453 | 0 |  | 0.73095 |  | 0.0037124 |  | 0.36986 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1162.5 | 2021/05/18 22:04:43 | 0
05/18/21 | 70 | 20210048.114 | 875 | 0.826 |  |  | 845 | -31.478 | Carbon | 4 | 49.2464 | 23.044 |  |  | 22.688 |  | 0 | 1.1502 |  | 0.010828 |  | 1.0712 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/18 22:04:43 | 95
05/18/21 | 70 | 20210048.114 | 875 | 0.826 |  |  | 5508 | -43.8 |  | 5 | 52.2732 | 103.096 |  |  | 101.53 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 22:04:43 | 95
05/18/21 | 70 | 20210048.114 | 875 | 0.826 |  |  | 5511 | -44 |  | 6 | 51.7618 | 103.415 |  |  | 101.845 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/18 22:04:43 | 95
05/18/21 | 71 | 20210048.115 | 877 | 0.832 | 6562 | 0.115 |  |  |  | 1 | 13.8423 | 120.905 |  |  |  | 120.036 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 1066.4 | 2021/05/18 22:14:28 | 0
05/18/21 | 71 | 20210048.115 | 877 | 0.832 | 6563 | 0 |  |  |  | 2 | 13.8411 | 120.883 |  |  |  | 120.014 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1194.5 | 2021/05/18 22:14:28 | 0
05/18/21 | 71 | 20210048.115 | 877 | 0.832 | 3197 | 12.098 |  |  | Nitrogen | 3 | 10.4747 | 76.749 |  |  |  | 76.191 | 0 |  | 0.73297 |  | 0.0037227 |  | 0.37089 | BGD 28: 42.8; BGD 29: 32.7; BGD 30: 1161.8 | 2021/05/18 22:14:28 | 0
05/18/21 | 71 | 20210048.115 | 877 | 0.832 |  |  | 947 | -35.144 | Carbon | 4 | 53.4562 | 25.901 |  |  | 25.502 |  | 0 | 1.1461 |  | 0.010787 |  | 1.0672 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/18 22:14:28 | 95
05/18/21 | 71 | 20210048.115 | 877 | 0.832 |  |  | 5512 | -43.8 |  | 5 | 51.9072 | 103.089 |  |  | 101.523 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/18 22:14:28 | 95
05/18/21 | 71 | 20210048.115 | 877 | 0.832 |  |  | 5503 | -44.005 |  | 6 | 51.6506 | 103.251 |  |  | 101.683 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/18 22:14:28 | 95
05/18/21 | 72 | 20210048.116 | 879 | 0.824 | 6563 | 0.086 |  |  |  | 1 | 13.9727 | 120.832 |  |  |  | 119.963 | 0 |  | 0.72428 |  | 0.0036785 |  | 0.3665 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 1065.4 | 2021/05/18 22:24:13 | 0
05/18/21 | 72 | 20210048.116 | 879 | 0.824 | 6565 | 0 |  |  |  | 2 | 13.9779 | 120.926 |  |  |  | 120.057 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1193.0 | 2021/05/18 22:24:13 | 0
05/18/21 | 72 | 20210048.116 | 879 | 0.824 | 4014 | 12.092 |  |  | Nitrogen | 3 | 12.3355 | 96.447 |  |  |  | 95.745 | 0 |  | 0.73297 |  | 0.0037227 |  | 0.37089 | BGD 28: 42.8; BGD 29: 32.6; BGD 30: 1163.0 | 2021/05/18 22:24:13 | 0
05/18/21 | 72 | 20210048.116 | 879 | 0.824 |  |  | 836 | -31.314 | Carbon | 4 | 48.7387 | 22.671 |  |  | 22.32 |  | 0 | 1.1504 |  | 0.01083 |  | 1.0714 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/18 22:24:13 | 95
05/18/21 | 72 | 20210048.116 | 879 | 0.824 |  |  | 5511 | -43.8 |  | 5 | 52.4421 | 103.069 |  |  | 101.504 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 22:24:13 | 95
05/18/21 | 72 | 20210048.116 | 879 | 0.824 |  |  | 5506 | -44.011 |  | 6 | 52.0569 | 103.31 |  |  | 101.741 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/18 22:24:13 | 95
05/18/21 | 73 | 20210048.117 | 881 | 0.845 | 6565 | 0.105 |  |  |  | 1 | 13.6272 | 120.865 |  |  |  | 119.996 | 0 |  | 0.72428 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 1068.0 | 2021/05/18 22:33:59 | 0
05/18/21 | 73 | 20210048.117 | 881 | 0.845 | 6569 | 0 |  |  |  | 2 | 13.6319 | 120.953 |  |  |  | 120.083 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1195.4 | 2021/05/18 22:33:59 | 0
05/18/21 | 73 | 20210048.117 | 881 | 0.845 | 4095 | 11.592 |  |  | Nitrogen | 3 | 12.2022 | 98.717 |  |  |  | 97.999 | 0 |  | 0.7326 |  | 0.0037208 |  | 0.3707 | BGD 28: 42.8; BGD 29: 32.6; BGD 30: 1165.7 | 2021/05/18 22:33:59 | 0
05/18/21 | 73 | 20210048.117 | 881 | 0.845 |  |  | 857 | -31.727 | Carbon | 4 | 48.5883 | 23.321 |  |  | 22.96 |  | 0 | 1.1499 |  | 0.010826 |  | 1.071 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 10.2 | 2021/05/18 22:33:59 | 95
05/18/21 | 73 | 20210048.117 | 881 | 0.845 |  |  | 5510 | -43.8 |  | 5 | 51.2039 | 103.027 |  |  | 101.463 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/18 22:33:59 | 95
05/18/21 | 73 | 20210048.117 | 881 | 0.845 |  |  | 5508 | -44.003 |  | 6 | 50.8059 | 103.283 |  |  | 101.714 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/18 22:33:59 | 95
05/18/21 | 74 | 20210048.118 | 883 | 0.841 | 6567 | 0.115 |  |  |  | 1 | 13.6889 | 120.808 |  |  |  | 119.939 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 1069.7 | 2021/05/18 22:43:44 | 0
05/18/21 | 74 | 20210048.118 | 883 | 0.841 | 6566 | 0 |  |  |  | 2 | 13.6938 | 120.898 |  |  |  | 120.029 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1197.0 | 2021/05/18 22:43:44 | 0
05/18/21 | 74 | 20210048.118 | 883 | 0.841 | 3957 | 12.331 |  |  | Nitrogen | 3 | 12.0054 | 95.416 |  |  |  | 94.721 | 0 |  | 0.73314 |  | 0.0037236 |  | 0.37097 | BGD 28: 42.8; BGD 29: 32.6; BGD 30: 1165.5 | 2021/05/18 22:43:44 | 0
05/18/21 | 74 | 20210048.118 | 883 | 0.841 |  |  | 848 | -31.341 | Carbon | 4 | 48.3633 | 23.041 |  |  | 22.685 |  | 0 | 1.1503 |  | 0.01083 |  | 1.0714 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/18 22:43:44 | 95
05/18/21 | 74 | 20210048.118 | 883 | 0.841 |  |  | 5512 | -43.8 |  | 5 | 51.4261 | 103.041 |  |  | 101.476 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/18 22:43:44 | 95
05/18/21 | 74 | 20210048.118 | 883 | 0.841 |  |  | 5512 | -44.015 |  | 6 | 50.8811 | 103.388 |  |  | 101.819 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/18 22:43:44 | 95
05/18/21 | 75 | 20210048.119 | 885 | 0.807 | 6558 | 0.128 |  |  |  | 1 | 14.267 | 120.832 |  |  |  | 119.964 | 0 |  | 0.72428 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.9; BGD 29: 32.9; BGD 30: 1069.8 | 2021/05/18 22:53:29 | 0
05/18/21 | 75 | 20210048.119 | 885 | 0.807 | 6559 | 0 |  |  |  | 2 | 14.2674 | 120.838 |  |  |  | 119.97 | 1 |  | 0.72419 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1197.0 | 2021/05/18 22:53:29 | 0
05/18/21 | 75 | 20210048.119 | 885 | 0.807 | 3250 | 13.858 |  |  | Nitrogen | 3 | 10.9535 | 78.285 |  |  |  | 77.714 | 0 |  | 0.73423 |  | 0.0037292 |  | 0.37153 | BGD 28: 42.8; BGD 29: 32.6; BGD 30: 1164.5 | 2021/05/18 22:53:29 | 0
05/18/21 | 75 | 20210048.119 | 885 | 0.807 |  |  | 902 | -32.93 | Carbon | 4 | 53.0549 | 24.626 |  |  | 24.246 |  | 0 | 1.1486 |  | 0.010812 |  | 1.0696 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/18 22:53:29 | 95
05/18/21 | 75 | 20210048.119 | 885 | 0.807 |  |  | 5513 | -43.8 |  | 5 | 53.4955 | 103.101 |  |  | 101.535 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/18 22:53:29 | 95
05/18/21 | 75 | 20210048.119 | 885 | 0.807 |  |  | 5507 | -44.012 |  | 6 | 53.238 | 103.258 |  |  | 101.691 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/18 22:53:29 | 95
05/18/21 | 76 | 20210048.120 | 887 | 0.849 | 6554 | 0.112 |  |  |  | 1 | 13.5532 | 120.681 |  |  |  | 119.814 | 0 |  | 0.72428 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1066.1 | 2021/05/18 23:03:18 | 0
05/18/21 | 76 | 20210048.120 | 887 | 0.849 | 6558 | 0 |  |  |  | 2 | 13.5624 | 120.853 |  |  |  | 119.984 | 1 |  | 0.72419 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1193.3 | 2021/05/18 23:03:18 | 0
05/18/21 | 76 | 20210048.120 | 887 | 0.849 | 4153 | 11.666 |  |  | Nitrogen | 3 | 12.2303 | 99.87 |  |  |  | 99.143 | 0 |  | 0.73264 |  | 0.0037211 |  | 0.37073 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1164.0 | 2021/05/18 23:03:18 | 0
05/18/21 | 76 | 20210048.120 | 887 | 0.849 |  |  | 857 | -30.451 | Carbon | 4 | 48.353 | 23.317 |  |  | 22.956 |  | 0 | 1.1513 |  | 0.01084 |  | 1.0724 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/18 23:03:18 | 95
05/18/21 | 76 | 20210048.120 | 887 | 0.849 |  |  | 5513 | -43.8 |  | 5 | 50.9354 | 103.045 |  |  | 101.48 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/18 23:03:18 | 95
05/18/21 | 76 | 20210048.120 | 887 | 0.849 |  |  | 5506 | -44.015 |  | 6 | 50.5482 | 103.294 |  |  | 101.726 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/18 23:03:18 | 95
05/18/21 | 77 | 20210048.121 | 889 | 0.811 | 6559 | 0.124 |  |  |  | 1 | 14.1961 | 120.823 |  |  |  | 119.954 | 0 |  | 0.72427 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.9; BGD 29: 32.8; BGD 30: 1069.4 | 2021/05/18 23:13:03 | 0
05/18/21 | 77 | 20210048.121 | 889 | 0.811 | 6565 | 0 |  |  |  | 2 | 14.1962 | 120.824 |  |  |  | 119.955 | 1 |  | 0.72418 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.9; BGD 29: 32.1; BGD 30: 1196.3 | 2021/05/18 23:13:03 | 0
05/18/21 | 77 | 20210048.121 | 889 | 0.811 | 3910 | 11.857 |  |  | Nitrogen | 3 | 12.3453 | 94.151 |  |  |  | 93.466 | 0 |  | 0.73277 |  | 0.0037218 |  | 0.3708 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1166.1 | 2021/05/18 23:13:03 | 0
05/18/21 | 77 | 20210048.121 | 889 | 0.811 |  |  | 819 | -30.548 | Carbon | 4 | 48.71 | 22.201 |  |  | 21.858 |  | 0 | 1.1512 |  | 0.010839 |  | 1.0722 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/18 23:13:03 | 95
05/18/21 | 77 | 20210048.121 | 889 | 0.811 |  |  | 5510 | -43.8 |  | 5 | 53.3602 | 103.021 |  |  | 101.457 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/18 23:13:03 | 95
05/18/21 | 77 | 20210048.121 | 889 | 0.811 |  |  | 5514 | -44.004 |  | 6 | 52.9479 | 103.275 |  |  | 101.707 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/18 23:13:03 | 95
05/18/21 | 78 | 20210048.122 | 891 | 0.801 | 6554 | 0.119 |  |  |  | 1 | 14.3664 | 120.7 |  |  |  | 119.832 | 0 |  | 0.72426 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1069.5 | 2021/05/18 23:22:48 | 0
05/18/21 | 78 | 20210048.122 | 891 | 0.801 | 6557 | 0 |  |  |  | 2 | 14.3746 | 120.844 |  |  |  | 119.976 | 1 |  | 0.72418 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1196.5 | 2021/05/18 23:22:48 | 0
05/18/21 | 78 | 20210048.122 | 891 | 0.801 | 3644 | 12.94 |  |  | Nitrogen | 3 | 11.9448 | 87.8 |  |  |  | 87.161 | 0 |  | 0.73355 |  | 0.0037258 |  | 0.3712 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1164.8 | 2021/05/18 23:22:48 | 0
05/18/21 | 78 | 20210048.122 | 891 | 0.801 |  |  | 820 | -31.594 | Carbon | 4 | 49.3689 | 22.23 |  |  | 21.886 |  | 0 | 1.1501 |  | 0.010827 |  | 1.0711 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/18 23:22:48 | 95
05/18/21 | 78 | 20210048.122 | 891 | 0.801 |  |  | 5514 | -43.8 |  | 5 | 53.8605 | 103.122 |  |  | 101.556 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/18 23:22:48 | 95
05/18/21 | 78 | 20210048.122 | 891 | 0.801 |  |  | 5500 | -44.026 |  | 6 | 53.4782 | 103.354 |  |  | 101.785 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/18 23:22:48 | 95
05/18/21 | 79 | 20210048.123 | 893 | 0.813 | 6562 | 0.116 |  |  |  | 1 | 14.1569 | 120.745 |  |  |  | 119.877 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1069.4 | 2021/05/18 23:32:34 | 0
05/18/21 | 79 | 20210048.123 | 893 | 0.813 | 6559 | 0 |  |  |  | 2 | 14.1624 | 120.844 |  |  |  | 119.975 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1196.6 | 2021/05/18 23:32:34 | 0
05/18/21 | 79 | 20210048.123 | 893 | 0.813 | 3699 | 11.211 |  |  | Nitrogen | 3 | 11.8919 | 89.194 |  |  |  | 88.545 | 0 |  | 0.73239 |  | 0.0037194 |  | 0.37056 | BGD 28: 42.7; BGD 29: 32.6; BGD 30: 1166.3 | 2021/05/18 23:32:34 | 0
05/18/21 | 79 | 20210048.123 | 893 | 0.813 |  |  | 813 | -32.459 | Carbon | 4 | 48.3006 | 22.034 |  |  | 21.694 |  | 0 | 1.1491 |  | 0.010817 |  | 1.0702 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/18 23:32:34 | 95
05/18/21 | 79 | 20210048.123 | 893 | 0.813 |  |  | 5511 | -43.8 |  | 5 | 53.0983 | 103.102 |  |  | 101.537 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/18 23:32:34 | 95
05/18/21 | 79 | 20210048.123 | 893 | 0.813 |  |  | 5504 | -44.02 |  | 6 | 52.9089 | 103.219 |  |  | 101.652 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/18 23:32:34 | 95
05/18/21 | 80 | 20210048.124 | 895 | 0.812 | 6558 | 0.09 |  |  |  | 1 | 14.1692 | 120.654 |  |  |  | 119.786 | 0 |  | 0.72426 |  | 0.0036785 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1069.0 | 2021/05/18 23:42:19 | 0
05/18/21 | 80 | 20210048.124 | 895 | 0.812 | 6556 | 0 |  |  |  | 2 | 14.1781 | 120.813 |  |  |  | 119.944 | 1 |  | 0.72419 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 31.9; BGD 30: 1196.1 | 2021/05/18 23:42:19 | 0
05/18/21 | 80 | 20210048.124 | 895 | 0.812 | 3650 | 12.26 |  |  | Nitrogen | 3 | 11.8139 | 88.147 |  |  |  | 87.506 | 0 |  | 0.73307 |  | 0.0037233 |  | 0.37095 | BGD 28: 42.6; BGD 29: 32.4; BGD 30: 1165.0 | 2021/05/18 23:42:19 | 0
05/18/21 | 80 | 20210048.124 | 895 | 0.812 |  |  | 871 | -32.612 | Carbon | 4 | 51.2429 | 23.726 |  |  | 23.36 |  | 0 | 1.1489 |  | 0.010816 |  | 1.07 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/18 23:42:19 | 95
05/18/21 | 80 | 20210048.124 | 895 | 0.812 |  |  | 5506 | -43.8 |  | 5 | 53.3622 | 102.979 |  |  | 101.416 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/18 23:42:19 | 95
05/18/21 | 80 | 20210048.124 | 895 | 0.812 |  |  | 5502 | -44.014 |  | 6 | 52.9346 | 103.243 |  |  | 101.676 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/18 23:42:19 | 95
05/18/21 | 81 | 20210048.125 | 897 | 0.841 | 6560 | 0.094 |  |  |  | 1 | 13.6877 | 120.784 |  |  |  | 119.916 | 0 |  | 0.72425 |  | 0.0036785 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1068.8 | 2021/05/18 23:52:05 | 0
05/18/21 | 81 | 20210048.125 | 897 | 0.841 | 6550 | 0 |  |  |  | 2 | 13.6853 | 120.74 |  |  |  | 119.872 | 1 |  | 0.72418 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.8; BGD 29: 32.0; BGD 30: 1195.9 | 2021/05/18 23:52:05 | 0
05/18/21 | 81 | 20210048.125 | 897 | 0.841 | 3892 | 11.603 |  |  | Nitrogen | 3 | 11.8641 | 93.643 |  |  |  | 92.962 | 0 |  | 0.73258 |  | 0.0037209 |  | 0.37071 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1166.1 | 2021/05/18 23:52:05 | 0
05/18/21 | 81 | 20210048.125 | 897 | 0.841 |  |  | 848 | -32.28 | Carbon | 4 | 48.3512 | 23.034 |  |  | 22.678 |  | 0 | 1.1492 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.3 | 2021/05/18 23:52:05 | 95
05/18/21 | 81 | 20210048.125 | 897 | 0.841 |  |  | 5504 | -43.8 |  | 5 | 51.5734 | 102.947 |  |  | 101.383 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/18 23:52:05 | 95
05/18/21 | 81 | 20210048.125 | 897 | 0.841 |  |  | 5500 | -44.015 |  | 6 | 51.2021 | 103.184 |  |  | 101.618 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/18 23:52:05 | 95
05/19/21 | 82 | 20210048.126 | 899 | 0.81 | 6552 | 0.107 |  |  |  | 1 | 14.1992 | 120.565 |  |  |  | 119.698 | 0 |  | 0.72424 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.7; BGD 30: 1069.8 | 2021/05/19 00:01:50 | 0
05/19/21 | 82 | 20210048.126 | 899 | 0.81 | 6552 | 0 |  |  |  | 2 | 14.2068 | 120.7 |  |  |  | 119.832 | 1 |  | 0.72417 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 31.9; BGD 30: 1196.6 | 2021/05/19 00:01:50 | 0
05/19/21 | 82 | 20210048.126 | 899 | 0.81 | 3525 | 9.629 |  |  | Nitrogen | 3 | 11.5388 | 84.803 |  |  |  | 84.187 | 0 |  | 0.73114 |  | 0.0037136 |  | 0.36999 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1165.0 | 2021/05/19 00:01:50 | 0
05/19/21 | 82 | 20210048.126 | 899 | 0.81 |  |  | 851 | -32.003 | Carbon | 4 | 50.3688 | 23.132 |  |  | 22.775 |  | 0 | 1.1495 |  | 0.010822 |  | 1.0707 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 00:01:50 | 95
05/19/21 | 82 | 20210048.126 | 899 | 0.81 |  |  | 5514 | -43.8 |  | 5 | 53.2714 | 103.117 |  |  | 101.551 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 00:01:50 | 95
05/19/21 | 82 | 20210048.126 | 899 | 0.81 |  |  | 5509 | -44.018 |  | 6 | 53.0417 | 103.258 |  |  | 101.69 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/19 00:01:50 | 95
05/19/21 | 83 | 20210048.127 | 901 | 0.811 | 6546 | 0.134 |  |  |  | 1 | 14.1877 | 120.672 |  |  |  | 119.804 | 0 |  | 0.72435 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1069.3 | 2021/05/19 00:11:35 | 0
05/19/21 | 83 | 20210048.127 | 901 | 0.811 | 6559 | 0 |  |  |  | 2 | 14.1955 | 120.811 |  |  |  | 119.943 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 32.0; BGD 30: 1195.9 | 2021/05/19 00:11:35 | 0
05/19/21 | 83 | 20210048.127 | 901 | 0.811 | 3097 | 11.444 |  |  | Nitrogen | 3 | 10.5055 | 74.392 |  |  |  | 73.851 | 0 |  | 0.73254 |  | 0.0037203 |  | 0.37065 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1163.4 | 2021/05/19 00:11:35 | 0
05/19/21 | 83 | 20210048.127 | 901 | 0.811 |  |  | 907 | -34.162 | Carbon | 4 | 52.96 | 24.729 |  |  | 24.348 |  | 0 | 1.1471 |  | 0.010798 |  | 1.0683 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/19 00:11:35 | 95
05/19/21 | 83 | 20210048.127 | 901 | 0.811 |  |  | 5504 | -43.8 |  | 5 | 53.429 | 102.979 |  |  | 101.415 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 00:11:35 | 95
05/19/21 | 83 | 20210048.127 | 901 | 0.811 |  |  | 5503 | -44.003 |  | 6 | 53.0692 | 103.201 |  |  | 101.634 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 00:11:35 | 95
05/19/21 | 84 | 20210048.128 | 903 | 0.832 | 6551 | 0.116 |  |  |  | 1 | 13.8272 | 120.628 |  |  |  | 119.761 | 0 |  | 0.72425 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.7; BGD 29: 32.7; BGD 30: 1066.3 | 2021/05/19 00:21:21 | 0
05/19/21 | 84 | 20210048.128 | 903 | 0.832 | 6555 | 0 |  |  |  | 2 | 13.8265 | 120.615 |  |  |  | 119.748 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 31.9; BGD 30: 1193.4 | 2021/05/19 00:21:21 | 0
05/19/21 | 84 | 20210048.128 | 903 | 0.832 | 3624 | 11.885 |  |  | Nitrogen | 3 | 11.4313 | 87.02 |  |  |  | 86.387 | 0 |  | 0.73277 |  | 0.0037219 |  | 0.37081 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1163.1 | 2021/05/19 00:21:21 | 0
05/19/21 | 84 | 20210048.128 | 903 | 0.832 |  |  | 897 | -33.744 | Carbon | 4 | 51.0738 | 24.384 |  |  | 24.008 |  | 0 | 1.1476 |  | 0.010803 |  | 1.0687 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.3 | 2021/05/19 00:21:21 | 95
05/19/21 | 84 | 20210048.128 | 903 | 0.832 |  |  | 5501 | -43.8 |  | 5 | 52.1655 | 102.925 |  |  | 101.362 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 00:21:21 | 95
05/19/21 | 84 | 20210048.128 | 903 | 0.832 |  |  | 5508 | -43.995 |  | 6 | 51.7547 | 103.185 |  |  | 101.618 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/19 00:21:21 | 95
05/19/21 | 85 | 20210048.129 | 905 | 0.8 | 6552 | 0.098 |  |  |  | 1 | 14.3743 | 120.522 |  |  |  | 119.655 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.8; BGD 29: 32.8; BGD 30: 1067.9 | 2021/05/19 00:31:06 | 0
05/19/21 | 85 | 20210048.129 | 905 | 0.8 | 6543 | 0 |  |  |  | 2 | 14.3774 | 120.576 |  |  |  | 119.709 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 31.9; BGD 30: 1194.2 | 2021/05/19 00:31:06 | 0
05/19/21 | 85 | 20210048.129 | 905 | 0.8 | 3140 | 11.456 |  |  | Nitrogen | 3 | 10.7542 | 75.394 |  |  |  | 74.845 | 0 |  | 0.73255 |  | 0.0037203 |  | 0.37066 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1163.0 | 2021/05/19 00:31:06 | 0
05/19/21 | 85 | 20210048.129 | 905 | 0.8 |  |  | 898 | -34.205 | Carbon | 4 | 53.2039 | 24.436 |  |  | 24.06 |  | 0 | 1.1471 |  | 0.010798 |  | 1.0682 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/19 00:31:06 | 95
05/19/21 | 85 | 20210048.129 | 905 | 0.8 |  |  | 5506 | -43.8 |  | 5 | 54.2833 | 102.906 |  |  | 101.343 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 00:31:06 | 95
05/19/21 | 85 | 20210048.129 | 905 | 0.8 |  |  | 5501 | -43.996 |  | 6 | 53.9118 | 103.132 |  |  | 101.566 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/19 00:31:06 | 95
05/19/21 | 86 | 20210048.130 | 907 | 0.846 | 6554 | 0.139 |  |  |  | 1 | 13.6012 | 120.68 |  |  |  | 119.813 | 0 |  | 0.72434 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.7; BGD 29: 32.7; BGD 30: 1066.0 | 2021/05/19 00:40:51 | 0
05/19/21 | 86 | 20210048.130 | 907 | 0.846 | 6554 | 0 |  |  |  | 2 | 13.5999 | 120.656 |  |  |  | 119.788 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 31.9; BGD 30: 1193.1 | 2021/05/19 00:40:51 | 0
05/19/21 | 86 | 20210048.130 | 907 | 0.846 | 3482 | 11.942 |  |  | Nitrogen | 3 | 10.9337 | 83.525 |  |  |  | 82.917 | 0 |  | 0.73288 |  | 0.0037221 |  | 0.37083 | BGD 28: 42.5; BGD 29: 32.4; BGD 30: 1163.0 | 2021/05/19 00:40:51 | 0
05/19/21 | 86 | 20210048.130 | 907 | 0.846 |  |  | 945 | -34.24 | Carbon | 4 | 52.4969 | 25.851 |  |  | 25.453 |  | 0 | 1.1471 |  | 0.010797 |  | 1.0682 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 00:40:51 | 95
05/19/21 | 86 | 20210048.130 | 907 | 0.846 |  |  | 5501 | -43.8 |  | 5 | 51.3777 | 102.876 |  |  | 101.314 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/19 00:40:51 | 95
05/19/21 | 86 | 20210048.130 | 907 | 0.846 |  |  | 5499 | -44.025 |  | 6 | 51.1062 | 103.051 |  |  | 101.487 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 9.7 | 2021/05/19 00:40:51 | 95
05/19/21 | 87 | 20210048.131 | 909 | 0.807 | 6541 | 0.112 |  |  |  | 1 | 14.2524 | 120.572 |  |  |  | 119.705 | 0 |  | 0.72423 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.7; BGD 29: 32.8; BGD 30: 1066.5 | 2021/05/19 00:50:37 | 0
05/19/21 | 87 | 20210048.131 | 909 | 0.807 | 6551 | 0 |  |  |  | 2 | 14.2494 | 120.518 |  |  |  | 119.651 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.7; BGD 29: 32.0; BGD 30: 1193.6 | 2021/05/19 00:50:37 | 0
05/19/21 | 87 | 20210048.131 | 909 | 0.807 | 3815 | 11.722 |  |  | Nitrogen | 3 | 12.2066 | 91.789 |  |  |  | 91.121 | 0 |  | 0.73264 |  | 0.0037213 |  | 0.37075 | BGD 28: 42.6; BGD 29: 32.5; BGD 30: 1163.4 | 2021/05/19 00:50:37 | 0
05/19/21 | 87 | 20210048.131 | 909 | 0.807 |  |  | 831 | -32.495 | Carbon | 4 | 49.554 | 22.548 |  |  | 22.2 |  | 0 | 1.149 |  | 0.010817 |  | 1.0701 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 10.3 | 2021/05/19 00:50:37 | 95
05/19/21 | 87 | 20210048.131 | 909 | 0.807 |  |  | 5498 | -43.8 |  | 5 | 53.8718 | 102.869 |  |  | 101.307 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/19 00:50:37 | 95
05/19/21 | 87 | 20210048.131 | 909 | 0.807 |  |  | 5494 | -44.008 |  | 6 | 53.5847 | 103.046 |  |  | 101.482 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/19 00:50:37 | 95
05/19/21 | 88 | 20210048.132 | 911 | 0.844 | 6539 | 0.102 |  |  |  | 1 | 13.6259 | 120.541 |  |  |  | 119.674 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.6; BGD 29: 32.7; BGD 30: 1068.2 | 2021/05/19 01:00:22 | 0
05/19/21 | 88 | 20210048.132 | 911 | 0.844 | 6549 | 0 |  |  |  | 2 | 13.6222 | 120.472 |  |  |  | 119.606 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.6; BGD 29: 31.9; BGD 30: 1195.0 | 2021/05/19 01:00:22 | 0
05/19/21 | 88 | 20210048.132 | 911 | 0.844 | 3993 | 12.154 |  |  | Nitrogen | 3 | 12.0041 | 95.944 |  |  |  | 95.245 | 0 |  | 0.73304 |  | 0.0037229 |  | 0.37091 | BGD 28: 42.5; BGD 29: 32.4; BGD 30: 1164.6 | 2021/05/19 01:00:22 | 0
05/19/21 | 88 | 20210048.132 | 911 | 0.844 |  |  | 866 | -32.566 | Carbon | 4 | 48.9827 | 23.529 |  |  | 23.166 |  | 0 | 1.1489 |  | 0.010816 |  | 1.07 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/19 01:00:22 | 95
05/19/21 | 88 | 20210048.132 | 911 | 0.844 |  |  | 5504 | -43.8 |  | 5 | 51.5116 | 102.868 |  |  | 101.306 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/19 01:00:22 | 95
05/19/21 | 88 | 20210048.132 | 911 | 0.844 |  |  | 5497 | -44.005 |  | 6 | 51.143 | 103.105 |  |  | 101.54 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/19 01:00:22 | 95
05/19/21 | 89 | 20210048.133 | 913 | 0.815 | 6541 | 0.111 |  |  |  | 1 | 14.0975 | 120.304 |  |  |  | 119.439 | 0 |  | 0.72423 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.7; BGD 29: 32.7; BGD 30: 1070.0 | 2021/05/19 01:10:07 | 0
05/19/21 | 89 | 20210048.133 | 913 | 0.815 | 6546 | 0 |  |  |  | 2 | 14.1067 | 120.469 |  |  |  | 119.602 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.6; BGD 29: 31.9; BGD 30: 1196.7 | 2021/05/19 01:10:07 | 0
05/19/21 | 89 | 20210048.133 | 913 | 0.815 | 3823 | 11.86 |  |  | Nitrogen | 3 | 12.0893 | 91.818 |  |  |  | 91.151 | 0 |  | 0.73274 |  | 0.0037218 |  | 0.3708 | BGD 28: 42.5; BGD 29: 32.4; BGD 30: 1166.5 | 2021/05/19 01:10:07 | 0
05/19/21 | 89 | 20210048.133 | 913 | 0.815 |  |  | 850 | -32.299 | Carbon | 4 | 49.9936 | 23.093 |  |  | 22.736 |  | 0 | 1.1492 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.3 | 2021/05/19 01:10:07 | 95
05/19/21 | 89 | 20210048.133 | 913 | 0.815 |  |  | 5505 | -43.8 |  | 5 | 53.4162 | 102.823 |  |  | 101.262 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/19 01:10:07 | 95
05/19/21 | 89 | 20210048.133 | 913 | 0.815 |  |  | 5488 | -44.008 |  | 6 | 53.0537 | 103.049 |  |  | 101.485 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.8 | 2021/05/19 01:10:07 | 95
05/19/21 | 90 | 20210048.134 | 915 | 0.838 | 6540 | 0.101 |  |  |  | 1 | 13.7152 | 120.388 |  |  |  | 119.523 | 0 |  | 0.72422 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.6; BGD 29: 32.6; BGD 30: 1070.5 | 2021/05/19 01:19:53 | 0
05/19/21 | 90 | 20210048.134 | 915 | 0.838 | 6543 | 0 |  |  |  | 2 | 13.7182 | 120.443 |  |  |  | 119.577 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.6; BGD 29: 31.8; BGD 30: 1197.1 | 2021/05/19 01:19:53 | 0
05/19/21 | 90 | 20210048.134 | 915 | 0.838 | 3783 | 11.953 |  |  | Nitrogen | 3 | 11.6749 | 90.825 |  |  |  | 90.164 | 0 |  | 0.7328 |  | 0.0037222 |  | 0.37084 | BGD 28: 42.5; BGD 29: 32.4; BGD 30: 1167.2 | 2021/05/19 01:19:53 | 0
05/19/21 | 90 | 20210048.134 | 915 | 0.838 |  |  | 887 | -33.38 | Carbon | 4 | 50.3475 | 24.158 |  |  | 23.785 |  | 0 | 1.148 |  | 0.010807 |  | 1.0691 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.3 | 2021/05/19 01:19:53 | 95
05/19/21 | 90 | 20210048.134 | 915 | 0.838 |  |  | 5498 | -43.8 |  | 5 | 51.9718 | 102.81 |  |  | 101.249 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.4 | 2021/05/19 01:19:53 | 95
05/19/21 | 90 | 20210048.134 | 915 | 0.838 |  |  | 5492 | -44.009 |  | 6 | 51.7695 | 102.939 |  |  | 101.376 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.8 | 2021/05/19 01:19:53 | 95
05/19/21 | 91 | 20210048.135 | 917 | 0.822 | 6537 | 0.112 |  |  |  | 1 | 13.9748 | 120.257 |  |  |  | 119.392 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.6; BGD 29: 32.6; BGD 30: 1071.5 | 2021/05/19 01:29:38 | 0
05/19/21 | 91 | 20210048.135 | 917 | 0.822 | 6540 | 0 |  |  |  | 2 | 13.9873 | 120.481 |  |  |  | 119.614 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.6; BGD 29: 31.9; BGD 30: 1197.6 | 2021/05/19 01:29:38 | 0
05/19/21 | 91 | 20210048.135 | 917 | 0.822 | 3981 | 12.022 |  |  | Nitrogen | 3 | 12.2988 | 95.613 |  |  |  | 94.917 | 0 |  | 0.73294 |  | 0.0037224 |  | 0.37086 | BGD 28: 42.5; BGD 29: 32.4; BGD 30: 1167.6 | 2021/05/19 01:29:38 | 0
05/19/21 | 91 | 20210048.135 | 917 | 0.822 |  |  | 833 | -32.137 | Carbon | 4 | 48.7249 | 22.592 |  |  | 22.244 |  | 0 | 1.1493 |  | 0.010821 |  | 1.0705 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.4 | 2021/05/19 01:29:38 | 95
05/19/21 | 91 | 20210048.135 | 917 | 0.822 |  |  | 5493 | -43.8 |  | 5 | 53.1218 | 102.722 |  |  | 101.163 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/19 01:29:38 | 95
05/19/21 | 91 | 20210048.135 | 917 | 0.822 |  |  | 5489 | -44.022 |  | 6 | 52.5338 | 103.092 |  |  | 101.527 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.8 | 2021/05/19 01:29:38 | 95
05/19/21 | 92 | 20210048.136 | 919 | 0.817 | 6538 | 0.123 |  |  |  | 1 | 14.0669 | 120.373 |  |  |  | 119.508 | 0 |  | 0.72431 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.5; BGD 29: 32.6; BGD 30: 1071.6 | 2021/05/19 01:39:23 | 0
05/19/21 | 92 | 20210048.136 | 919 | 0.817 | 6544 | 0 |  |  |  | 2 | 14.0759 | 120.535 |  |  |  | 119.668 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.5; BGD 29: 31.8; BGD 30: 1198.4 | 2021/05/19 01:39:23 | 0
05/19/21 | 92 | 20210048.136 | 919 | 0.817 | 3644 | 11.517 |  |  | Nitrogen | 3 | 11.6918 | 87.586 |  |  |  | 86.949 | 0 |  | 0.73256 |  | 0.0037206 |  | 0.37068 | BGD 28: 42.4; BGD 29: 32.3; BGD 30: 1168.0 | 2021/05/19 01:39:23 | 0
05/19/21 | 92 | 20210048.136 | 919 | 0.817 |  |  | 856 | -33.092 | Carbon | 4 | 50.1252 | 23.244 |  |  | 22.885 |  | 0 | 1.1482 |  | 0.01081 |  | 1.0695 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/19 01:39:23 | 95
05/19/21 | 92 | 20210048.136 | 919 | 0.817 |  |  | 5491 | -43.8 |  | 5 | 53.5461 | 102.66 |  |  | 101.101 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/19 01:39:23 | 95
05/19/21 | 92 | 20210048.136 | 919 | 0.817 |  |  | 5486 | -44.021 |  | 6 | 53.0032 | 103 |  |  | 101.436 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.8 | 2021/05/19 01:39:23 | 95
05/19/21 | 93 | 20210048.136 | Replicate H8 | 0.817 | 6532 | 0.127 |  |  |  | 1 | 14.0616 | 120.279 |  |  |  | 119.414 | 0 |  | 0.72422 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.6; BGD 29: 32.6; BGD 30: 1070.8 | 2021/05/19 01:49:09 | 0
05/19/21 | 93 | 20210048.136 | Replicate H8 | 0.817 | 6536 | 0 |  |  |  | 2 | 14.0642 | 120.325 |  |  |  | 119.46 | 1 |  | 0.72413 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.6; BGD 29: 31.8; BGD 30: 1197.8 | 2021/05/19 01:49:09 | 0
05/19/21 | 93 | 20210048.136 | Replicate H8 | 0.817 | 3677 | 11.473 |  |  | Nitrogen | 3 | 11.764 | 88.4 |  |  |  | 87.758 | 0 |  | 0.73243 |  | 0.0037204 |  | 0.37066 | BGD 28: 42.4; BGD 29: 32.4; BGD 30: 1167.6 | 2021/05/19 01:49:09 | 0
05/19/21 | 93 | 20210048.136 | Replicate H8 | 0.817 |  |  | 857 | -33.199 | Carbon | 4 | 50.2 | 23.289 |  |  | 22.929 |  | 0 | 1.148 |  | 0.010809 |  | 1.0693 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 10.4 | 2021/05/19 01:49:09 | 95
05/19/21 | 93 | 20210048.136 | Replicate H8 | 0.817 |  |  | 5492 | -43.8 |  | 5 | 53.3994 | 102.752 |  |  | 101.192 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.4 | 2021/05/19 01:49:09 | 95
05/19/21 | 93 | 20210048.136 | Replicate H8 | 0.817 |  |  | 5486 | -44.017 |  | 6 | 53.1984 | 102.878 |  |  | 101.316 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.8 | 2021/05/19 01:49:09 | 95
05/19/21 | 94 | 36-UWSIF-UT Glut 1-20210048.211 | 36-UWSIF-UT Glut 1- | 0.744 | 6533 | 0.105 |  |  |  | 1 | 15.4336 | 120.154 |  |  |  | 119.29 | 0 |  | 0.72422 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.6; BGD 29: 32.6; BGD 30: 1071.0 | 2021/05/19 01:58:54 | 0
05/19/21 | 94 | 36-UWSIF-UT Glut 1-20210048.211 | 36-UWSIF-UT Glut 1- | 0.744 | 6535 | 0 |  |  |  | 2 | 15.4461 | 120.357 |  |  |  | 119.492 | 1 |  | 0.72414 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.5; BGD 29: 31.8; BGD 30: 1197.6 | 2021/05/19 01:58:54 | 0
05/19/21 | 94 | 36-UWSIF-UT Glut 1-20210048.211 | 36-UWSIF-UT Glut 1- | 0.744 | 2271 | -2.888 |  |  | Nitrogen | 3 | 9.0473 | 54.902 |  |  |  | 54.508 | 0 |  | 0.72205 |  | 0.0036676 |  | 0.36542 | BGD 28: 42.4; BGD 29: 32.3; BGD 30: 1165.0 | 2021/05/19 01:58:54 | 0
05/19/21 | 94 | 36-UWSIF-UT Glut 1-20210048.211 | 36-UWSIF-UT Glut 1- | 0.744 |  |  | 645 | -29.709 | Carbon | 4 | 43.1676 | 17.241 |  |  | 16.974 |  | 0 | 1.1519 |  | 0.010848 |  | 1.0732 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 01:58:54 | 95
05/19/21 | 94 | 36-UWSIF-UT Glut 1-20210048.211 | 36-UWSIF-UT Glut 1- | 0.744 |  |  | 5496 | -43.8 |  | 5 | 58.8347 | 102.64 |  |  | 101.082 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/19 01:58:54 | 95
05/19/21 | 94 | 36-UWSIF-UT Glut 1-20210048.211 | 36-UWSIF-UT Glut 1- | 0.744 |  |  | 5488 | -44.013 |  | 6 | 58.4009 | 102.888 |  |  | 101.326 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 01:58:54 | 95
05/19/21 | 95 | 36-UWSIF-UT Glut 1-20210048.212 | 36-UWSIF-UT Glut 1- | 0.755 | 6533 | 0.121 |  |  |  | 1 | 15.2109 | 120.188 |  |  |  | 119.324 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36652 | BGD 28: 42.2; BGD 29: 32.3; BGD 30: 1064.2 | 2021/05/19 02:08:40 | 0
05/19/21 | 95 | 36-UWSIF-UT Glut 1-20210048.212 | 36-UWSIF-UT Glut 1- | 0.755 | 6541 | 0 |  |  |  | 2 | 15.2229 | 120.387 |  |  |  | 119.521 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.3; BGD 29: 31.6; BGD 30: 1190.7 | 2021/05/19 02:08:40 | 0
05/19/21 | 95 | 36-UWSIF-UT Glut 1-20210048.212 | 36-UWSIF-UT Glut 1- | 0.755 | 2305 | -2.906 |  |  | Nitrogen | 3 | 9.0132 | 55.643 |  |  |  | 55.244 | 0 |  | 0.72211 |  | 0.0036675 |  | 0.36541 | BGD 28: 42.3; BGD 29: 32.2; BGD 30: 1158.7 | 2021/05/19 02:08:40 | 0
05/19/21 | 95 | 36-UWSIF-UT Glut 1-20210048.212 | 36-UWSIF-UT Glut 1- | 0.755 |  |  | 654 | -29.902 | Carbon | 4 | 43.0198 | 17.473 |  |  | 17.202 |  | 0 | 1.1517 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 10.1 | 2021/05/19 02:08:40 | 95
05/19/21 | 95 | 36-UWSIF-UT Glut 1-20210048.212 | 36-UWSIF-UT Glut 1- | 0.755 |  |  | 5488 | -43.8 |  | 5 | 57.9736 | 102.642 |  |  | 101.084 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.0 | 2021/05/19 02:08:40 | 95
05/19/21 | 95 | 36-UWSIF-UT Glut 1-20210048.212 | 36-UWSIF-UT Glut 1- | 0.755 |  |  | 5486 | -44.015 |  | 6 | 57.5994 | 102.859 |  |  | 101.298 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 02:08:40 | 95
05/19/21 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 39-UWSIF-UW Glut 2- | 0.73 | 6533 | 0.106 |  |  |  | 1 | 15.7315 | 120.184 |  |  |  | 119.319 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.3; BGD 29: 32.3; BGD 30: 1060.4 | 2021/05/19 02:18:25 | 0
05/19/21 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 39-UWSIF-UW Glut 2- | 0.73 | 6530 | 0 |  |  |  | 2 | 15.7354 | 120.246 |  |  |  | 119.382 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.3; BGD 29: 31.5; BGD 30: 1187.1 | 2021/05/19 02:18:25 | 0
05/19/21 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 39-UWSIF-UW Glut 2- | 0.73 | 2418 | 29.767 |  |  | Nitrogen | 3 | 9.7475 | 58.816 |  |  |  | 58.381 | 0 |  | 0.74578 |  | 0.0037877 |  | 0.37734 | BGD 28: 42.1; BGD 29: 32.1; BGD 30: 1155.0 | 2021/05/19 02:18:25 | 0
05/19/21 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 39-UWSIF-UW Glut 2- | 0.73 |  |  | 680 | 22.186 | Carbon | 4 | 46.3965 | 18.369 |  |  | 18.075 |  | 0 | 1.2092 |  | 0.011428 |  | 1.1299 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 10.0 | 2021/05/19 02:18:25 | 95
05/19/21 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 39-UWSIF-UW Glut 2- | 0.73 |  |  | 5485 | -43.8 |  | 5 | 60.0708 | 102.58 |  |  | 101.022 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/19 02:18:25 | 95
05/19/21 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 39-UWSIF-UW Glut 2- | 0.73 |  |  | 5489 | -44 |  | 6 | 59.5577 | 102.867 |  |  | 101.306 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.5 | 2021/05/19 02:18:25 | 95
05/19/21 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 39-UWSIF-UW Glut 2- | 0.768 | 6534 | 0.093 |  |  |  | 1 | 14.9462 | 120.067 |  |  |  | 119.204 | 0 |  | 0.72431 |  | 0.0036785 |  | 0.36651 | BGD 28: 42.2; BGD 29: 32.3; BGD 30: 1057.4 | 2021/05/19 02:28:10 | 0
05/19/21 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 39-UWSIF-UW Glut 2- | 0.768 | 6536 | 0 |  |  |  | 2 | 14.9597 | 120.295 |  |  |  | 119.43 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.2; BGD 29: 31.5; BGD 30: 1184.4 | 2021/05/19 02:28:10 | 0
05/19/21 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 39-UWSIF-UW Glut 2- | 0.768 | 2523 | 29.664 |  |  | Nitrogen | 3 | 9.5503 | 61.105 |  |  |  | 60.653 | 0 |  | 0.74572 |  | 0.0037873 |  | 0.3773 | BGD 28: 42.2; BGD 29: 32.1; BGD 30: 1153.6 | 2021/05/19 02:28:10 | 0
05/19/21 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 39-UWSIF-UW Glut 2- | 0.768 |  |  | 711 | 22.303 | Carbon | 4 | 45.7308 | 19.192 |  |  | 18.885 |  | 0 | 1.2093 |  | 0.01143 |  | 1.13 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 02:28:10 | 95
05/19/21 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 39-UWSIF-UW Glut 2- | 0.768 |  |  | 5486 | -43.8 |  | 5 | 56.961 | 102.661 |  |  | 101.102 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/19 02:28:10 | 95
05/19/21 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 39-UWSIF-UW Glut 2- | 0.768 |  |  | 5483 | -44.028 |  | 6 | 56.6498 | 102.844 |  |  | 101.283 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 02:28:10 | 95
05/19/21 | 98 | 01-UWSIF-Liver-20210048.27 | 01-UWSIF-Liver- | 0.825 | 6524 | 0.115 |  |  |  | 1 | 13.9159 | 120.11 |  |  |  | 119.246 | 0 |  | 0.72421 |  | 0.0036786 |  | 0.36651 | BGD 28: 42.2; BGD 29: 32.4; BGD 30: 1056.5 | 2021/05/19 02:37:55 | 0
05/19/21 | 98 | 01-UWSIF-Liver-20210048.27 | 01-UWSIF-Liver- | 0.825 | 6529 | 0 |  |  |  | 2 | 13.9229 | 120.236 |  |  |  | 119.371 | 1 |  | 0.72413 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.2; BGD 29: 31.5; BGD 30: 1182.9 | 2021/05/19 02:37:55 | 0
05/19/21 | 98 | 01-UWSIF-Liver-20210048.27 | 01-UWSIF-Liver- | 0.825 | 3496 | 8.476 |  |  | Nitrogen | 3 | 11.2451 | 83.884 |  |  |  | 83.276 | 0 |  | 0.73027 |  | 0.0037094 |  | 0.36957 | BGD 28: 42.1; BGD 29: 32.1; BGD 30: 1152.6 | 2021/05/19 02:37:55 | 0
05/19/21 | 98 | 01-UWSIF-Liver-20210048.27 | 01-UWSIF-Liver- | 0.825 |  |  | 877 | -19.405 | Carbon | 4 | 50.6243 | 23.841 |  |  | 23.47 |  | 0 | 1.1632 |  | 0.010963 |  | 1.0844 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 02:37:55 | 95
05/19/21 | 98 | 01-UWSIF-Liver-20210048.27 | 01-UWSIF-Liver- | 0.825 |  |  | 5492 | -43.8 |  | 5 | 52.9868 | 102.685 |  |  | 101.126 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/19 02:37:55 | 95
05/19/21 | 98 | 01-UWSIF-Liver-20210048.27 | 01-UWSIF-Liver- | 0.825 |  |  | 5484 | -44.005 |  | 6 | 52.7011 | 102.866 |  |  | 101.305 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/19 02:37:55 | 95
05/19/21 | 99 | 01-UWSIF-Liver-20210048.28 | 01-UWSIF-Liver- | 0.736 | 6530 | 0.126 |  |  |  | 1 | 15.6067 | 120.238 |  |  |  | 119.374 | 0 |  | 0.72439 |  | 0.0036787 |  | 0.36652 | BGD 28: 42.5; BGD 29: 32.6; BGD 30: 1061.2 | 2021/05/19 02:47:41 | 0
05/19/21 | 99 | 01-UWSIF-Liver-20210048.28 | 01-UWSIF-Liver- | 0.736 | 6533 | 0 |  |  |  | 2 | 15.6108 | 120.305 |  |  |  | 119.44 | 1 |  | 0.7243 |  | 0.0036782 |  | 0.36647 | BGD 28: 41.5; BGD 29: 31.8; BGD 30: 1188.1 | 2021/05/19 02:47:41 | 0
05/19/21 | 99 | 01-UWSIF-Liver-20210048.28 | 01-UWSIF-Liver- | 0.736 | 3105 | 8.488 |  |  | Nitrogen | 3 | 11.6233 | 74.808 |  |  |  | 74.265 | 0 |  | 0.73045 |  | 0.0037094 |  | 0.36957 | BGD 28: 42.3; BGD 29: 32.3; BGD 30: 1156.4 | 2021/05/19 02:47:41 | 0
05/19/21 | 99 | 01-UWSIF-Liver-20210048.28 | 01-UWSIF-Liver- | 0.736 |  |  | 783 | -19.351 | Carbon | 4 | 51.6775 | 21.169 |  |  | 20.839 |  | 0 | 1.1633 |  | 0.010964 |  | 1.0845 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/19 02:47:41 | 95
05/19/21 | 99 | 01-UWSIF-Liver-20210048.28 | 01-UWSIF-Liver- | 0.736 |  |  | 5497 | -43.8 |  | 5 | 59.4964 | 102.628 |  |  | 101.069 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 02:47:41 | 95
05/19/21 | 99 | 01-UWSIF-Liver-20210048.28 | 01-UWSIF-Liver- | 0.736 |  |  | 5484 | -44.024 |  | 6 | 59.2433 | 102.771 |  |  | 101.211 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 02:47:41 | 95

## Original 3
Date | Row | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Component | Peak Nr | Amt% | Area All | Comment | Preparation | Area 44 | Area 28 | Is Ref _ | rR 45CO2/44CO2 | rR 29N2/28N2 | R 13C/12C | R 15N/14N | AT% 13C/12C | AT% 15N/14N | BGD 28 | BGD 29 | BGD 30 | BGD 44 | BGD 45 | BGD 46 | Time Code | Sample Dilution
05/19/21 | 1 | 36-UWSIF-UT Glut 1-20210048.31 | 36-UWSIF-UT Glut 1- | 0.795 | 6471 | 0.119 |  |  |  | 1 | 14.3803 | 119.067 | JOB: 2021-0048 Tray 3 Nelson |  |  | 118.21 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36652 | BGD 28: 38.0; BGD 29: 28.3; BGD 30: 809.9 | 2021/05/19 10:41:07 | 0
05/19/21 | 1 | 36-UWSIF-UT Glut 1-20210048.31 | 36-UWSIF-UT Glut 1- | 0.795 | 6462 | 0 |  |  |  | 2 | 14.3784 | 119.033 | JOB: 2021-0048 Tray 3 Nelson |  |  | 118.177 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 37.1; BGD 29: 27.7; BGD 30: 919.1 | 2021/05/19 10:41:07 | 0
05/19/21 | 1 | 36-UWSIF-UT Glut 1-20210048.31 | 36-UWSIF-UT Glut 1- | 0.795 | 2410 | -2.778 |  |  | Nitrogen | 3 | 8.9016 | 58.414 | JOB: 2021-0048 Tray 3 Nelson |  |  | 57.995 | 0 |  | 0.72222 |  | 0.003668 |  | 0.36546 | BGD 28: 38.1; BGD 29: 28.4; BGD 30: 883.2 | 2021/05/19 10:41:07 | 0
05/19/21 | 1 | 36-UWSIF-UT Glut 1-20210048.31 | 36-UWSIF-UT Glut 1- | 0.795 |  |  | 672 | -29.798 | Carbon | 4 | 42.1726 | 18.147 | JOB: 2021-0048 Tray 3 Nelson |  | 17.866 |  | 0 | 1.1515 |  | 0.010847 |  | 1.0731 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 7.5 | 2021/05/19 10:41:07 | 95
05/19/21 | 1 | 36-UWSIF-UT Glut 1-20210048.31 | 36-UWSIF-UT Glut 1- | 0.795 |  |  | 5428 | -43.8 |  | 5 | 57.1542 | 101.338 | JOB: 2021-0048 Tray 3 Nelson |  | 99.8 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.5; BGD 45: 0.6; BGD 46: 5.5 | 2021/05/19 10:41:07 | 95
05/19/21 | 1 | 36-UWSIF-UT Glut 1-20210048.31 | 36-UWSIF-UT Glut 1- | 0.795 |  |  | 5429 | -44.114 |  | 6 | 56.3847 | 101.822 | JOB: 2021-0048 Tray 3 Nelson |  | 100.277 |  | 0 | 1.1344 |  | 0.010687 |  | 1.0574 |  | BGD 44: 1.3; BGD 45: 1.5; BGD 46: 6.9 | 2021/05/19 10:41:07 | 95
05/19/21 | 2 | 36-UWSIF-UT Glut 1-20210048.32 | 36-UWSIF-UT Glut 1- | 0.72 | 6454 | 0.118 |  |  |  | 1 | 15.8586 | 118.764 | ANALYST: sme |  |  | 117.91 | 0 |  | 0.72437 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.7; BGD 29: 29.9; BGD 30: 812.0 | 2021/05/19 10:50:52 | 0
05/19/21 | 2 | 36-UWSIF-UT Glut 1-20210048.32 | 36-UWSIF-UT Glut 1- | 0.72 | 6459 | 0 |  |  |  | 2 | 15.8681 | 118.91 | ANALYST: sme |  |  | 118.055 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.8; BGD 29: 29.2; BGD 30: 929.2 | 2021/05/19 10:50:52 | 0
05/19/21 | 2 | 36-UWSIF-UT Glut 1-20210048.32 | 36-UWSIF-UT Glut 1- | 0.72 | 2125 | -2.828 |  |  | Nitrogen | 3 | 9.0011 | 52.42 | ANALYST: sme |  |  | 52.044 | 0 |  | 0.72224 |  | 0.0036678 |  | 0.36544 | BGD 28: 39.8; BGD 29: 29.8; BGD 30: 897.8 | 2021/05/19 10:50:52 | 0
05/19/21 | 2 | 36-UWSIF-UT Glut 1-20210048.32 | 36-UWSIF-UT Glut 1- | 0.72 |  |  | 604 | -30.029 | Carbon | 4 | 42.6905 | 16.372 | ANALYST: sme |  | 16.119 |  | 0 | 1.1515 |  | 0.010844 |  | 1.0728 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.2 | 2021/05/19 10:50:52 | 95
05/19/21 | 2 | 36-UWSIF-UT Glut 1-20210048.32 | 36-UWSIF-UT Glut 1- | 0.72 |  |  | 5419 | -43.8 |  | 5 | 63.1162 | 101.333 | ANALYST: sme |  | 99.794 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.5; BGD 45: 0.7; BGD 46: 5.5 | 2021/05/19 10:50:52 | 95
05/19/21 | 2 | 36-UWSIF-UT Glut 1-20210048.32 | 36-UWSIF-UT Glut 1- | 0.72 |  |  | 5414 | -44.055 |  | 6 | 62.6243 | 101.614 | ANALYST: sme |  | 100.071 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.3; BGD 45: 1.6; BGD 46: 6.9 | 2021/05/19 10:50:52 | 95
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 | 6432 | 0.126 |  |  |  | 1 | 9.52 | 118.658 | SAMPLE COUNT: 0207 |  |  | 117.805 | 0 |  | 0.72447 |  | 0.0036787 |  | 0.36652 | BGD 28: 40.5; BGD 29: 30.5; BGD 30: 826.3 | 2021/05/19 11:00:37 | 0
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 | 6451 | 0 |  |  |  | 2 | 9.52 | 118.795 | SAMPLE COUNT: 0207 |  |  | 117.941 | 1 |  | 0.72438 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.5; BGD 29: 29.8; BGD 30: 943.7 | 2021/05/19 11:00:37 | 0
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 | 1249 | -2.831 |  |  | Nitrogen | 3 | 9.52 | 31.047 | SAMPLE COUNT: 0207 |  |  | 30.825 | 0 |  | 0.72233 |  | 0.0036678 |  | 0.36544 | BGD 28: 40.4; BGD 29: 30.4; BGD 30: 910.2 | 2021/05/19 11:00:37 | 0
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 |  |  | 357 | -29.73 | Carbon | 4 | 40.81 | 9.514 | SAMPLE COUNT: 0207 |  | 9.367 |  | 0 | 1.152 |  | 0.010848 |  | 1.0731 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/19 11:00:37 | 95
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 |  |  | 5411 | -43.8 |  | 5 | 40.81 | 101.143 | SAMPLE COUNT: 0207 |  | 99.607 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.5; BGD 45: 0.6; BGD 46: 5.4 | 2021/05/19 11:00:37 | 95
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 |  |  | 5403 | -44.041 |  | 6 | 40.81 | 101.426 | SAMPLE COUNT: 0207 |  | 99.886 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.2; BGD 45: 1.5; BGD 46: 6.8 | 2021/05/19 11:00:37 | 95
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 | 6434 | 0.11 |  |  |  | 1 | 9.52 | 118.477 | HE PSI: 1600 |  |  | 117.625 | 0 |  | 0.72447 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.5; BGD 29: 30.6; BGD 30: 836.6 | 2021/05/19 11:10:22 | 0
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 | 6439 | 0 |  |  |  | 2 | 9.52 | 118.527 | HE PSI: 1600 |  |  | 117.675 | 1 |  | 0.72439 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.5; BGD 29: 29.8; BGD 30: 954.4 | 2021/05/19 11:10:22 | 0
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 | 2223 | -2.847 |  |  | Nitrogen | 3 | 9.52 | 54.634 | HE PSI: 1600 |  |  | 54.242 | 0 |  | 0.72233 |  | 0.0036677 |  | 0.36543 | BGD 28: 40.5; BGD 29: 30.5; BGD 30: 923.8 | 2021/05/19 11:10:22 | 0
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 |  |  | 628 | -29.877 | Carbon | 4 | 40.81 | 17.058 | HE PSI: 1600 |  | 16.794 |  | 0 | 1.1519 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.3 | 2021/05/19 11:10:22 | 95
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 |  |  | 5405 | -43.8 |  | 5 | 40.81 | 101.056 | HE PSI: 1600 |  | 99.522 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 5.6 | 2021/05/19 11:10:22 | 95
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 |  |  | 5402 | -44.057 |  | 6 | 40.81 | 101.3 | HE PSI: 1600 |  | 99.762 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.3; BGD 45: 1.6; BGD 46: 7.0 | 2021/05/19 11:10:22 | 95
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 | 6427 | 0.115 |  |  |  | 1 | 9.52 | 118.3 | O2 PSI: 400 |  |  | 117.449 | 0 |  | 0.72437 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.7; BGD 29: 30.8; BGD 30: 853.4 | 2021/05/19 11:20:07 | 0
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 | 6427 | 0 |  |  |  | 2 | 9.52 | 118.354 | O2 PSI: 400 |  |  | 117.503 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.8; BGD 29: 30.1; BGD 30: 971.3 | 2021/05/19 11:20:07 | 0
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 | 4785 | -2.893 |  |  | Nitrogen | 3 | 9.52 | 115.219 | O2 PSI: 400 |  |  | 114.393 | 0 |  | 0.7222 |  | 0.0036676 |  | 0.36542 | BGD 28: 40.8; BGD 29: 30.7; BGD 30: 945.7 | 2021/05/19 11:20:07 | 0
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 |  |  | 1297 | -30.115 | Carbon | 4 | 40.81 | 36.875 | O2 PSI: 400 |  | 36.306 |  | 0 | 1.1515 |  | 0.010844 |  | 1.0727 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 7.7 | 2021/05/19 11:20:07 | 95
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 |  |  | 5397 | -43.8 |  | 5 | 40.81 | 100.966 | O2 PSI: 400 |  | 99.433 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 6.3 | 2021/05/19 11:20:07 | 95
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 |  |  | 5390 | -43.962 |  | 6 | 40.81 | 101.163 | O2 PSI: 400 |  | 99.627 |  | 0 | 1.1352 |  | 0.010689 |  | 1.0576 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 7.5 | 2021/05/19 11:20:07 | 95
05/19/21 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 39-UWSIF-UW Glut 2- | 0.8 | 6425 | 0.123 |  |  |  | 1 | 14.8943 | 118.267 | MASS 28: |  |  | 117.416 | 0 |  | 0.72439 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.4; BGD 29: 31.6; BGD 30: 880.5 | 2021/05/19 11:29:52 | 0
05/19/21 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 39-UWSIF-UW Glut 2- | 0.8 | 6415 | 0 |  |  |  | 2 | 14.8936 | 118.257 | MASS 28: |  |  | 117.407 | 1 |  | 0.7243 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.3; BGD 29: 30.7; BGD 30: 998.0 | 2021/05/19 11:29:52 | 0
05/19/21 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 39-UWSIF-UW Glut 2- | 0.8 | 2580 | 29.724 |  |  | Nitrogen | 3 | 9.6909 | 63.315 | MASS 28: |  |  | 62.846 | 0 |  | 0.74583 |  | 0.0037875 |  | 0.37732 | BGD 28: 41.1; BGD 29: 31.2; BGD 30: 967.6 | 2021/05/19 11:29:52 | 0
05/19/21 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 39-UWSIF-UW Glut 2- | 0.8 |  |  | 728 | 22.144 | Carbon | 4 | 46.6632 | 19.899 | MASS 28: |  | 19.581 |  | 0 | 1.2094 |  | 0.011428 |  | 1.1299 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/19 11:29:52 | 95
05/19/21 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 39-UWSIF-UW Glut 2- | 0.8 |  |  | 5402 | -43.8 |  | 5 | 53.1915 | 100.995 | MASS 28: |  | 99.462 |  | 1 | 1.1355 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.0 | 2021/05/19 11:29:52 | 95
05/19/21 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 39-UWSIF-UW Glut 2- | 0.8 |  |  | 5398 | -44.014 |  | 6 | 52.7267 | 101.262 | MASS 28: |  | 99.725 |  | 0 | 1.1352 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 7.4 | 2021/05/19 11:29:52 | 95
05/19/21 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 39-UWSIF-UW Glut 2- | 0.829 | 6423 | 0.117 |  |  |  | 1 | 14.3757 | 118.304 | MASS 29: |  |  | 117.453 | 0 |  | 0.72439 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 889.1 | 2021/05/19 11:39:36 | 0
05/19/21 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 39-UWSIF-UW Glut 2- | 0.829 | 6431 | 0 |  |  |  | 2 | 14.3791 | 118.356 | MASS 29: |  |  | 117.505 | 1 |  | 0.7243 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.4; BGD 30: 1007.3 | 2021/05/19 11:39:36 | 0
05/19/21 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 39-UWSIF-UW Glut 2- | 0.829 | 2680 | 29.762 |  |  | Nitrogen | 3 | 9.63 | 65.705 | MASS 29: |  |  | 65.218 | 0 |  | 0.74586 |  | 0.0037877 |  | 0.37734 | BGD 28: 41.0; BGD 29: 31.0; BGD 30: 977.2 | 2021/05/19 11:39:36 | 0
05/19/21 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 39-UWSIF-UW Glut 2- | 0.829 |  |  | 755 | 22.232 | Carbon | 4 | 46.3801 | 20.642 | MASS 29: |  | 20.312 |  | 0 | 1.2094 |  | 0.011429 |  | 1.13 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/19 11:39:36 | 95
05/19/21 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 39-UWSIF-UW Glut 2- | 0.829 |  |  | 5403 | -43.8 |  | 5 | 51.2441 | 101.047 | MASS 29: |  | 99.513 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.2 | 2021/05/19 11:39:36 | 95
05/19/21 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 39-UWSIF-UW Glut 2- | 0.829 |  |  | 5396 | -44.032 |  | 6 | 50.9214 | 101.239 | MASS 29: |  | 99.702 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 7.5 | 2021/05/19 11:39:36 | 95
05/19/21 | 8 | 01-UWSIF-Liver-20210048.31 | 01-UWSIF-Liver- | 0.811 | 6428 | 0.106 |  |  |  | 1 | 14.6951 | 118.31 | MASS 30: |  |  | 117.459 | 0 |  | 0.72438 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 899.4 | 2021/05/19 11:49:21 | 0
05/19/21 | 8 | 01-UWSIF-Liver-20210048.31 | 01-UWSIF-Liver- | 0.811 | 6423 | 0 |  |  |  | 2 | 14.692 | 118.263 | MASS 30: |  |  | 117.412 | 1 |  | 0.7243 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.4; BGD 30: 1017.5 | 2021/05/19 11:49:21 | 0
05/19/21 | 8 | 01-UWSIF-Liver-20210048.31 | 01-UWSIF-Liver- | 0.811 | 3340 | 8.509 |  |  | Nitrogen | 3 | 11.5789 | 81.455 | MASS 30: |  |  | 80.864 | 0 |  | 0.73046 |  | 0.0037095 |  | 0.36958 | BGD 28: 40.9; BGD 29: 31.0; BGD 30: 988.1 | 2021/05/19 11:49:21 | 0
05/19/21 | 8 | 01-UWSIF-Liver-20210048.31 | 01-UWSIF-Liver- | 0.811 |  |  | 844 | -19.515 | Carbon | 4 | 51.8245 | 23.115 | MASS 30: |  | 22.755 |  | 0 | 1.1632 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 8.0 | 2021/05/19 11:49:21 | 95
05/19/21 | 8 | 01-UWSIF-Liver-20210048.31 | 01-UWSIF-Liver- | 0.811 |  |  | 5403 | -43.8 |  | 5 | 52.5086 | 100.973 | MASS 30: |  | 99.44 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.3 | 2021/05/19 11:49:21 | 95
05/19/21 | 8 | 01-UWSIF-Liver-20210048.31 | 01-UWSIF-Liver- | 0.811 |  |  | 5396 | -44.016 |  | 6 | 52.1084 | 101.206 | MASS 30: |  | 99.669 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 7.6 | 2021/05/19 11:49:21 | 95
05/19/21 | 9 | 01-UWSIF-Liver-20210048.32 | 01-UWSIF-Liver- | 0.719 | 6417 | 0.108 |  |  |  | 1 | 16.5605 | 118.111 |  |  |  | 117.262 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.3; BGD 30: 911.0 | 2021/05/19 11:59:06 | 0
05/19/21 | 9 | 01-UWSIF-Liver-20210048.32 | 01-UWSIF-Liver- | 0.719 | 6414 | 0 |  |  |  | 2 | 16.5649 | 118.17 |  |  |  | 117.32 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.5; BGD 30: 1029.6 | 2021/05/19 11:59:06 | 0
05/19/21 | 9 | 01-UWSIF-Liver-20210048.32 | 01-UWSIF-Liver- | 0.719 | 2940 | 8.452 |  |  | Nitrogen | 3 | 11.9001 | 71.856 |  |  |  | 71.335 | 0 |  | 0.73033 |  | 0.0037093 |  | 0.36956 | BGD 28: 41.0; BGD 29: 31.1; BGD 30: 1000.1 | 2021/05/19 11:59:06 | 0
05/19/21 | 9 | 01-UWSIF-Liver-20210048.32 | 01-UWSIF-Liver- | 0.719 |  |  | 747 | -19.55 | Carbon | 4 | 52.8517 | 20.342 |  |  | 20.026 |  | 0 | 1.1633 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.0 | 2021/05/19 11:59:06 | 95
05/19/21 | 9 | 01-UWSIF-Liver-20210048.32 | 01-UWSIF-Liver- | 0.719 |  |  | 5396 | -43.8 |  | 5 | 59.3247 | 100.923 |  |  | 99.39 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.3 | 2021/05/19 11:59:06 | 95
05/19/21 | 9 | 01-UWSIF-Liver-20210048.32 | 01-UWSIF-Liver- | 0.719 |  |  | 5392 | -44.015 |  | 6 | 58.8258 | 101.18 |  |  | 99.644 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 7.6 | 2021/05/19 11:59:06 | 95
05/19/21 | 10 | 20210048.137 | 921 | 0.846 | 6409 | 0.122 |  |  |  | 1 | 14.0707 | 118.053 |  |  |  | 117.204 | 0 |  | 0.7244 |  | 0.0036786 |  | 0.36652 | BGD 28: 40.9; BGD 29: 31.1; BGD 30: 919.8 | 2021/05/19 12:08:51 | 0
05/19/21 | 10 | 20210048.137 | 921 | 0.846 | 6419 | 0 |  |  |  | 2 | 14.0805 | 118.206 |  |  |  | 117.356 | 1 |  | 0.72431 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.9; BGD 29: 30.4; BGD 30: 1038.8 | 2021/05/19 12:08:51 | 0
05/19/21 | 10 | 20210048.137 | 921 | 0.846 | 3401 | 11.977 |  |  | Nitrogen | 3 | 11.2392 | 82.883 |  |  |  | 82.28 | 0 |  | 0.73298 |  | 0.0037223 |  | 0.37084 | BGD 28: 40.9; BGD 29: 31.0; BGD 30: 1010.8 | 2021/05/19 12:08:51 | 0
05/19/21 | 10 | 20210048.137 | 921 | 0.846 |  |  | 716 | -31.01 | Carbon | 4 | 43.3316 | 19.459 |  |  | 19.159 |  | 0 | 1.1506 |  | 0.010833 |  | 1.0717 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.2 | 2021/05/19 12:08:51 | 95
05/19/21 | 10 | 20210048.137 | 921 | 0.846 |  |  | 5400 | -43.8 |  | 5 | 50.5688 | 100.831 |  |  | 99.3 |  | 1 | 1.1354 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.4 | 2021/05/19 12:08:51 | 95
05/19/21 | 10 | 20210048.137 | 921 | 0.846 |  |  | 5389 | -44.031 |  | 6 | 50.1734 | 101.072 |  |  | 99.537 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 7.7 | 2021/05/19 12:08:51 | 95
05/19/21 | 11 | 20210048.138 | 923 | 0.842 | 6411 | 0.065 |  |  |  | 1 | 14.1461 | 118.186 |  |  |  | 117.336 | 0 |  | 0.72446 |  | 0.0036784 |  | 0.3665 | BGD 28: 41.2; BGD 29: 31.4; BGD 30: 932.2 | 2021/05/19 12:18:36 | 0
05/19/21 | 11 | 20210048.138 | 923 | 0.842 | 6414 | 0 |  |  |  | 2 | 14.1488 | 118.227 |  |  |  | 117.377 | 1 |  | 0.72442 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.6; BGD 30: 1051.6 | 2021/05/19 12:18:36 | 0
05/19/21 | 11 | 20210048.138 | 923 | 0.842 | 3675 | 12.085 |  |  | Nitrogen | 3 | 11.8668 | 88.957 |  |  |  | 88.309 | 0 |  | 0.73317 |  | 0.0037227 |  | 0.37088 | BGD 28: 41.1; BGD 29: 31.2; BGD 30: 1024.1 | 2021/05/19 12:18:36 | 0
05/19/21 | 11 | 20210048.138 | 923 | 0.842 |  |  | 884 | -33.528 | Carbon | 4 | 51.7641 | 24.239 |  |  | 23.865 |  | 0 | 1.1478 |  | 0.010805 |  | 1.069 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 8.3 | 2021/05/19 12:18:36 | 95
05/19/21 | 11 | 20210048.138 | 923 | 0.842 |  |  | 5392 | -43.8 |  | 5 | 50.6836 | 100.907 |  |  | 99.375 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.6 | 2021/05/19 12:18:36 | 95
05/19/21 | 11 | 20210048.138 | 923 | 0.842 |  |  | 5384 | -44.011 |  | 6 | 50.5601 | 100.982 |  |  | 99.449 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 7.9 | 2021/05/19 12:18:36 | 95
05/19/21 | 12 | 20210048.138 | Replicate A11 | 0.829 | 6418 | 0.097 |  |  |  | 1 | 14.3658 | 118.153 |  |  |  | 117.303 | 0 |  | 0.72438 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.5; BGD 30: 944.1 | 2021/05/19 12:28:21 | 0
05/19/21 | 12 | 20210048.138 | Replicate A11 | 0.829 | 6411 | 0 |  |  |  | 2 | 14.3729 | 118.261 |  |  |  | 117.41 | 1 |  | 0.7243 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.8; BGD 30: 1064.3 | 2021/05/19 12:28:21 | 0
05/19/21 | 12 | 20210048.138 | Replicate A11 | 0.829 | 3366 | 12.112 |  |  | Nitrogen | 3 | 11.3466 | 81.646 |  |  |  | 81.052 | 0 |  | 0.73308 |  | 0.0037228 |  | 0.37089 | BGD 28: 41.3; BGD 29: 31.3; BGD 30: 1034.8 | 2021/05/19 12:28:21 | 0
05/19/21 | 12 | 20210048.138 | Replicate A11 | 0.829 |  |  | 887 | -34.148 | Carbon | 4 | 52.7432 | 24.341 |  |  | 23.966 |  | 0 | 1.1471 |  | 0.010798 |  | 1.0683 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.5 | 2021/05/19 12:28:21 | 95
05/19/21 | 12 | 20210048.138 | Replicate A11 | 0.829 |  |  | 5395 | -43.8 |  | 5 | 51.5869 | 100.842 |  |  | 99.311 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/19 12:28:21 | 95
05/19/21 | 12 | 20210048.138 | Replicate A11 | 0.829 |  |  | 5378 | -44.012 |  | 6 | 51.3232 | 101 |  |  | 99.467 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.0 | 2021/05/19 12:28:21 | 95
05/19/21 | 13 | 20210048.139 | 925 | 0.801 | 6411 | 0.114 |  |  |  | 1 | 14.862 | 118.064 |  |  |  | 117.215 | 0 |  | 0.72437 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.6; BGD 30: 954.1 | 2021/05/19 12:38:06 | 0
05/19/21 | 13 | 20210048.139 | 925 | 0.801 | 6410 | 0 |  |  |  | 2 | 14.8658 | 118.121 |  |  |  | 117.272 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.8; BGD 30: 1074.6 | 2021/05/19 12:38:06 | 0
05/19/21 | 13 | 20210048.139 | 925 | 0.801 | 3166 | 11.553 |  |  | Nitrogen | 3 | 11.2179 | 76.691 |  |  |  | 76.133 | 0 |  | 0.73266 |  | 0.0037207 |  | 0.37069 | BGD 28: 41.4; BGD 29: 31.4; BGD 30: 1045.3 | 2021/05/19 12:38:06 | 0
05/19/21 | 13 | 20210048.139 | 925 | 0.801 |  |  | 852 | -33.665 | Carbon | 4 | 52.8411 | 23.326 |  |  | 22.967 |  | 0 | 1.1477 |  | 0.010804 |  | 1.0688 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.5 | 2021/05/19 12:38:06 | 95
05/19/21 | 13 | 20210048.139 | 925 | 0.801 |  |  | 5394 | -43.8 |  | 5 | 53.4281 | 100.82 |  |  | 99.29 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/19 12:38:06 | 95
05/19/21 | 13 | 20210048.139 | 925 | 0.801 |  |  | 5389 | -44.014 |  | 6 | 52.9832 | 101.077 |  |  | 99.543 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.1 | 2021/05/19 12:38:06 | 95
05/19/21 | 14 | 20210048.140 | 927 | 0.839 | 6415 | 0.127 |  |  |  | 1 | 14.1899 | 118.08 |  |  |  | 117.231 | 0 |  | 0.72438 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.5; BGD 29: 31.7; BGD 30: 961.0 | 2021/05/19 12:47:51 | 0
05/19/21 | 14 | 20210048.140 | 927 | 0.839 | 6410 | 0 |  |  |  | 2 | 14.192 | 118.113 |  |  |  | 117.263 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.5; BGD 29: 30.9; BGD 30: 1081.3 | 2021/05/19 12:47:51 | 0
05/19/21 | 14 | 20210048.140 | 927 | 0.839 | 3573 | 12.023 |  |  | Nitrogen | 3 | 11.6777 | 86.472 |  |  |  | 85.843 | 0 |  | 0.733 |  | 0.0037224 |  | 0.37086 | BGD 28: 41.3; BGD 29: 31.4; BGD 30: 1053.1 | 2021/05/19 12:47:51 | 0
05/19/21 | 14 | 20210048.140 | 927 | 0.839 |  |  | 879 | -33.66 | Carbon | 4 | 51.7163 | 24.096 |  |  | 23.724 |  | 0 | 1.1477 |  | 0.010804 |  | 1.0688 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.6 | 2021/05/19 12:47:51 | 95
05/19/21 | 14 | 20210048.140 | 927 | 0.839 |  |  | 5393 | -43.8 |  | 5 | 50.869 | 100.905 |  |  | 99.373 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/19 12:47:51 | 95
05/19/21 | 14 | 20210048.140 | 927 | 0.839 |  |  | 5388 | -44.039 |  | 6 | 50.4896 | 101.134 |  |  | 99.598 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.2 | 2021/05/19 12:47:51 | 95
05/19/21 | 15 | 20210048.141 | 929 | 0.843 | 6409 | 0.098 |  |  |  | 1 | 14.1225 | 118.079 |  |  |  | 117.23 | 0 |  | 0.72428 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.5; BGD 29: 31.7; BGD 30: 967.4 | 2021/05/19 12:57:36 | 0
05/19/21 | 15 | 20210048.141 | 929 | 0.843 | 6426 | 0 |  |  |  | 2 | 14.128 | 118.165 |  |  |  | 117.315 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.5; BGD 29: 30.9; BGD 30: 1088.2 | 2021/05/19 12:57:36 | 0
05/19/21 | 15 | 20210048.141 | 929 | 0.843 | 3486 | 11.08 |  |  | Nitrogen | 3 | 11.422 | 84.362 |  |  |  | 83.749 | 0 |  | 0.73223 |  | 0.003719 |  | 0.37052 | BGD 28: 41.4; BGD 29: 31.4; BGD 30: 1058.5 | 2021/05/19 12:57:36 | 0
05/19/21 | 15 | 20210048.141 | 929 | 0.843 |  |  | 902 | -33.344 | Carbon | 4 | 52.4275 | 24.69 |  |  | 24.309 |  | 0 | 1.148 |  | 0.010807 |  | 1.0692 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.7 | 2021/05/19 12:57:36 | 95
05/19/21 | 15 | 20210048.141 | 929 | 0.843 |  |  | 5398 | -43.8 |  | 5 | 50.6094 | 100.916 |  |  | 99.383 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/19 12:57:36 | 95
05/19/21 | 15 | 20210048.141 | 929 | 0.843 |  |  | 5391 | -44.003 |  | 6 | 50.2478 | 101.135 |  |  | 99.6 |  | 0 | 1.1351 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.3 | 2021/05/19 12:57:36 | 95
05/19/21 | 16 | 20210048.142 | 931 | 0.812 | 6407 | 0.124 |  |  |  | 1 | 14.6612 | 118.072 |  |  |  | 117.223 | 0 |  | 0.72428 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.4; BGD 29: 31.7; BGD 30: 972.3 | 2021/05/19 13:07:21 | 0
05/19/21 | 16 | 20210048.142 | 931 | 0.812 | 6416 | 0 |  |  |  | 2 | 14.6598 | 118.051 |  |  |  | 117.202 | 1 |  | 0.72419 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.5; BGD 29: 30.9; BGD 30: 1093.4 | 2021/05/19 13:07:21 | 0
05/19/21 | 16 | 20210048.142 | 931 | 0.812 | 3675 | 10.748 |  |  | Nitrogen | 3 | 12.3239 | 89.154 |  |  |  | 88.506 | 0 |  | 0.73197 |  | 0.0037177 |  | 0.3704 | BGD 28: 41.3; BGD 29: 31.4; BGD 30: 1062.5 | 2021/05/19 13:07:21 | 0
05/19/21 | 16 | 20210048.142 | 931 | 0.812 |  |  | 799 | -30.192 | Carbon | 4 | 49.392 | 21.768 |  |  | 21.431 |  | 0 | 1.1515 |  | 0.010843 |  | 1.0726 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.7 | 2021/05/19 13:07:21 | 95
05/19/21 | 16 | 20210048.142 | 931 | 0.812 |  |  | 5396 | -43.8 |  | 5 | 52.5587 | 100.906 |  |  | 99.374 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/19 13:07:21 | 95
05/19/21 | 16 | 20210048.142 | 931 | 0.812 |  |  | 5400 | -44.02 |  | 6 | 52.1709 | 101.132 |  |  | 99.597 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.3 | 2021/05/19 13:07:21 | 95
05/19/21 | 17 | 20210048.143 | 933 | 0.83 | 6404 | 0.109 |  |  |  | 1 | 14.3373 | 117.982 |  |  |  | 117.133 | 0 |  | 0.72438 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.3; BGD 29: 31.5; BGD 30: 976.8 | 2021/05/19 13:17:06 | 0
05/19/21 | 17 | 20210048.143 | 933 | 0.83 | 6417 | 0 |  |  |  | 2 | 14.3502 | 118.179 |  |  |  | 117.329 | 1 |  | 0.7243 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.3; BGD 29: 30.8; BGD 30: 1097.2 | 2021/05/19 13:17:06 | 0
05/19/21 | 17 | 20210048.143 | 933 | 0.83 | 3740 | 9.55 |  |  | Nitrogen | 3 | 12.1802 | 90.494 |  |  |  | 89.837 | 0 |  | 0.73122 |  | 0.0037133 |  | 0.36996 | BGD 28: 41.3; BGD 29: 31.4; BGD 30: 1067.4 | 2021/05/19 13:17:06 | 0
05/19/21 | 17 | 20210048.143 | 933 | 0.83 |  |  | 849 | -32.288 | Carbon | 4 | 50.737 | 23.173 |  |  | 22.815 |  | 0 | 1.1492 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.7 | 2021/05/19 13:17:06 | 95
05/19/21 | 17 | 20210048.143 | 933 | 0.83 |  |  | 5401 | -43.8 |  | 5 | 51.2429 | 101.011 |  |  | 99.477 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/19 13:17:06 | 95
05/19/21 | 17 | 20210048.143 | 933 | 0.83 |  |  | 5398 | -44.008 |  | 6 | 50.9678 | 101.175 |  |  | 99.639 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.3 | 2021/05/19 13:17:06 | 95
05/19/21 | 18 | 20210048.144 | 935 | 0.825 | 6415 | 0.124 |  |  |  | 1 | 14.4225 | 117.955 |  |  |  | 117.107 | 0 |  | 0.72428 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.4; BGD 29: 31.6; BGD 30: 980.2 | 2021/05/19 13:26:51 | 0
05/19/21 | 18 | 20210048.144 | 935 | 0.825 | 6420 | 0 |  |  |  | 2 | 14.4363 | 118.166 |  |  |  | 117.316 | 1 |  | 0.72419 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.3; BGD 29: 30.8; BGD 30: 1100.9 | 2021/05/19 13:26:51 | 0
05/19/21 | 18 | 20210048.144 | 935 | 0.825 | 3828 | 11.789 |  |  | Nitrogen | 3 | 12.4702 | 92.87 |  |  |  | 92.194 | 0 |  | 0.73272 |  | 0.0037216 |  | 0.37078 | BGD 28: 41.2; BGD 29: 31.4; BGD 30: 1071.3 | 2021/05/19 13:26:51 | 0
05/19/21 | 18 | 20210048.144 | 935 | 0.825 |  |  | 827 | -32.141 | Carbon | 4 | 50.1334 | 22.64 |  |  | 22.291 |  | 0 | 1.1494 |  | 0.010821 |  | 1.0705 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/19 13:26:51 | 95
05/19/21 | 18 | 20210048.144 | 935 | 0.825 |  |  | 5394 | -43.8 |  | 5 | 51.7255 | 100.909 |  |  | 99.377 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.0 | 2021/05/19 13:26:51 | 95
05/19/21 | 18 | 20210048.144 | 935 | 0.825 |  |  | 5394 | -44.021 |  | 6 | 51.2986 | 101.162 |  |  | 99.626 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/19 13:26:51 | 95
05/19/21 | 19 | 20210048.145 | 937 | 0.84 | 6402 | 0.101 |  |  |  | 1 | 14.1739 | 118.094 |  |  |  | 117.244 | 0 |  | 0.72446 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.6; BGD 30: 983.5 | 2021/05/19 13:36:36 | 0
05/19/21 | 19 | 20210048.145 | 937 | 0.84 | 6416 | 0 |  |  |  | 2 | 14.1719 | 118.063 |  |  |  | 117.214 | 1 |  | 0.72439 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.8; BGD 30: 1104.4 | 2021/05/19 13:36:36 | 0
05/19/21 | 19 | 20210048.145 | 937 | 0.84 | 2573 | 12.043 |  |  | Nitrogen | 3 | 9.107 | 62.264 |  |  |  | 61.811 | 0 |  | 0.73311 |  | 0.0037225 |  | 0.37087 | BGD 28: 41.2; BGD 29: 31.3; BGD 30: 1070.0 | 2021/05/19 13:36:36 | 0
05/19/21 | 19 | 20210048.145 | 937 | 0.84 |  |  | 994 | -35.9 | Carbon | 4 | 56.9059 | 27.468 |  |  | 27.046 |  | 0 | 1.1451 |  | 0.010779 |  | 1.0664 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 8.7 | 2021/05/19 13:36:36 | 95
05/19/21 | 19 | 20210048.145 | 937 | 0.84 |  |  | 5397 | -43.8 |  | 5 | 50.8994 | 100.85 |  |  | 99.318 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.1 | 2021/05/19 13:36:36 | 95
05/19/21 | 19 | 20210048.145 | 937 | 0.84 |  |  | 5384 | -43.992 |  | 6 | 50.6211 | 101.018 |  |  | 99.485 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/19 13:36:36 | 95
05/19/21 | 20 | 20210048.146 | 939 | 0.832 | 6418 | 0.112 |  |  |  | 1 | 14.3082 | 118.064 |  |  |  | 117.215 | 0 |  | 0.72439 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.5; BGD 29: 31.7; BGD 30: 982.6 | 2021/05/19 13:46:21 | 0
05/19/21 | 20 | 20210048.146 | 939 | 0.832 | 6410 | 0 |  |  |  | 2 | 14.3104 | 118.098 |  |  |  | 117.248 | 1 |  | 0.72431 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.9; BGD 30: 1103.8 | 2021/05/19 13:46:21 | 0
05/19/21 | 20 | 20210048.146 | 939 | 0.832 | 3539 | 11.959 |  |  | Nitrogen | 3 | 11.6937 | 85.611 |  |  |  | 84.988 | 0 |  | 0.73297 |  | 0.0037222 |  | 0.37084 | BGD 28: 41.3; BGD 29: 31.5; BGD 30: 1072.9 | 2021/05/19 13:46:21 | 0
05/19/21 | 20 | 20210048.146 | 939 | 0.832 |  |  | 876 | -33.604 | Carbon | 4 | 51.8605 | 23.92 |  |  | 23.551 |  | 0 | 1.1477 |  | 0.010804 |  | 1.0689 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/19 13:46:21 | 95
05/19/21 | 20 | 20210048.146 | 939 | 0.832 |  |  | 5402 | -43.8 |  | 5 | 51.2088 | 100.958 |  |  | 99.425 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/19 13:46:21 | 95
05/19/21 | 20 | 20210048.146 | 939 | 0.832 |  |  | 5391 | -44.002 |  | 6 | 50.8568 | 101.168 |  |  | 99.632 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/19 13:46:21 | 95
05/19/21 | 21 | 20210048.147 | 941 | 0.814 | 6405 | 0.111 |  |  |  | 1 | 14.6095 | 117.837 |  |  |  | 116.99 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.6; BGD 30: 984.8 | 2021/05/19 13:56:06 | 0
05/19/21 | 21 | 20210048.147 | 941 | 0.814 | 6399 | 0 |  |  |  | 2 | 14.6193 | 117.984 |  |  |  | 117.135 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.8; BGD 30: 1105.3 | 2021/05/19 13:56:06 | 0
05/19/21 | 21 | 20210048.147 | 941 | 0.814 | 3149 | 12.63 |  |  | Nitrogen | 3 | 10.9866 | 76.204 |  |  |  | 75.65 | 0 |  | 0.73335 |  | 0.0037247 |  | 0.37108 | BGD 28: 41.3; BGD 29: 31.4; BGD 30: 1073.4 | 2021/05/19 13:56:06 | 0
05/19/21 | 21 | 20210048.147 | 941 | 0.814 |  |  | 909 | -34.17 | Carbon | 4 | 54.5721 | 24.858 |  |  | 24.474 |  | 0 | 1.1471 |  | 0.010798 |  | 1.0683 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.8 | 2021/05/19 13:56:06 | 95
05/19/21 | 21 | 20210048.147 | 941 | 0.814 |  |  | 5401 | -43.8 |  | 5 | 52.38 | 100.935 |  |  | 99.402 |  | 1 | 1.1353 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.1 | 2021/05/19 13:56:06 | 95
05/19/21 | 21 | 20210048.147 | 941 | 0.814 |  |  | 5393 | -44.011 |  | 6 | 51.9218 | 101.203 |  |  | 99.666 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/19 13:56:06 | 95
05/19/21 | 22 | 20210048.148 | 943 | 0.805 | 6402 | 0.097 |  |  |  | 1 | 14.784 | 118.003 |  |  |  | 117.155 | 0 |  | 0.72437 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.6; BGD 30: 985.6 | 2021/05/19 14:05:51 | 0
05/19/21 | 22 | 20210048.148 | 943 | 0.805 | 6410 | 0 |  |  |  | 2 | 14.7904 | 118.098 |  |  |  | 117.249 | 1 |  | 0.7243 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.9; BGD 30: 1106.6 | 2021/05/19 14:05:51 | 0
05/19/21 | 22 | 20210048.148 | 943 | 0.805 | 3678 | 12.729 |  |  | Nitrogen | 3 | 12.4236 | 89.075 |  |  |  | 88.427 | 0 |  | 0.73352 |  | 0.003725 |  | 0.37112 | BGD 28: 41.3; BGD 29: 31.4; BGD 30: 1077.3 | 2021/05/19 14:05:51 | 0
05/19/21 | 22 | 20210048.148 | 943 | 0.805 |  |  | 821 | -31.437 | Carbon | 4 | 50.9016 | 22.371 |  |  | 22.025 |  | 0 | 1.1501 |  | 0.010829 |  | 1.0713 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/19 14:05:51 | 95
05/19/21 | 22 | 20210048.148 | 943 | 0.805 |  |  | 5393 | -43.8 |  | 5 | 53.0957 | 100.859 |  |  | 99.328 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/19 14:05:51 | 95
05/19/21 | 22 | 20210048.148 | 943 | 0.805 |  |  | 5387 | -43.974 |  | 6 | 52.707 | 101.085 |  |  | 99.55 |  | 0 | 1.135 |  | 0.010689 |  | 1.0576 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/19 14:05:51 | 95
05/19/21 | 23 | 20210048.149 | 945 | 0.84 | 6398 | 0.128 |  |  |  | 1 | 14.1673 | 117.993 |  |  |  | 117.144 | 0 |  | 0.72437 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.4; BGD 29: 31.6; BGD 30: 988.8 | 2021/05/19 14:15:36 | 0
05/19/21 | 23 | 20210048.149 | 945 | 0.84 | 6405 | 0 |  |  |  | 2 | 14.1657 | 117.968 |  |  |  | 117.119 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.9; BGD 30: 1109.2 | 2021/05/19 14:15:36 | 0
05/19/21 | 23 | 20210048.149 | 945 | 0.84 | 3436 | 12.043 |  |  | Nitrogen | 3 | 11.333 | 83.022 |  |  |  | 82.418 | 0 |  | 0.733 |  | 0.0037225 |  | 0.37087 | BGD 28: 41.2; BGD 29: 31.4; BGD 30: 1078.5 | 2021/05/19 14:15:36 | 0
05/19/21 | 23 | 20210048.149 | 945 | 0.84 |  |  | 915 | -34.253 | Carbon | 4 | 53.2184 | 25.068 |  |  | 24.682 |  | 0 | 1.147 |  | 0.010797 |  | 1.0682 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/19 14:15:36 | 95
05/19/21 | 23 | 20210048.149 | 945 | 0.84 |  |  | 5394 | -43.8 |  | 5 | 50.7931 | 100.914 |  |  | 99.382 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/19 14:15:36 | 95
05/19/21 | 23 | 20210048.149 | 945 | 0.84 |  |  | 5390 | -44.035 |  | 6 | 50.4986 | 101.092 |  |  | 99.557 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.5 | 2021/05/19 14:15:36 | 95
05/19/21 | 24 | 20210048.150 | 947 | 0.829 | 6397 | 0.094 |  |  |  | 1 | 14.3548 | 117.985 |  |  |  | 117.136 | 0 |  | 0.72436 |  | 0.0036785 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.6; BGD 30: 990.6 | 2021/05/19 14:25:21 | 0
05/19/21 | 24 | 20210048.150 | 947 | 0.829 | 6408 | 0 |  |  |  | 2 | 14.354 | 117.972 |  |  |  | 117.124 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.3; BGD 29: 30.8; BGD 30: 1111.5 | 2021/05/19 14:25:21 | 0
05/19/21 | 24 | 20210048.150 | 947 | 0.829 | 3487 | 12.022 |  |  | Nitrogen | 3 | 11.6286 | 84.502 |  |  |  | 83.887 | 0 |  | 0.733 |  | 0.0037224 |  | 0.37086 | BGD 28: 41.3; BGD 29: 31.4; BGD 30: 1081.2 | 2021/05/19 14:25:21 | 0
05/19/21 | 24 | 20210048.150 | 947 | 0.829 |  |  | 881 | -33.772 | Carbon | 4 | 52.2585 | 24.047 |  |  | 23.676 |  | 0 | 1.1475 |  | 0.010803 |  | 1.0687 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/19 14:25:21 | 95
05/19/21 | 24 | 20210048.150 | 947 | 0.829 |  |  | 5392 | -43.8 |  | 5 | 51.5791 | 100.847 |  |  | 99.316 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/19 14:25:21 | 95
05/19/21 | 24 | 20210048.150 | 947 | 0.829 |  |  | 5389 | -43.999 |  | 6 | 51.264 | 101.035 |  |  | 99.501 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.6 | 2021/05/19 14:25:21 | 95
05/19/21 | 25 | 20210048.151 | 949 | 0.826 | 6400 | 0.109 |  |  |  | 1 | 14.3858 | 117.664 |  |  |  | 116.818 | 0 |  | 0.72426 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.7; BGD 30: 992.9 | 2021/05/19 14:35:06 | 0
05/19/21 | 25 | 20210048.151 | 949 | 0.826 | 6397 | 0 |  |  |  | 2 | 14.4045 | 117.947 |  |  |  | 117.099 | 1 |  | 0.72418 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.9; BGD 30: 1114.2 | 2021/05/19 14:35:06 | 0
05/19/21 | 25 | 20210048.151 | 949 | 0.826 | 3706 | 12.087 |  |  | Nitrogen | 3 | 12.1554 | 89.587 |  |  |  | 88.935 | 0 |  | 0.73294 |  | 0.0037227 |  | 0.37089 | BGD 28: 41.2; BGD 29: 31.4; BGD 30: 1084.8 | 2021/05/19 14:35:06 | 0
05/19/21 | 25 | 20210048.151 | 949 | 0.826 |  |  | 851 | -32.762 | Carbon | 4 | 51.0088 | 23.189 |  |  | 22.831 |  | 0 | 1.1486 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/19 14:35:06 | 95
05/19/21 | 25 | 20210048.151 | 949 | 0.826 |  |  | 5392 | -43.8 |  | 5 | 51.7296 | 100.869 |  |  | 99.337 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/19 14:35:06 | 95
05/19/21 | 25 | 20210048.151 | 949 | 0.826 |  |  | 5387 | -44.022 |  | 6 | 51.5251 | 100.991 |  |  | 99.458 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.6 | 2021/05/19 14:35:06 | 95
05/19/21 | 26 | 20210048.152 | 951 | 0.843 | 6399 | 0.113 |  |  |  | 1 | 14.0972 | 117.687 |  |  |  | 116.841 | 0 |  | 0.72426 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.7; BGD 30: 996.8 | 2021/05/19 14:44:51 | 0
05/19/21 | 26 | 20210048.152 | 951 | 0.843 | 6390 | 0 |  |  |  | 2 | 14.101 | 117.745 |  |  |  | 116.899 | 1 |  | 0.72417 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.9; BGD 30: 1117.5 | 2021/05/19 14:44:51 | 0
05/19/21 | 26 | 20210048.152 | 951 | 0.843 | 3951 | 12.309 |  |  | Nitrogen | 3 | 12.4059 | 95.196 |  |  |  | 94.504 | 0 |  | 0.73309 |  | 0.0037235 |  | 0.37097 | BGD 28: 41.3; BGD 29: 31.4; BGD 30: 1088.4 | 2021/05/19 14:44:51 | 0
05/19/21 | 26 | 20210048.152 | 951 | 0.843 |  |  | 846 | -31.093 | Carbon | 4 | 49.6484 | 22.989 |  |  | 22.634 |  | 0 | 1.1505 |  | 0.010833 |  | 1.0716 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/19 14:44:51 | 95
05/19/21 | 26 | 20210048.152 | 951 | 0.843 |  |  | 5393 | -43.8 |  | 5 | 50.6559 | 100.888 |  |  | 99.356 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/19 14:44:51 | 95
05/19/21 | 26 | 20210048.152 | 951 | 0.843 |  |  | 5386 | -44.017 |  | 6 | 50.4332 | 101.023 |  |  | 99.489 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.6 | 2021/05/19 14:44:51 | 95
05/19/21 | 27 | 20210048.153 | 953 | 0.827 | 6386 | 0.096 |  |  |  | 1 | 14.3694 | 117.679 |  |  |  | 116.833 | 0 |  | 0.72426 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.7; BGD 30: 1000.8 | 2021/05/19 14:54:36 | 0
05/19/21 | 27 | 20210048.153 | 953 | 0.827 | 6396 | 0 |  |  |  | 2 | 14.3728 | 117.73 |  |  |  | 116.884 | 1 |  | 0.72419 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.9; BGD 30: 1122.2 | 2021/05/19 14:54:36 | 0
05/19/21 | 27 | 20210048.153 | 953 | 0.827 | 3473 | 13.363 |  |  | Nitrogen | 3 | 11.6219 | 84.146 |  |  |  | 83.533 | 0 |  | 0.73387 |  | 0.0037274 |  | 0.37135 | BGD 28: 41.2; BGD 29: 31.4; BGD 30: 1091.0 | 2021/05/19 14:54:36 | 0
05/19/21 | 27 | 20210048.153 | 953 | 0.827 |  |  | 869 | -32.644 | Carbon | 4 | 51.9134 | 23.763 |  |  | 23.396 |  | 0 | 1.1488 |  | 0.010815 |  | 1.07 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.0 | 2021/05/19 14:54:36 | 95
05/19/21 | 27 | 20210048.153 | 953 | 0.827 |  |  | 5392 | -43.8 |  | 5 | 51.7682 | 100.809 |  |  | 99.278 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/19 14:54:36 | 95
05/19/21 | 27 | 20210048.153 | 953 | 0.827 |  |  | 5382 | -44.021 |  | 6 | 51.375 | 101.043 |  |  | 99.509 |  | 0 | 1.135 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.6 | 2021/05/19 14:54:36 | 95
05/19/21 | 28 | 20210048.154 | 955 | 0.803 | 6398 | 0.105 |  |  |  | 1 | 14.8066 | 117.793 |  |  |  | 116.946 | 0 |  | 0.72446 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.7; BGD 30: 1003.3 | 2021/05/19 15:04:21 | 0
05/19/21 | 28 | 20210048.154 | 955 | 0.803 | 6407 | 0 |  |  |  | 2 | 14.8167 | 117.941 |  |  |  | 117.093 | 1 |  | 0.72438 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.9; BGD 30: 1125.0 | 2021/05/19 15:04:21 | 0
05/19/21 | 28 | 20210048.154 | 955 | 0.803 | 3584 | 9.979 |  |  | Nitrogen | 3 | 12.1993 | 86.453 |  |  |  | 85.825 | 0 |  | 0.73161 |  | 0.0037149 |  | 0.37012 | BGD 28: 41.3; BGD 29: 31.4; BGD 30: 1095.4 | 2021/05/19 15:04:21 | 0
05/19/21 | 28 | 20210048.154 | 955 | 0.803 |  |  | 824 | -31.884 | Carbon | 4 | 51.0927 | 22.407 |  |  | 22.061 |  | 0 | 1.1496 |  | 0.010824 |  | 1.0708 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.1 | 2021/05/19 15:04:21 | 95
05/19/21 | 28 | 20210048.154 | 955 | 0.803 |  |  | 5395 | -43.8 |  | 5 | 53.1888 | 100.882 |  |  | 99.35 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/19 15:04:21 | 95
05/19/21 | 28 | 20210048.154 | 955 | 0.803 |  |  | 5382 | -44.015 |  | 6 | 52.8921 | 101.053 |  |  | 99.519 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 15:04:21 | 95
05/19/21 | 29 | 20210048.155 | 957 | 0.846 | 6416 | 0.117 |  |  |  | 1 | 14.0659 | 117.978 |  |  |  | 117.13 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.6; BGD 30: 1005.5 | 2021/05/19 15:14:06 | 0
05/19/21 | 29 | 20210048.155 | 957 | 0.846 | 6411 | 0 |  |  |  | 2 | 14.0694 | 118.032 |  |  |  | 117.184 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.4; BGD 29: 30.9; BGD 30: 1127.7 | 2021/05/19 15:14:06 | 0
05/19/21 | 29 | 20210048.155 | 957 | 0.846 | 4029 | 11.627 |  |  | Nitrogen | 3 | 12.5367 | 97.266 |  |  |  | 96.558 | 0 |  | 0.73269 |  | 0.003721 |  | 0.37072 | BGD 28: 41.2; BGD 29: 31.4; BGD 30: 1099.2 | 2021/05/19 15:14:06 | 0
05/19/21 | 29 | 20210048.155 | 957 | 0.846 |  |  | 835 | -31.969 | Carbon | 4 | 49.0284 | 22.724 |  |  | 22.373 |  | 0 | 1.1495 |  | 0.010823 |  | 1.0707 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.2 | 2021/05/19 15:14:06 | 95
05/19/21 | 29 | 20210048.155 | 957 | 0.846 |  |  | 5403 | -43.8 |  | 5 | 50.276 | 101.01 |  |  | 99.476 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 15:14:06 | 95
05/19/21 | 29 | 20210048.155 | 957 | 0.846 |  |  | 5406 | -44.02 |  | 6 | 49.5525 | 101.447 |  |  | 99.907 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 15:14:06 | 95
05/19/21 | 30 | 20210048.156 | 959 | 0.846 | 6412 | 0.101 |  |  |  | 1 | 14.0587 | 117.866 |  |  |  | 117.018 | 0 |  | 0.72424 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.3; BGD 29: 31.6; BGD 30: 1007.5 | 2021/05/19 15:23:51 | 0
05/19/21 | 30 | 20210048.156 | 959 | 0.846 | 6408 | 0 |  |  |  | 2 | 14.0734 | 118.094 |  |  |  | 117.245 | 1 |  | 0.72416 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.8; BGD 30: 1128.9 | 2021/05/19 15:23:51 | 0
05/19/21 | 30 | 20210048.156 | 959 | 0.846 | 4035 | 12.117 |  |  | Nitrogen | 3 | 12.5638 | 97.591 |  |  |  | 96.881 | 0 |  | 0.73294 |  | 0.0037228 |  | 0.3709 | BGD 28: 41.2; BGD 29: 31.4; BGD 30: 1100.1 | 2021/05/19 15:23:51 | 0
05/19/21 | 30 | 20210048.156 | 959 | 0.846 |  |  | 849 | -29.876 | Carbon | 4 | 49.7504 | 23.157 |  |  | 22.799 |  | 0 | 1.1518 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.2 | 2021/05/19 15:23:51 | 95
05/19/21 | 30 | 20210048.156 | 959 | 0.846 |  |  | 5409 | -43.8 |  | 5 | 50.0982 | 101.118 |  |  | 99.582 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 15:23:51 | 95
05/19/21 | 30 | 20210048.156 | 959 | 0.846 |  |  | 5404 | -44.005 |  | 6 | 49.6338 | 101.398 |  |  | 99.859 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 15:23:51 | 95
05/19/21 | 31 | 20210048.157 | 961 | 0.824 | 6410 | 0.105 |  |  |  | 1 | 14.4563 | 118.203 |  |  |  | 117.353 | 0 |  | 0.72443 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.4; BGD 29: 31.7; BGD 30: 1008.7 | 2021/05/19 15:33:36 | 0
05/19/21 | 31 | 20210048.157 | 961 | 0.824 | 6416 | 0 |  |  |  | 2 | 14.4574 | 118.22 |  |  |  | 117.369 | 1 |  | 0.72435 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.3; BGD 29: 30.8; BGD 30: 1130.4 | 2021/05/19 15:33:36 | 0
05/19/21 | 31 | 20210048.157 | 961 | 0.824 | 3329 | 13.074 |  |  | Nitrogen | 3 | 11.3168 | 80.671 |  |  |  | 80.083 | 0 |  | 0.73382 |  | 0.0037263 |  | 0.37124 | BGD 28: 41.2; BGD 29: 31.4; BGD 30: 1099.0 | 2021/05/19 15:33:36 | 0
05/19/21 | 31 | 20210048.157 | 961 | 0.824 |  |  | 881 | -32.89 | Carbon | 4 | 52.5952 | 24.058 |  |  | 23.687 |  | 0 | 1.1484 |  | 0.010812 |  | 1.0697 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.1 | 2021/05/19 15:33:36 | 95
05/19/21 | 31 | 20210048.157 | 961 | 0.824 |  |  | 5411 | -43.8 |  | 5 | 51.4203 | 101.127 |  |  | 99.591 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.4 | 2021/05/19 15:33:36 | 95
05/19/21 | 31 | 20210048.157 | 961 | 0.824 |  |  | 5405 | -44.008 |  | 6 | 51.0515 | 101.344 |  |  | 99.806 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.8 | 2021/05/19 15:33:36 | 95
05/19/21 | 32 | 20210048.158 | 963 | 0.827 | 6404 | 0.1 |  |  |  | 1 | 14.3907 | 118.003 |  |  |  | 117.154 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.2; BGD 29: 31.6; BGD 30: 1006.8 | 2021/05/19 15:43:21 | 0
05/19/21 | 32 | 20210048.158 | 963 | 0.827 | 6424 | 0 |  |  |  | 2 | 14.3955 | 118.077 |  |  |  | 117.227 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.8; BGD 30: 1128.6 | 2021/05/19 15:43:21 | 0
05/19/21 | 32 | 20210048.158 | 963 | 0.827 | 3793 | 9.809 |  |  | Nitrogen | 3 | 12.3542 | 91.917 |  |  |  | 91.249 | 0 |  | 0.73136 |  | 0.0037143 |  | 0.37005 | BGD 28: 41.1; BGD 29: 31.4; BGD 30: 1097.6 | 2021/05/19 15:43:21 | 0
05/19/21 | 32 | 20210048.158 | 963 | 0.827 |  |  | 840 | -31.099 | Carbon | 4 | 50.4096 | 22.873 |  |  | 22.519 |  | 0 | 1.1504 |  | 0.010833 |  | 1.0716 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.2 | 2021/05/19 15:43:21 | 95
05/19/21 | 32 | 20210048.158 | 963 | 0.827 |  |  | 5409 | -43.8 |  | 5 | 51.1634 | 101.168 |  |  | 99.632 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 15:43:21 | 95
05/19/21 | 32 | 20210048.158 | 963 | 0.827 |  |  | 5407 | -44.023 |  | 6 | 50.9874 | 101.272 |  |  | 99.735 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.8 | 2021/05/19 15:43:21 | 95
05/19/21 | 33 | 20210048.158 | Replicate C8 | 0.821 | 6405 | 0.137 |  |  |  | 1 | 14.4967 | 118.015 |  |  |  | 117.167 | 0 |  | 0.72433 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.2; BGD 29: 31.5; BGD 30: 1008.5 | 2021/05/19 15:53:06 | 0
05/19/21 | 33 | 20210048.158 | Replicate C8 | 0.821 | 6412 | 0 |  |  |  | 2 | 14.4995 | 118.057 |  |  |  | 117.208 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1129.6 | 2021/05/19 15:53:06 | 0
05/19/21 | 33 | 20210048.158 | Replicate C8 | 0.821 | 3738 | 9.841 |  |  | Nitrogen | 3 | 12.3156 | 90.514 |  |  |  | 89.857 | 0 |  | 0.73136 |  | 0.0037144 |  | 0.37006 | BGD 28: 41.1; BGD 29: 31.3; BGD 30: 1100.8 | 2021/05/19 15:53:06 | 0
05/19/21 | 33 | 20210048.158 | Replicate C8 | 0.821 |  |  | 818 | -30.989 | Carbon | 4 | 49.7408 | 22.274 |  |  | 21.93 |  | 0 | 1.1505 |  | 0.010834 |  | 1.0718 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/19 15:53:06 | 95
05/19/21 | 33 | 20210048.158 | Replicate C8 | 0.821 |  |  | 5406 | -43.8 |  | 5 | 51.6503 | 101.102 |  |  | 99.567 |  | 1 | 1.1352 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 15:53:06 | 95
05/19/21 | 33 | 20210048.158 | Replicate C8 | 0.821 |  |  | 5403 | -44.015 |  | 6 | 51.289 | 101.314 |  |  | 99.776 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 15:53:06 | 95
05/19/21 | 34 | 36-UWSIF-UT Glut 1-20210048.36 | 36-UWSIF-UT Glut 1- | 0.801 | 6417 | 0.124 |  |  |  | 1 | 14.8585 | 118.012 |  |  |  | 117.163 | 0 |  | 0.72433 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1008.9 | 2021/05/19 16:02:52 | 0
05/19/21 | 34 | 36-UWSIF-UT Glut 1-20210048.36 | 36-UWSIF-UT Glut 1- | 0.801 | 6409 | 0 |  |  |  | 2 | 14.862 | 118.064 |  |  |  | 117.215 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1130.3 | 2021/05/19 16:02:52 | 0
05/19/21 | 34 | 36-UWSIF-UT Glut 1-20210048.36 | 36-UWSIF-UT Glut 1- | 0.801 | 2382 | -2.943 |  |  | Nitrogen | 3 | 9.0475 | 58.231 |  |  |  | 57.813 | 0 |  | 0.72211 |  | 0.0036674 |  | 0.3654 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1097.9 | 2021/05/19 16:02:52 | 0
05/19/21 | 34 | 36-UWSIF-UT Glut 1-20210048.36 | 36-UWSIF-UT Glut 1- | 0.801 |  |  | 674 | -29.867 | Carbon | 4 | 43.2377 | 18.159 |  |  | 17.878 |  | 0 | 1.1518 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/19 16:02:52 | 95
05/19/21 | 34 | 36-UWSIF-UT Glut 1-20210048.36 | 36-UWSIF-UT Glut 1- | 0.801 |  |  | 5423 | -43.8 |  | 5 | 52.5155 | 101.345 |  |  | 99.806 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.2 | 2021/05/19 16:02:52 | 95
05/19/21 | 34 | 36-UWSIF-UT Glut 1-20210048.36 | 36-UWSIF-UT Glut 1- | 0.801 |  |  | 5413 | -44.017 |  | 6 | 52.3114 | 101.461 |  |  | 99.921 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.6 | 2021/05/19 16:02:52 | 95
05/19/21 | 35 | 36-UWSIF-UT Glut 1-20210048.37 | 36-UWSIF-UT Glut 1- | 0.786 | 6423 | 0.117 |  |  |  | 1 | 15.1632 | 118.32 |  |  |  | 117.469 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.0; BGD 29: 31.4; BGD 30: 1005.0 | 2021/05/19 16:12:37 | 0
05/19/21 | 35 | 36-UWSIF-UT Glut 1-20210048.37 | 36-UWSIF-UT Glut 1- | 0.786 | 6421 | 0 |  |  |  | 2 | 15.1571 | 118.231 |  |  |  | 117.381 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.6; BGD 30: 1126.3 | 2021/05/19 16:12:37 | 0
05/19/21 | 35 | 36-UWSIF-UT Glut 1-20210048.37 | 36-UWSIF-UT Glut 1- | 0.786 | 2342 | -2.875 |  |  | Nitrogen | 3 | 9.0782 | 57.136 |  |  |  | 56.726 | 0 |  | 0.72215 |  | 0.0036676 |  | 0.36542 | BGD 28: 40.9; BGD 29: 31.1; BGD 30: 1093.9 | 2021/05/19 16:12:37 | 0
05/19/21 | 35 | 36-UWSIF-UT Glut 1-20210048.37 | 36-UWSIF-UT Glut 1- | 0.786 |  |  | 663 | -29.855 | Carbon | 4 | 43.4959 | 17.878 |  |  | 17.602 |  | 0 | 1.1518 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/19 16:12:37 | 95
05/19/21 | 35 | 36-UWSIF-UT Glut 1-20210048.37 | 36-UWSIF-UT Glut 1- | 0.786 |  |  | 5421 | -43.8 |  | 5 | 53.5027 | 101.353 |  |  | 99.815 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.2 | 2021/05/19 16:12:37 | 95
05/19/21 | 35 | 36-UWSIF-UT Glut 1-20210048.37 | 36-UWSIF-UT Glut 1- | 0.786 |  |  | 5412 | -44.003 |  | 6 | 53.1859 | 101.531 |  |  | 99.989 |  | 0 | 1.1349 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.6 | 2021/05/19 16:12:37 | 95
05/19/21 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 39-UWSIF-UW Glut 2- | 0.779 | 6422 | 0.124 |  |  |  | 1 | 15.2952 | 118.259 |  |  |  | 117.408 | 0 |  | 0.72432 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1001.0 | 2021/05/19 16:22:22 | 0
05/19/21 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 39-UWSIF-UW Glut 2- | 0.779 | 6440 | 0 |  |  |  | 2 | 15.3049 | 118.399 |  |  |  | 117.548 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.9; BGD 29: 30.6; BGD 30: 1122.6 | 2021/05/19 16:22:22 | 0
05/19/21 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 39-UWSIF-UW Glut 2- | 0.779 | 2511 | 29.645 |  |  | Nitrogen | 3 | 9.7051 | 61.356 |  |  |  | 60.902 | 0 |  | 0.7457 |  | 0.0037872 |  | 0.37729 | BGD 28: 40.9; BGD 29: 31.1; BGD 30: 1091.4 | 2021/05/19 16:22:22 | 0
05/19/21 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 39-UWSIF-UW Glut 2- | 0.779 |  |  | 706 | 22.272 | Carbon | 4 | 46.4372 | 19.145 |  |  | 18.838 |  | 0 | 1.2093 |  | 0.011429 |  | 1.13 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/19 16:22:22 | 95
05/19/21 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 39-UWSIF-UW Glut 2- | 0.779 |  |  | 5420 | -43.8 |  | 5 | 53.789 | 101.461 |  |  | 99.921 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/19 16:22:22 | 95
05/19/21 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 39-UWSIF-UW Glut 2- | 0.779 |  |  | 5430 | -44.02 |  | 6 | 53.2892 | 101.737 |  |  | 100.193 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 16:22:22 | 95
05/19/21 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 39-UWSIF-UW Glut 2- | 0.726 | 6434 | 0.097 |  |  |  | 1 | 16.4291 | 118.492 |  |  |  | 117.639 | 0 |  | 0.72438 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1001.3 | 2021/05/19 16:32:07 | 0
05/19/21 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 39-UWSIF-UW Glut 2- | 0.726 | 6445 | 0 |  |  |  | 2 | 16.4433 | 118.683 |  |  |  | 117.829 | 1 |  | 0.72431 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1123.1 | 2021/05/19 16:32:07 | 0
05/19/21 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 39-UWSIF-UW Glut 2- | 0.726 | 2350 | 29.79 |  |  | Nitrogen | 3 | 9.8879 | 57.559 |  |  |  | 57.132 | 0 |  | 0.74589 |  | 0.0037878 |  | 0.37735 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 1091.7 | 2021/05/19 16:32:07 | 0
05/19/21 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 39-UWSIF-UW Glut 2- | 0.726 |  |  | 660 | 22.207 | Carbon | 4 | 46.9978 | 17.836 |  |  | 17.55 |  | 0 | 1.2092 |  | 0.011428 |  | 1.1299 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/19 16:32:07 | 95
05/19/21 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 39-UWSIF-UW Glut 2- | 0.726 |  |  | 5447 | -43.8 |  | 5 | 57.579 | 101.532 |  |  | 99.99 |  | 1 | 1.1351 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/19 16:32:07 | 95
05/19/21 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 39-UWSIF-UW Glut 2- | 0.726 |  |  | 5427 | -44.033 |  | 6 | 57.0078 | 101.825 |  |  | 100.279 |  | 0 | 1.1348 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 16:32:07 | 95
05/19/21 | 38 | 01-UWSIF-Liver-20210048.33 | 01-UWSIF-Liver- | 0.75 | 6452 | 0.106 |  |  |  | 1 | 15.9136 | 118.634 |  |  |  | 117.781 | 0 |  | 0.72439 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.0; BGD 29: 31.4; BGD 30: 999.7 | 2021/05/19 16:41:52 | 0
05/19/21 | 38 | 01-UWSIF-Liver-20210048.33 | 01-UWSIF-Liver- | 0.75 | 6441 | 0 |  |  |  | 2 | 15.9157 | 118.663 |  |  |  | 117.81 | 1 |  | 0.72432 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.6; BGD 30: 1121.7 | 2021/05/19 16:41:52 | 0
05/19/21 | 38 | 01-UWSIF-Liver-20210048.33 | 01-UWSIF-Liver- | 0.75 | 3087 | 8.461 |  |  | Nitrogen | 3 | 11.7577 | 74.783 |  |  |  | 74.241 | 0 |  | 0.73044 |  | 0.0037093 |  | 0.36956 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1089.5 | 2021/05/19 16:41:52 | 0
05/19/21 | 38 | 01-UWSIF-Liver-20210048.33 | 01-UWSIF-Liver- | 0.75 |  |  | 778 | -19.317 | Carbon | 4 | 52.1065 | 21.067 |  |  | 20.739 |  | 0 | 1.1631 |  | 0.010964 |  | 1.0845 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/19 16:41:52 | 95
05/19/21 | 38 | 01-UWSIF-Liver-20210048.33 | 01-UWSIF-Liver- | 0.75 |  |  | 5425 | -43.8 |  | 5 | 55.6476 | 101.579 |  |  | 100.037 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/19 16:41:52 | 95
05/19/21 | 38 | 01-UWSIF-Liver-20210048.33 | 01-UWSIF-Liver- | 0.75 |  |  | 5432 | -44.006 |  | 6 | 55.3023 | 101.762 |  |  | 100.218 |  | 0 | 1.1346 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 16:41:52 | 95
05/19/21 | 39 | 01-UWSIF-Liver-20210048.34 | 01-UWSIF-Liver- | 0.79 | 6427 | 0.083 |  |  |  | 1 | 15.0983 | 118.494 |  |  |  | 117.642 | 0 |  | 0.72427 |  | 0.0036785 |  | 0.3665 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1000.9 | 2021/05/19 16:51:37 | 0
05/19/21 | 39 | 01-UWSIF-Liver-20210048.34 | 01-UWSIF-Liver- | 0.79 | 6439 | 0 |  |  |  | 2 | 15.1038 | 118.574 |  |  |  | 117.721 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1122.7 | 2021/05/19 16:51:37 | 0
05/19/21 | 39 | 01-UWSIF-Liver-20210048.34 | 01-UWSIF-Liver- | 0.79 | 3251 | 8.439 |  |  | Nitrogen | 3 | 11.5954 | 78.722 |  |  |  | 78.151 | 0 |  | 0.73033 |  | 0.0037092 |  | 0.36955 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 1090.3 | 2021/05/19 16:51:37 | 0
05/19/21 | 39 | 01-UWSIF-Liver-20210048.34 | 01-UWSIF-Liver- | 0.79 |  |  | 817 | -19.31 | Carbon | 4 | 51.5793 | 22.212 |  |  | 21.866 |  | 0 | 1.1631 |  | 0.010964 |  | 1.0845 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.1 | 2021/05/19 16:51:37 | 95
05/19/21 | 39 | 01-UWSIF-Liver-20210048.34 | 01-UWSIF-Liver- | 0.79 |  |  | 5438 | -43.8 |  | 5 | 52.513 | 101.756 |  |  | 100.212 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/19 16:51:37 | 95
05/19/21 | 39 | 01-UWSIF-Liver-20210048.34 | 01-UWSIF-Liver- | 0.79 |  |  | 5433 | -44.023 |  | 6 | 52.1922 | 101.934 |  |  | 100.387 |  | 0 | 1.1346 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 16:51:37 | 95
05/19/21 | 40 | 20210048.159 | 965 | 0.825 | 6447 | 0.131 |  |  |  | 1 | 14.4708 | 118.693 |  |  |  | 117.84 | 0 |  | 0.72439 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1001.9 | 2021/05/19 17:01:22 | 0
05/19/21 | 40 | 20210048.159 | 965 | 0.825 | 6457 | 0 |  |  |  | 2 | 14.4755 | 118.766 |  |  |  | 117.912 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1123.3 | 2021/05/19 17:01:22 | 0
05/19/21 | 40 | 20210048.159 | 965 | 0.825 | 3728 | 12.89 |  |  | Nitrogen | 3 | 12.2208 | 90.134 |  |  |  | 89.478 | 0 |  | 0.73363 |  | 0.0037256 |  | 0.37118 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 1093.6 | 2021/05/19 17:01:22 | 0
05/19/21 | 40 | 20210048.159 | 965 | 0.825 |  |  | 848 | -30.661 | Carbon | 4 | 50.8922 | 23.084 |  |  | 22.727 |  | 0 | 1.1506 |  | 0.010837 |  | 1.0721 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.2 | 2021/05/19 17:01:22 | 95
05/19/21 | 40 | 20210048.159 | 965 | 0.825 |  |  | 5448 | -43.8 |  | 5 | 50.1998 | 101.806 |  |  | 100.26 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 17:01:22 | 95
05/19/21 | 40 | 20210048.159 | 965 | 0.825 |  |  | 5441 | -44.002 |  | 6 | 49.818 | 102.027 |  |  | 100.479 |  | 0 | 1.1346 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 17:01:22 | 95
05/19/21 | 41 | 20210048.160 | 967 | 0.807 | 6457 | 0.118 |  |  |  | 1 | 14.7959 | 118.728 |  |  |  | 117.874 | 0 |  | 0.72428 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.2; BGD 29: 31.5; BGD 30: 1002.5 | 2021/05/19 17:11:07 | 0
05/19/21 | 41 | 20210048.160 | 967 | 0.807 | 6460 | 0 |  |  |  | 2 | 14.8067 | 118.89 |  |  |  | 118.036 | 1 |  | 0.72419 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1124.2 | 2021/05/19 17:11:07 | 0
05/19/21 | 41 | 20210048.160 | 967 | 0.807 | 3330 | 11.46 |  |  | Nitrogen | 3 | 11.5458 | 80.58 |  |  |  | 79.994 | 0 |  | 0.73249 |  | 0.0037204 |  | 0.37066 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1093.2 | 2021/05/19 17:11:07 | 0
05/19/21 | 41 | 20210048.160 | 967 | 0.807 |  |  | 890 | -34.016 | Carbon | 4 | 54.073 | 24.277 |  |  | 23.903 |  | 0 | 1.1468 |  | 0.0108 |  | 1.0684 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/19 17:11:07 | 95
05/19/21 | 41 | 20210048.160 | 967 | 0.807 |  |  | 5460 | -43.8 |  | 5 | 51.0292 | 101.97 |  |  | 100.423 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.4 | 2021/05/19 17:11:07 | 95
05/19/21 | 41 | 20210048.160 | 967 | 0.807 |  |  | 5441 | -44.017 |  | 6 | 50.9232 | 102.03 |  |  | 100.482 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.8 | 2021/05/19 17:11:07 | 95
05/19/21 | 42 | 20210048.161 | 969 | 0.809 | 6466 | 0.126 |  |  |  | 1 | 14.7746 | 118.959 |  |  |  | 118.103 | 0 |  | 0.72445 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.2; BGD 29: 31.6; BGD 30: 1002.3 | 2021/05/19 17:20:52 | 0
05/19/21 | 42 | 20210048.161 | 969 | 0.809 | 6465 | 0 |  |  |  | 2 | 14.7781 | 119.012 |  |  |  | 118.156 | 1 |  | 0.72435 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1123.3 | 2021/05/19 17:20:52 | 0
05/19/21 | 42 | 20210048.161 | 969 | 0.809 | 3572 | 11.974 |  |  | Nitrogen | 3 | 12.1233 | 86.601 |  |  |  | 85.971 | 0 |  | 0.73303 |  | 0.0037222 |  | 0.37084 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1094.7 | 2021/05/19 17:20:52 | 0
05/19/21 | 42 | 20210048.161 | 969 | 0.809 |  |  | 830 | -33.071 | Carbon | 4 | 50.9984 | 22.569 |  |  | 22.22 |  | 0 | 1.1478 |  | 0.010811 |  | 1.0695 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.2 | 2021/05/19 17:20:52 | 95
05/19/21 | 42 | 20210048.161 | 969 | 0.809 |  |  | 5465 | -43.8 |  | 5 | 51.0832 | 101.868 |  |  | 100.322 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 17:20:52 | 95
05/19/21 | 42 | 20210048.161 | 969 | 0.809 |  |  | 5446 | -44.053 |  | 6 | 50.7304 | 102.068 |  |  | 100.519 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.8 | 2021/05/19 17:20:52 | 95
05/19/21 | 43 | 20210048.162 | 971 | 0.808 | 6466 | 0.127 |  |  |  | 1 | 14.7943 | 118.98 |  |  |  | 118.124 | 0 |  | 0.72444 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1002.2 | 2021/05/19 17:30:37 | 0
05/19/21 | 43 | 20210048.162 | 971 | 0.808 | 6462 | 0 |  |  |  | 2 | 14.7894 | 118.907 |  |  |  | 118.051 | 1 |  | 0.72434 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1123.2 | 2021/05/19 17:30:37 | 0
05/19/21 | 43 | 20210048.162 | 971 | 0.808 | 3329 | 11.648 |  |  | Nitrogen | 3 | 11.5125 | 80.398 |  |  |  | 79.813 | 0 |  | 0.73278 |  | 0.003721 |  | 0.37073 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1093.3 | 2021/05/19 17:30:37 | 0
05/19/21 | 43 | 20210048.162 | 971 | 0.808 |  |  | 867 | -33.489 | Carbon | 4 | 52.8379 | 23.59 |  |  | 23.226 |  | 0 | 1.1473 |  | 0.010806 |  | 1.069 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.3 | 2021/05/19 17:30:37 | 95
05/19/21 | 43 | 20210048.162 | 971 | 0.808 |  |  | 5447 | -43.8 |  | 5 | 51.1399 | 101.872 |  |  | 100.326 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 17:30:37 | 95
05/19/21 | 43 | 20210048.162 | 971 | 0.808 |  |  | 5443 | -44.011 |  | 6 | 50.8357 | 102.044 |  |  | 100.496 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.8 | 2021/05/19 17:30:37 | 95
05/19/21 | 44 | 20210048.163 | 973 | 0.826 | 6463 | 0.087 |  |  |  | 1 | 14.4725 | 118.989 |  |  |  | 118.133 | 0 |  | 0.7244 |  | 0.0036785 |  | 0.3665 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1002.2 | 2021/05/19 17:40:22 | 0
05/19/21 | 44 | 20210048.163 | 973 | 0.826 | 6465 | 0 |  |  |  | 2 | 14.4727 | 118.992 |  |  |  | 118.137 | 1 |  | 0.72434 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1122.0 | 2021/05/19 17:40:22 | 0
05/19/21 | 44 | 20210048.163 | 973 | 0.826 | 3271 | 11.517 |  |  | Nitrogen | 3 | 11.1255 | 79.066 |  |  |  | 78.491 | 0 |  | 0.73268 |  | 0.0037206 |  | 0.37068 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 1092.0 | 2021/05/19 17:40:22 | 0
05/19/21 | 44 | 20210048.163 | 973 | 0.826 |  |  | 909 | -34.023 | Carbon | 4 | 53.7396 | 24.833 |  |  | 24.45 |  | 0 | 1.1467 |  | 0.0108 |  | 1.0684 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.2 | 2021/05/19 17:40:22 | 95
05/19/21 | 44 | 20210048.163 | 973 | 0.826 |  |  | 5439 | -43.8 |  | 5 | 49.9886 | 101.893 |  |  | 100.347 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.4 | 2021/05/19 17:40:22 | 95
05/19/21 | 44 | 20210048.163 | 973 | 0.826 |  |  | 5448 | -44.006 |  | 6 | 49.6207 | 102.106 |  |  | 100.557 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/19 17:40:22 | 95
05/19/21 | 45 | 20210048.164 | 975 | 0.833 | 6462 | 0.116 |  |  |  | 1 | 14.3498 | 118.972 |  |  |  | 118.116 | 0 |  | 0.72439 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.6; BGD 30: 1001.1 | 2021/05/19 17:50:08 | 0
05/19/21 | 45 | 20210048.164 | 975 | 0.833 | 6466 | 0 |  |  |  | 2 | 14.3443 | 118.887 |  |  |  | 118.032 | 1 |  | 0.72431 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1121.2 | 2021/05/19 17:50:08 | 0
05/19/21 | 45 | 20210048.164 | 975 | 0.833 | 3733 | 11.822 |  |  | Nitrogen | 3 | 12.1111 | 90.218 |  |  |  | 89.562 | 0 |  | 0.73287 |  | 0.0037217 |  | 0.37079 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1093.6 | 2021/05/19 17:50:08 | 0
05/19/21 | 45 | 20210048.164 | 975 | 0.833 |  |  | 852 | -32.845 | Carbon | 4 | 50.5936 | 23.197 |  |  | 22.839 |  | 0 | 1.148 |  | 0.010813 |  | 1.0697 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.3 | 2021/05/19 17:50:08 | 95
05/19/21 | 45 | 20210048.164 | 975 | 0.833 |  |  | 5444 | -43.8 |  | 5 | 49.6849 | 101.825 |  |  | 100.28 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 17:50:08 | 95
05/19/21 | 45 | 20210048.164 | 975 | 0.833 |  |  | 5454 | -44.036 |  | 6 | 49.1172 | 102.157 |  |  | 100.607 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.8 | 2021/05/19 17:50:08 | 95
05/19/21 | 46 | 20210048.165 | 749 | 0.842 | 6452 | 0.121 |  |  |  | 1 | 14.1824 | 118.752 |  |  |  | 117.898 | 0 |  | 0.72428 |  | 0.0036786 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.6; BGD 30: 1001.3 | 2021/05/19 17:59:53 | 0
05/19/21 | 46 | 20210048.165 | 749 | 0.842 | 6452 | 0 |  |  |  | 2 | 14.1796 | 118.709 |  |  |  | 117.856 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1122.3 | 2021/05/19 17:59:53 | 0
05/19/21 | 46 | 20210048.165 | 749 | 0.842 | 4079 | 11.936 |  |  | Nitrogen | 3 | 12.7066 | 98.59 |  |  |  | 97.873 | 0 |  | 0.73284 |  | 0.0037221 |  | 0.37083 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1094.1 | 2021/05/19 17:59:53 | 0
05/19/21 | 46 | 20210048.165 | 749 | 0.842 |  |  | 838 | -31.845 | Carbon | 4 | 49.3396 | 22.77 |  |  | 22.419 |  | 0 | 1.149 |  | 0.010824 |  | 1.0708 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/19 17:59:53 | 95
05/19/21 | 46 | 20210048.165 | 749 | 0.842 |  |  | 5449 | -43.8 |  | 5 | 49.23 | 101.78 |  |  | 100.235 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/19 17:59:53 | 95
05/19/21 | 46 | 20210048.165 | 749 | 0.842 |  |  | 5439 | -44.032 |  | 6 | 48.8467 | 102.007 |  |  | 100.459 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.8 | 2021/05/19 17:59:53 | 95
05/19/21 | 47 | 20210048.166 | 751 | 0.808 | 6450 | 0.105 |  |  |  | 1 | 14.7724 | 118.651 |  |  |  | 117.798 | 0 |  | 0.72436 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1002.7 | 2021/05/19 18:09:38 | 0
05/19/21 | 47 | 20210048.166 | 751 | 0.808 | 6454 | 0 |  |  |  | 2 | 14.7807 | 118.776 |  |  |  | 117.922 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1123.4 | 2021/05/19 18:09:38 | 0
05/19/21 | 47 | 20210048.166 | 751 | 0.808 | 3503 | 11.085 |  |  | Nitrogen | 3 | 11.9548 | 84.742 |  |  |  | 84.126 | 0 |  | 0.73232 |  | 0.003719 |  | 0.37052 | BGD 28: 40.9; BGD 29: 31.3; BGD 30: 1094.9 | 2021/05/19 18:09:38 | 0
05/19/21 | 47 | 20210048.166 | 751 | 0.808 |  |  | 832 | -31.533 | Carbon | 4 | 51.1225 | 22.603 |  |  | 22.254 |  | 0 | 1.1494 |  | 0.010828 |  | 1.0712 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/19 18:09:38 | 95
05/19/21 | 47 | 20210048.166 | 751 | 0.808 |  |  | 5439 | -43.8 |  | 5 | 51.4622 | 101.688 |  |  | 100.145 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 18:09:38 | 95
05/19/21 | 47 | 20210048.166 | 751 | 0.808 |  |  | 5442 | -44.028 |  | 6 | 51.0767 | 101.908 |  |  | 100.361 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.9 | 2021/05/19 18:09:38 | 95
05/19/21 | 48 | 20210048.167 | 753 | 0.827 | 6453 | 0.116 |  |  |  | 1 | 14.433 | 118.65 |  |  |  | 117.797 | 0 |  | 0.72436 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.0; BGD 29: 31.5; BGD 30: 1002.9 | 2021/05/19 18:19:23 | 0
05/19/21 | 48 | 20210048.167 | 753 | 0.827 | 6448 | 0 |  |  |  | 2 | 14.44 | 118.758 |  |  |  | 117.904 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.7; BGD 30: 1123.5 | 2021/05/19 18:19:23 | 0
05/19/21 | 48 | 20210048.167 | 753 | 0.827 | 3552 | 12.24 |  |  | Nitrogen | 3 | 11.789 | 85.867 |  |  |  | 85.242 | 0 |  | 0.73314 |  | 0.0037232 |  | 0.37094 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1095.1 | 2021/05/19 18:19:23 | 0
05/19/21 | 48 | 20210048.167 | 753 | 0.827 |  |  | 844 | -32.914 | Carbon | 4 | 50.554 | 22.957 |  |  | 22.603 |  | 0 | 1.1478 |  | 0.010812 |  | 1.0697 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/19 18:19:23 | 95
05/19/21 | 48 | 20210048.167 | 753 | 0.827 |  |  | 5442 | -43.8 |  | 5 | 50.243 | 101.71 |  |  | 100.166 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.4 | 2021/05/19 18:19:23 | 95
05/19/21 | 48 | 20210048.167 | 753 | 0.827 |  |  | 5439 | -44.03 |  | 6 | 49.8951 | 101.912 |  |  | 100.366 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.8 | 2021/05/19 18:19:23 | 95
05/19/21 | 49 | 20210048.168 | 755 | 0.808 | 6444 | 0.108 |  |  |  | 1 | 14.7727 | 118.655 |  |  |  | 117.802 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.0; BGD 29: 31.5; BGD 30: 1004.0 | 2021/05/19 18:29:08 | 0
05/19/21 | 49 | 20210048.168 | 755 | 0.808 | 6449 | 0 |  |  |  | 2 | 14.7775 | 118.727 |  |  |  | 117.873 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.7; BGD 30: 1124.7 | 2021/05/19 18:29:08 | 0
05/19/21 | 49 | 20210048.168 | 755 | 0.808 | 3639 | 12.506 |  |  | Nitrogen | 3 | 12.2654 | 87.908 |  |  |  | 87.268 | 0 |  | 0.73331 |  | 0.0037242 |  | 0.37104 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 1095.2 | 2021/05/19 18:29:08 | 0
05/19/21 | 49 | 20210048.168 | 755 | 0.808 |  |  | 821 | -32.251 | Carbon | 4 | 50.4976 | 22.25 |  |  | 21.906 |  | 0 | 1.1485 |  | 0.01082 |  | 1.0704 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/19 18:29:08 | 95
05/19/21 | 49 | 20210048.168 | 755 | 0.808 |  |  | 5435 | -43.8 |  | 5 | 51.5529 | 101.636 |  |  | 100.094 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/19 18:29:08 | 95
05/19/21 | 49 | 20210048.168 | 755 | 0.808 |  |  | 5434 | -44.012 |  | 6 | 51.1114 | 101.888 |  |  | 100.342 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.8 | 2021/05/19 18:29:08 | 95
05/19/21 | 50 | 20210048.169 | 757 | 0.838 | 6445 | 0.121 |  |  |  | 1 | 14.2365 | 118.541 |  |  |  | 117.688 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36652 | BGD 28: 41.0; BGD 29: 31.5; BGD 30: 1005.9 | 2021/05/19 18:38:53 | 0
05/19/21 | 50 | 20210048.169 | 757 | 0.838 | 6447 | 0 |  |  |  | 2 | 14.2456 | 118.682 |  |  |  | 117.829 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.7; BGD 30: 1127.1 | 2021/05/19 18:38:53 | 0
05/19/21 | 50 | 20210048.169 | 757 | 0.838 | 3540 | 11.715 |  |  | Nitrogen | 3 | 11.5995 | 85.501 |  |  |  | 84.879 | 0 |  | 0.73273 |  | 0.0037213 |  | 0.37075 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1098.3 | 2021/05/19 18:38:53 | 0
05/19/21 | 50 | 20210048.169 | 757 | 0.838 |  |  | 877 | -32.679 | Carbon | 4 | 51.453 | 23.898 |  |  | 23.529 |  | 0 | 1.1479 |  | 0.010815 |  | 1.0699 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/19 18:38:53 | 95
05/19/21 | 50 | 20210048.169 | 757 | 0.838 |  |  | 5435 | -43.8 |  | 5 | 49.8751 | 101.537 |  |  | 99.996 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/19 18:38:53 | 95
05/19/21 | 50 | 20210048.169 | 757 | 0.838 |  |  | 5431 | -44.032 |  | 6 | 49.5261 | 101.744 |  |  | 100.2 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 8.9 | 2021/05/19 18:38:53 | 95
05/19/21 | 51 | 20210048.170 | 759 | 0.84 | 6434 | 0.114 |  |  |  | 1 | 14.1968 | 118.45 |  |  |  | 117.598 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1008.5 | 2021/05/19 18:48:38 | 0
05/19/21 | 51 | 20210048.170 | 759 | 0.84 | 6434 | 0 |  |  |  | 2 | 14.1993 | 118.489 |  |  |  | 117.637 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.7; BGD 30: 1128.0 | 2021/05/19 18:48:38 | 0
05/19/21 | 51 | 20210048.170 | 759 | 0.84 | 3943 | 12.063 |  |  | Nitrogen | 3 | 12.4543 | 95.244 |  |  |  | 94.551 | 0 |  | 0.73298 |  | 0.0037226 |  | 0.37088 | BGD 28: 40.9; BGD 29: 31.3; BGD 30: 1101.0 | 2021/05/19 18:48:38 | 0
05/19/21 | 51 | 20210048.170 | 759 | 0.84 |  |  | 846 | -29.339 | Carbon | 4 | 49.831 | 22.993 |  |  | 22.637 |  | 0 | 1.1516 |  | 0.010852 |  | 1.0736 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.4 | 2021/05/19 18:48:38 | 95
05/19/21 | 51 | 20210048.170 | 759 | 0.84 |  |  | 5426 | -43.8 |  | 5 | 50.05 | 101.361 |  |  | 99.823 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/19 18:48:38 | 95
05/19/21 | 51 | 20210048.170 | 759 | 0.84 |  |  | 5419 | -44.023 |  | 6 | 49.6007 | 101.629 |  |  | 100.088 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.0 | 2021/05/19 18:48:38 | 95
05/19/21 | 52 | 20210048.171 | 761 | 0.816 | 6422 | 0.151 |  |  |  | 1 | 14.6071 | 118.34 |  |  |  | 117.489 | 0 |  | 0.72432 |  | 0.0036788 |  | 0.36653 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1010.9 | 2021/05/19 18:58:23 | 0
05/19/21 | 52 | 20210048.171 | 761 | 0.816 | 6426 | 0 |  |  |  | 2 | 14.6081 | 118.355 |  |  |  | 117.504 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1131.4 | 2021/05/19 18:58:23 | 0
05/19/21 | 52 | 20210048.171 | 761 | 0.816 | 2699 | 12.088 |  |  | Nitrogen | 3 | 9.7121 | 65.097 |  |  |  | 64.623 | 0 |  | 0.73297 |  | 0.0037227 |  | 0.37089 | BGD 28: 40.9; BGD 29: 31.3; BGD 30: 1100.6 | 2021/05/19 18:58:23 | 0
05/19/21 | 52 | 20210048.171 | 761 | 0.816 |  |  | 932 | -33.478 | Carbon | 4 | 55.4353 | 25.469 |  |  | 25.077 |  | 0 | 1.147 |  | 0.010806 |  | 1.069 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.4 | 2021/05/19 18:58:23 | 95
05/19/21 | 52 | 20210048.171 | 761 | 0.816 |  |  | 5426 | -43.8 |  | 5 | 51.3446 | 101.464 |  |  | 99.925 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.6 | 2021/05/19 18:58:23 | 95
05/19/21 | 52 | 20210048.171 | 761 | 0.816 |  |  | 5419 | -44.024 |  | 6 | 51.2703 | 101.507 |  |  | 99.968 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.0 | 2021/05/19 18:58:23 | 95
05/19/21 | 53 | 20210048.172 | 763 | 0.803 | 6426 | 0.121 |  |  |  | 1 | 14.84 | 118.287 |  |  |  | 117.436 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.6; BGD 30: 1009.3 | 2021/05/19 19:08:08 | 0
05/19/21 | 53 | 20210048.172 | 763 | 0.803 | 6428 | 0 |  |  |  | 2 | 14.8394 | 118.279 |  |  |  | 117.428 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1130.4 | 2021/05/19 19:08:08 | 0
05/19/21 | 53 | 20210048.172 | 763 | 0.803 | 3527 | 12.382 |  |  | Nitrogen | 3 | 12.0647 | 85.097 |  |  |  | 84.477 | 0 |  | 0.73319 |  | 0.0037237 |  | 0.37099 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1104.0 | 2021/05/19 19:08:08 | 0
05/19/21 | 53 | 20210048.172 | 763 | 0.803 |  |  | 820 | -29.649 | Carbon | 4 | 50.7672 | 22.225 |  |  | 21.881 |  | 0 | 1.1513 |  | 0.010849 |  | 1.0732 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.4 | 2021/05/19 19:08:08 | 95
05/19/21 | 53 | 20210048.172 | 763 | 0.803 |  |  | 5416 | -43.8 |  | 5 | 52.6157 | 101.213 |  |  | 99.677 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/19 19:08:08 | 95
05/19/21 | 53 | 20210048.172 | 763 | 0.803 |  |  | 5408 | -44.018 |  | 6 | 52.2632 | 101.415 |  |  | 99.876 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.0 | 2021/05/19 19:08:08 | 95
05/19/21 | 54 | 20210048.173 | 765 | 0.817 | 6413 | 0.113 |  |  |  | 1 | 14.5687 | 118.031 |  |  |  | 117.182 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1011.8 | 2021/05/19 19:17:53 | 0
05/19/21 | 54 | 20210048.173 | 765 | 0.817 | 6414 | 0 |  |  |  | 2 | 14.5702 | 118.053 |  |  |  | 117.204 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1132.9 | 2021/05/19 19:17:53 | 0
05/19/21 | 54 | 20210048.173 | 765 | 0.817 | 3676 | 11.275 |  |  | Nitrogen | 3 | 12.2046 | 88.689 |  |  |  | 88.044 | 0 |  | 0.73239 |  | 0.0037197 |  | 0.37059 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1105.8 | 2021/05/19 19:17:53 | 0
05/19/21 | 54 | 20210048.173 | 765 | 0.817 |  |  | 818 | -31.919 | Carbon | 4 | 49.7868 | 22.162 |  |  | 21.82 |  | 0 | 1.1488 |  | 0.010823 |  | 1.0707 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/19 19:17:53 | 95
05/19/21 | 54 | 20210048.173 | 765 | 0.817 |  |  | 5408 | -43.8 |  | 5 | 51.8992 | 101.104 |  |  | 99.57 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.6 | 2021/05/19 19:17:53 | 95
05/19/21 | 54 | 20210048.173 | 765 | 0.817 |  |  | 5409 | -44.011 |  | 6 | 51.4024 | 101.394 |  |  | 99.856 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.1 | 2021/05/19 19:17:53 | 95
05/19/21 | 55 | 20210048.174 | 767 | 0.827 | 6410 | 0.122 |  |  |  | 1 | 14.3886 | 117.97 |  |  |  | 117.122 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1015.1 | 2021/05/19 19:27:39 | 0
05/19/21 | 55 | 20210048.174 | 767 | 0.827 | 6409 | 0 |  |  |  | 2 | 14.3879 | 117.96 |  |  |  | 117.112 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1136.1 | 2021/05/19 19:27:39 | 0
05/19/21 | 55 | 20210048.174 | 767 | 0.827 | 3747 | 11.866 |  |  | Nitrogen | 3 | 12.2009 | 90.238 |  |  |  | 89.582 | 0 |  | 0.73281 |  | 0.0037218 |  | 0.3708 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1108.8 | 2021/05/19 19:27:39 | 0
05/19/21 | 55 | 20210048.174 | 767 | 0.827 |  |  | 825 | -32.187 | Carbon | 4 | 49.51 | 22.349 |  |  | 22.004 |  | 0 | 1.1485 |  | 0.01082 |  | 1.0705 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/19 19:27:39 | 95
05/19/21 | 55 | 20210048.174 | 767 | 0.827 |  |  | 5405 | -43.8 |  | 5 | 51.3115 | 101.081 |  |  | 99.547 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.6 | 2021/05/19 19:27:39 | 95
05/19/21 | 55 | 20210048.174 | 767 | 0.827 |  |  | 5406 | -44.022 |  | 6 | 50.9695 | 101.283 |  |  | 99.747 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/19 19:27:39 | 95
05/19/21 | 56 | 20210048.175 | 769 | 0.824 | 6408 | 0.104 |  |  |  | 1 | 14.4353 | 117.885 |  |  |  | 117.037 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1018.6 | 2021/05/19 19:37:24 | 0
05/19/21 | 56 | 20210048.175 | 769 | 0.824 | 6404 | 0 |  |  |  | 2 | 14.4434 | 118.007 |  |  |  | 117.159 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1139.5 | 2021/05/19 19:37:24 | 0
05/19/21 | 56 | 20210048.175 | 769 | 0.824 | 3328 | 11.862 |  |  | Nitrogen | 3 | 11.2589 | 80.102 |  |  |  | 79.52 | 0 |  | 0.73282 |  | 0.0037218 |  | 0.3708 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 1109.6 | 2021/05/19 19:37:24 | 0
05/19/21 | 56 | 20210048.175 | 769 | 0.824 |  |  | 851 | -33.38 | Carbon | 4 | 50.9908 | 23.105 |  |  | 22.749 |  | 0 | 1.1472 |  | 0.010807 |  | 1.0691 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.5 | 2021/05/19 19:37:24 | 95
05/19/21 | 56 | 20210048.175 | 769 | 0.824 |  |  | 5409 | -43.8 |  | 5 | 51.5333 | 101.06 |  |  | 99.526 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/19 19:37:24 | 95
05/19/21 | 56 | 20210048.175 | 769 | 0.824 |  |  | 5404 | -44.022 |  | 6 | 51.3564 | 101.164 |  |  | 99.63 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/19 19:37:24 | 95
05/19/21 | 57 | 20210048.176 | 771 | 0.837 | 6400 | 0.121 |  |  |  | 1 | 14.2007 | 117.725 |  |  |  | 116.879 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1019.9 | 2021/05/19 19:47:09 | 0
05/19/21 | 57 | 20210048.176 | 771 | 0.837 | 6401 | 0 |  |  |  | 2 | 14.212 | 117.898 |  |  |  | 117.05 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.8; BGD 30: 1141.4 | 2021/05/19 19:47:09 | 0
05/19/21 | 57 | 20210048.176 | 771 | 0.837 | 3801 | 12.088 |  |  | Nitrogen | 3 | 12.1609 | 91.408 |  |  |  | 90.743 | 0 |  | 0.73297 |  | 0.0037227 |  | 0.37089 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1114.2 | 2021/05/19 19:47:09 | 0
05/19/21 | 57 | 20210048.176 | 771 | 0.837 |  |  | 831 | -29.962 | Carbon | 4 | 49.1454 | 22.482 |  |  | 22.135 |  | 0 | 1.151 |  | 0.010845 |  | 1.0729 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/19 19:47:09 | 95
05/19/21 | 57 | 20210048.176 | 771 | 0.837 |  |  | 5404 | -43.8 |  | 5 | 50.936 | 100.938 |  |  | 99.406 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/19 19:47:09 | 95
05/19/21 | 57 | 20210048.176 | 771 | 0.837 |  |  | 5399 | -44.039 |  | 6 | 50.5215 | 101.187 |  |  | 99.652 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.2 | 2021/05/19 19:47:09 | 95
05/19/21 | 58 | 20210048.177 | 773 | 0.83 | 6393 | 0.125 |  |  |  | 1 | 14.314 | 117.627 |  |  |  | 116.781 | 0 |  | 0.72429 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1023.4 | 2021/05/19 19:56:54 | 0
05/19/21 | 58 | 20210048.177 | 773 | 0.83 | 6401 | 0 |  |  |  | 2 | 14.3266 | 117.818 |  |  |  | 116.97 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1145.0 | 2021/05/19 19:56:54 | 0
05/19/21 | 58 | 20210048.177 | 773 | 0.83 | 3793 | 11.284 |  |  | Nitrogen | 3 | 12.2489 | 91.247 |  |  |  | 90.584 | 0 |  | 0.73237 |  | 0.0037197 |  | 0.37059 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1117.8 | 2021/05/19 19:56:54 | 0
05/19/21 | 58 | 20210048.177 | 773 | 0.83 |  |  | 830 | -31.856 | Carbon | 4 | 49.5969 | 22.504 |  |  | 22.156 |  | 0 | 1.1489 |  | 0.010824 |  | 1.0708 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.7 | 2021/05/19 19:56:54 | 95
05/19/21 | 58 | 20210048.177 | 773 | 0.83 |  |  | 5394 | -43.8 |  | 5 | 51.5903 | 100.803 |  |  | 99.273 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.7 | 2021/05/19 19:56:54 | 95
05/19/21 | 58 | 20210048.177 | 773 | 0.83 |  |  | 5387 | -44.013 |  | 6 | 51.3103 | 100.971 |  |  | 99.439 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.2 | 2021/05/19 19:56:54 | 95
05/19/21 | 59 | 20210048.178 | 775 | 0.82 | 6389 | 0.12 |  |  |  | 1 | 14.486 | 117.588 |  |  |  | 116.742 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36652 | BGD 28: 41.2; BGD 29: 31.6; BGD 30: 1026.6 | 2021/05/19 20:06:39 | 0
05/19/21 | 59 | 20210048.178 | 775 | 0.82 | 6395 | 0 |  |  |  | 2 | 14.4986 | 117.777 |  |  |  | 116.93 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.8; BGD 30: 1148.6 | 2021/05/19 20:06:39 | 0
05/19/21 | 59 | 20210048.178 | 775 | 0.82 | 3233 | 11.987 |  |  | Nitrogen | 3 | 11.0612 | 77.67 |  |  |  | 77.105 | 0 |  | 0.73289 |  | 0.0037223 |  | 0.37085 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1119.7 | 2021/05/19 20:06:39 | 0
05/19/21 | 59 | 20210048.178 | 775 | 0.82 |  |  | 906 | -33.734 | Carbon | 4 | 53.8618 | 24.668 |  |  | 24.288 |  | 0 | 1.1468 |  | 0.010803 |  | 1.0688 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/19 20:06:39 | 95
05/19/21 | 59 | 20210048.178 | 775 | 0.82 |  |  | 5392 | -43.8 |  | 5 | 52.2264 | 100.799 |  |  | 99.269 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.8 | 2021/05/19 20:06:39 | 95
05/19/21 | 59 | 20210048.178 | 775 | 0.82 |  |  | 5393 | -44.02 |  | 6 | 51.7425 | 101.085 |  |  | 99.551 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.2 | 2021/05/19 20:06:39 | 95
05/19/21 | 60 | 20210048.178 | Replicate E11 | 0.836 | 6389 | 0.1 |  |  |  | 1 | 14.2123 | 117.641 |  |  |  | 116.796 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.2; BGD 29: 31.6; BGD 30: 1027.5 | 2021/05/19 20:16:24 | 0
05/19/21 | 60 | 20210048.178 | Replicate E11 | 0.836 | 6396 | 0 |  |  |  | 2 | 14.2179 | 117.728 |  |  |  | 116.881 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.8; BGD 30: 1149.4 | 2021/05/19 20:16:24 | 0
05/19/21 | 60 | 20210048.178 | Replicate E11 | 0.836 | 3739 | 12.051 |  |  | Nitrogen | 3 | 12.0391 | 89.905 |  |  |  | 89.251 | 0 |  | 0.73296 |  | 0.0037225 |  | 0.37087 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1122.4 | 2021/05/19 20:16:24 | 0
05/19/21 | 60 | 20210048.178 | Replicate E11 | 0.836 |  |  | 883 | -32.671 | Carbon | 4 | 51.8037 | 24.036 |  |  | 23.665 |  | 0 | 1.148 |  | 0.010815 |  | 1.0699 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.8 | 2021/05/19 20:16:24 | 95
05/19/21 | 60 | 20210048.178 | Replicate E11 | 0.836 |  |  | 5391 | -43.8 |  | 5 | 51.1093 | 100.87 |  |  | 99.339 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.9 | 2021/05/19 20:16:24 | 95
05/19/21 | 60 | 20210048.178 | Replicate E11 | 0.836 |  |  | 5392 | -44.021 |  | 6 | 50.7989 | 101.057 |  |  | 99.524 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.3 | 2021/05/19 20:16:24 | 95
05/19/21 | 61 | 20210048.179 | 777 | 0.845 | 6391 | 0.11 |  |  |  | 1 | 14.0603 | 117.632 |  |  |  | 116.786 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.2; BGD 29: 31.6; BGD 30: 1031.3 | 2021/05/19 20:26:09 | 0
05/19/21 | 61 | 20210048.179 | 777 | 0.845 | 6395 | 0 |  |  |  | 2 | 14.0661 | 117.722 |  |  |  | 116.876 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.8; BGD 30: 1153.7 | 2021/05/19 20:26:09 | 0
05/19/21 | 61 | 20210048.179 | 777 | 0.845 | 3927 | 11.57 |  |  | Nitrogen | 3 | 12.3154 | 94.484 |  |  |  | 93.797 | 0 |  | 0.7326 |  | 0.0037208 |  | 0.3707 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1127.0 | 2021/05/19 20:26:09 | 0
05/19/21 | 61 | 20210048.179 | 777 | 0.845 |  |  | 830 | -30.568 | Carbon | 4 | 48.7228 | 22.507 |  |  | 22.16 |  | 0 | 1.1503 |  | 0.010838 |  | 1.0722 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.8 | 2021/05/19 20:26:09 | 95
05/19/21 | 61 | 20210048.179 | 777 | 0.845 |  |  | 5390 | -43.8 |  | 5 | 50.7318 | 100.768 |  |  | 99.239 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/19 20:26:09 | 95
05/19/21 | 61 | 20210048.179 | 777 | 0.845 |  |  | 5389 | -44.018 |  | 6 | 50.4021 | 100.969 |  |  | 99.437 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.3 | 2021/05/19 20:26:09 | 95
05/19/21 | 62 | 20210048.180 | 779 | 0.85 | 6387 | 0.121 |  |  |  | 1 | 13.9786 | 117.649 |  |  |  | 116.803 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.6; BGD 30: 1034.1 | 2021/05/19 20:35:55 | 0
05/19/21 | 62 | 20210048.180 | 779 | 0.85 | 6388 | 0 |  |  |  | 2 | 13.9772 | 117.626 |  |  |  | 116.78 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1156.5 | 2021/05/19 20:35:55 | 0
05/19/21 | 62 | 20210048.180 | 779 | 0.85 | 3312 | 11.1 |  |  | Nitrogen | 3 | 10.8596 | 79.55 |  |  |  | 78.972 | 0 |  | 0.73226 |  | 0.003719 |  | 0.37052 | BGD 28: 40.9; BGD 29: 31.3; BGD 30: 1127.2 | 2021/05/19 20:35:55 | 0
05/19/21 | 62 | 20210048.180 | 779 | 0.85 |  |  | 903 | -32.516 | Carbon | 4 | 51.8466 | 24.596 |  |  | 24.217 |  | 0 | 1.1482 |  | 0.010817 |  | 1.0701 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.7 | 2021/05/19 20:35:55 | 95
05/19/21 | 62 | 20210048.180 | 779 | 0.85 |  |  | 5387 | -43.8 |  | 5 | 50.3605 | 100.813 |  |  | 99.283 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 7.9 | 2021/05/19 20:35:55 | 95
05/19/21 | 62 | 20210048.180 | 779 | 0.85 |  |  | 5387 | -44.037 |  | 6 | 50.1016 | 100.972 |  |  | 99.44 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/19 20:35:55 | 95
05/19/21 | 63 | 20210048.181 | 781 | 0.81 | 6387 | 0.127 |  |  |  | 1 | 14.6684 | 117.641 |  |  |  | 116.795 | 0 |  | 0.72431 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.2; BGD 29: 31.6; BGD 30: 1034.8 | 2021/05/19 20:45:40 | 0
05/19/21 | 63 | 20210048.181 | 781 | 0.81 | 6395 | 0 |  |  |  | 2 | 14.6714 | 117.685 |  |  |  | 116.839 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.8; BGD 30: 1157.2 | 2021/05/19 20:45:40 | 0
05/19/21 | 63 | 20210048.181 | 781 | 0.81 | 3642 | 13.238 |  |  | Nitrogen | 3 | 12.2179 | 87.73 |  |  |  | 87.091 | 0 |  | 0.73381 |  | 0.0037269 |  | 0.3713 | BGD 28: 41.1; BGD 29: 31.3; BGD 30: 1129.6 | 2021/05/19 20:45:40 | 0
05/19/21 | 63 | 20210048.181 | 781 | 0.81 |  |  | 822 | -32.603 | Carbon | 4 | 50.3789 | 22.253 |  |  | 21.91 |  | 0 | 1.1481 |  | 0.010816 |  | 1.07 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.9 | 2021/05/19 20:45:40 | 95
05/19/21 | 63 | 20210048.181 | 781 | 0.81 |  |  | 5385 | -43.8 |  | 5 | 52.9323 | 100.763 |  |  | 99.234 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.9 | 2021/05/19 20:45:40 | 95
05/19/21 | 63 | 20210048.181 | 781 | 0.81 |  |  | 5394 | -44.004 |  | 6 | 52.4082 | 101.069 |  |  | 99.536 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.3 | 2021/05/19 20:45:40 | 95
05/19/21 | 64 | 36-UWSIF-UT Glut 1-20210048.38 | 36-UWSIF-UT Glut 1- | 0.813 | 6391 | 0.141 |  |  |  | 1 | 14.6094 | 117.568 |  |  |  | 116.723 | 0 |  | 0.72432 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1036.7 | 2021/05/19 20:55:25 | 0
05/19/21 | 64 | 36-UWSIF-UT Glut 1-20210048.38 | 36-UWSIF-UT Glut 1- | 0.813 | 6391 | 0 |  |  |  | 2 | 14.6162 | 117.669 |  |  |  | 116.823 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1159.1 | 2021/05/19 20:55:25 | 0
05/19/21 | 64 | 36-UWSIF-UT Glut 1-20210048.38 | 36-UWSIF-UT Glut 1- | 0.813 | 2426 | -2.897 |  |  | Nitrogen | 3 | 8.9468 | 58.494 |  |  |  | 58.075 | 0 |  | 0.72212 |  | 0.0036675 |  | 0.36541 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1128.7 | 2021/05/19 20:55:25 | 0
05/19/21 | 64 | 36-UWSIF-UT Glut 1-20210048.38 | 36-UWSIF-UT Glut 1- | 0.813 |  |  | 689 | -29.809 | Carbon | 4 | 43.1301 | 18.433 |  |  | 18.148 |  | 0 | 1.1513 |  | 0.010847 |  | 1.0731 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.7 | 2021/05/19 20:55:25 | 95
05/19/21 | 64 | 36-UWSIF-UT Glut 1-20210048.38 | 36-UWSIF-UT Glut 1- | 0.813 |  |  | 5391 | -43.8 |  | 5 | 52.6756 | 100.799 |  |  | 99.27 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/19 20:55:25 | 95
05/19/21 | 64 | 36-UWSIF-UT Glut 1-20210048.38 | 36-UWSIF-UT Glut 1- | 0.813 |  |  | 5380 | -44.026 |  | 6 | 52.319 | 101.008 |  |  | 99.476 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.2 | 2021/05/19 20:55:25 | 95
05/19/21 | 65 | 36-UWSIF-UT Glut 1-20210048.39 | 36-UWSIF-UT Glut 1- | 0.777 | 6384 | 0.108 |  |  |  | 1 | 15.281 | 117.492 |  |  |  | 116.647 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.9; BGD 29: 31.4; BGD 30: 1031.5 | 2021/05/19 21:05:10 | 0
05/19/21 | 65 | 36-UWSIF-UT Glut 1-20210048.39 | 36-UWSIF-UT Glut 1- | 0.777 | 6390 | 0 |  |  |  | 2 | 15.2892 | 117.61 |  |  |  | 116.764 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.6; BGD 30: 1154.4 | 2021/05/19 21:05:10 | 0
05/19/21 | 65 | 36-UWSIF-UT Glut 1-20210048.39 | 36-UWSIF-UT Glut 1- | 0.777 | 2310 | -2.858 |  |  | Nitrogen | 3 | 9.0249 | 55.938 |  |  |  | 55.537 | 0 |  | 0.72216 |  | 0.0036677 |  | 0.36543 | BGD 28: 40.8; BGD 29: 31.2; BGD 30: 1122.4 | 2021/05/19 21:05:10 | 0
05/19/21 | 65 | 36-UWSIF-UT Glut 1-20210048.39 | 36-UWSIF-UT Glut 1- | 0.777 |  |  | 654 | -29.891 | Carbon | 4 | 43.2535 | 17.516 |  |  | 17.245 |  | 0 | 1.1512 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 9.7 | 2021/05/19 21:05:10 | 95
05/19/21 | 65 | 36-UWSIF-UT Glut 1-20210048.39 | 36-UWSIF-UT Glut 1- | 0.777 |  |  | 5390 | -43.8 |  | 5 | 55.2587 | 100.719 |  |  | 99.191 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/19 21:05:10 | 95
05/19/21 | 65 | 36-UWSIF-UT Glut 1-20210048.39 | 36-UWSIF-UT Glut 1- | 0.777 |  |  | 5383 | -44.021 |  | 6 | 54.7318 | 101.015 |  |  | 99.482 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.2 | 2021/05/19 21:05:10 | 95
05/19/21 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 39-UWSIF-UW Glut 2- | 0.766 | 6387 | 0.103 |  |  |  | 1 | 15.5076 | 117.593 |  |  |  | 116.748 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.9; BGD 29: 31.3; BGD 30: 1029.3 | 2021/05/19 21:14:55 | 0
05/19/21 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 39-UWSIF-UW Glut 2- | 0.766 | 6388 | 0 |  |  |  | 2 | 15.5135 | 117.676 |  |  |  | 116.83 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.9; BGD 29: 30.6; BGD 30: 1152.1 | 2021/05/19 21:14:55 | 0
05/19/21 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 39-UWSIF-UW Glut 2- | 0.766 | 2463 | 29.66 |  |  | Nitrogen | 3 | 9.6519 | 59.681 |  |  |  | 59.239 | 0 |  | 0.74571 |  | 0.0037873 |  | 0.3773 | BGD 28: 40.8; BGD 29: 31.1; BGD 30: 1122.6 | 2021/05/19 21:14:55 | 0
05/19/21 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 39-UWSIF-UW Glut 2- | 0.766 |  |  | 699 | 22.281 | Carbon | 4 | 46.5229 | 18.798 |  |  | 18.497 |  | 0 | 1.2088 |  | 0.011429 |  | 1.13 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/19 21:14:55 | 95
05/19/21 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 39-UWSIF-UW Glut 2- | 0.766 |  |  | 5387 | -43.8 |  | 5 | 56.027 | 100.733 |  |  | 99.204 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/19 21:14:55 | 95
05/19/21 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 39-UWSIF-UW Glut 2- | 0.766 |  |  | 5394 | -44.03 |  | 6 | 55.5158 | 101.016 |  |  | 99.483 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.2 | 2021/05/19 21:14:55 | 95
05/19/21 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 39-UWSIF-UW Glut 2- | 0.757 | 6384 | 0.093 |  |  |  | 1 | 15.6875 | 117.531 |  |  |  | 116.686 | 0 |  | 0.72432 |  | 0.0036785 |  | 0.36651 | BGD 28: 40.9; BGD 29: 31.4; BGD 30: 1028.1 | 2021/05/19 21:24:40 | 0
05/19/21 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 39-UWSIF-UW Glut 2- | 0.757 | 6389 | 0 |  |  |  | 2 | 15.6977 | 117.673 |  |  |  | 116.827 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.6; BGD 30: 1151.5 | 2021/05/19 21:24:40 | 0
05/19/21 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 39-UWSIF-UW Glut 2- | 0.757 | 2434 | 29.59 |  |  | Nitrogen | 3 | 9.6574 | 58.86 |  |  |  | 58.424 | 0 |  | 0.74568 |  | 0.003787 |  | 0.37728 | BGD 28: 40.8; BGD 29: 31.1; BGD 30: 1121.1 | 2021/05/19 21:24:40 | 0
05/19/21 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 39-UWSIF-UW Glut 2- | 0.757 |  |  | 691 | 22.319 | Carbon | 4 | 46.5903 | 18.563 |  |  | 18.266 |  | 0 | 1.2088 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/19 21:24:40 | 95
05/19/21 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 39-UWSIF-UW Glut 2- | 0.757 |  |  | 5392 | -43.8 |  | 5 | 56.6271 | 100.769 |  |  | 99.24 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/19 21:24:40 | 95
05/19/21 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 39-UWSIF-UW Glut 2- | 0.757 |  |  | 5396 | -44.042 |  | 6 | 56.0345 | 101.092 |  |  | 99.559 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/19 21:24:40 | 95
05/19/21 | 68 | 01-UWSIF-Liver-20210048.35 | 01-UWSIF-Liver- | 0.771 | 6388 | 0.113 |  |  |  | 1 | 15.4095 | 117.628 |  |  |  | 116.782 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.9; BGD 29: 31.4; BGD 30: 1027.4 | 2021/05/19 21:34:25 | 0
05/19/21 | 68 | 01-UWSIF-Liver-20210048.35 | 01-UWSIF-Liver- | 0.771 | 6393 | 0 |  |  |  | 2 | 15.411 | 117.649 |  |  |  | 116.803 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.9; BGD 29: 30.6; BGD 30: 1150.2 | 2021/05/19 21:34:25 | 0
05/19/21 | 68 | 01-UWSIF-Liver-20210048.35 | 01-UWSIF-Liver- | 0.771 | 3149 | 8.504 |  |  | Nitrogen | 3 | 11.5505 | 75.773 |  |  |  | 75.224 | 0 |  | 0.7304 |  | 0.0037095 |  | 0.36958 | BGD 28: 40.8; BGD 29: 31.1; BGD 30: 1120.5 | 2021/05/19 21:34:25 | 0
05/19/21 | 68 | 01-UWSIF-Liver-20210048.35 | 01-UWSIF-Liver- | 0.771 |  |  | 798 | -19.47 | Carbon | 4 | 51.6303 | 21.562 |  |  | 21.226 |  | 0 | 1.1626 |  | 0.010963 |  | 1.0844 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.7 | 2021/05/19 21:34:25 | 95
05/19/21 | 68 | 01-UWSIF-Liver-20210048.35 | 01-UWSIF-Liver- | 0.771 |  |  | 5385 | -43.8 |  | 5 | 55.5628 | 100.789 |  |  | 99.26 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.9 | 2021/05/19 21:34:25 | 95
05/19/21 | 68 | 01-UWSIF-Liver-20210048.35 | 01-UWSIF-Liver- | 0.771 |  |  | 5384 | -44.006 |  | 6 | 55.3002 | 100.935 |  |  | 99.404 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.2 | 2021/05/19 21:34:25 | 95
05/19/21 | 69 | 01-UWSIF-Liver-20210048.36 | 01-UWSIF-Liver- | 0.765 | 6388 | 0.096 |  |  |  | 1 | 15.5237 | 117.535 |  |  |  | 116.69 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1030.2 | 2021/05/19 21:44:11 | 0
05/19/21 | 69 | 01-UWSIF-Liver-20210048.36 | 01-UWSIF-Liver- | 0.765 | 6390 | 0 |  |  |  | 2 | 15.5322 | 117.653 |  |  |  | 116.807 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1153.1 | 2021/05/19 21:44:11 | 0
05/19/21 | 69 | 01-UWSIF-Liver-20210048.36 | 01-UWSIF-Liver- | 0.765 | 3126 | 8.477 |  |  | Nitrogen | 3 | 11.5822 | 75.26 |  |  |  | 74.715 | 0 |  | 0.73039 |  | 0.0037094 |  | 0.36957 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1124.4 | 2021/05/19 21:44:11 | 0
05/19/21 | 69 | 01-UWSIF-Liver-20210048.36 | 01-UWSIF-Liver- | 0.765 |  |  | 790 | -19.398 | Carbon | 4 | 51.6403 | 21.355 |  |  | 21.023 |  | 0 | 1.1627 |  | 0.010963 |  | 1.0844 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.7 | 2021/05/19 21:44:11 | 95
05/19/21 | 69 | 01-UWSIF-Liver-20210048.36 | 01-UWSIF-Liver- | 0.765 |  |  | 5398 | -43.8 |  | 5 | 55.8087 | 100.894 |  |  | 99.363 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.9 | 2021/05/19 21:44:11 | 95
05/19/21 | 69 | 01-UWSIF-Liver-20210048.36 | 01-UWSIF-Liver- | 0.765 |  |  | 5391 | -43.996 |  | 6 | 55.6343 | 100.99 |  |  | 99.458 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.3 | 2021/05/19 21:44:11 | 95
05/19/21 | 70 | 20210048.182 | 783 | 0.808 | 6387 | 0.086 |  |  |  | 1 | 14.6962 | 117.514 |  |  |  | 116.669 | 0 |  | 0.72431 |  | 0.0036785 |  | 0.3665 | BGD 28: 41.0; BGD 29: 31.5; BGD 30: 1031.6 | 2021/05/19 21:53:56 | 0
05/19/21 | 70 | 20210048.182 | 783 | 0.808 | 6383 | 0 |  |  |  | 2 | 14.6987 | 117.552 |  |  |  | 116.707 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.7; BGD 30: 1154.4 | 2021/05/19 21:53:56 | 0
05/19/21 | 70 | 20210048.182 | 783 | 0.808 | 3525 | 12.036 |  |  | Nitrogen | 3 | 11.9579 | 84.773 |  |  |  | 84.156 | 0 |  | 0.73297 |  | 0.0037225 |  | 0.37087 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1125.2 | 2021/05/19 21:53:56 | 0
05/19/21 | 70 | 20210048.182 | 783 | 0.808 |  |  | 817 | -29.733 | Carbon | 4 | 50.2286 | 22.098 |  |  | 21.757 |  | 0 | 1.1513 |  | 0.010848 |  | 1.0731 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.9 | 2021/05/19 21:53:56 | 95
05/19/21 | 70 | 20210048.182 | 783 | 0.808 |  |  | 5384 | -43.8 |  | 5 | 53.1294 | 100.724 |  |  | 99.196 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/19 21:53:56 | 95
05/19/21 | 70 | 20210048.182 | 783 | 0.808 |  |  | 5388 | -44.012 |  | 6 | 52.8615 | 100.881 |  |  | 99.35 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/19 21:53:56 | 95
05/19/21 | 71 | 20210048.183 | 785 | 0.801 | 6384 | 0.101 |  |  |  | 1 | 14.8327 | 117.633 |  |  |  | 116.787 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1034.6 | 2021/05/19 22:03:41 | 0
05/19/21 | 71 | 20210048.183 | 785 | 0.801 | 6388 | 0 |  |  |  | 2 | 14.8304 | 117.599 |  |  |  | 116.753 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1157.7 | 2021/05/19 22:03:41 | 0
05/19/21 | 71 | 20210048.183 | 785 | 0.801 | 3042 | 12.168 |  |  | Nitrogen | 3 | 10.8179 | 73.063 |  |  |  | 72.532 | 0 |  | 0.73306 |  | 0.003723 |  | 0.37091 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1126.8 | 2021/05/19 22:03:41 | 0
05/19/21 | 71 | 20210048.183 | 785 | 0.801 |  |  | 875 | -32.422 | Carbon | 4 | 53.6453 | 23.79 |  |  | 23.423 |  | 0 | 1.1484 |  | 0.010818 |  | 1.0702 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.8 | 2021/05/19 22:03:41 | 95
05/19/21 | 71 | 20210048.183 | 785 | 0.801 |  |  | 5392 | -43.8 |  | 5 | 53.5707 | 100.738 |  |  | 99.209 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/19 22:03:41 | 95
05/19/21 | 71 | 20210048.183 | 785 | 0.801 |  |  | 5384 | -44.022 |  | 6 | 53.2927 | 100.899 |  |  | 99.368 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/19 22:03:41 | 95
05/19/21 | 72 | 20210048.184 | 787 | 0.815 | 6388 | 0.091 |  |  |  | 1 | 14.5777 | 117.629 |  |  |  | 116.783 | 0 |  | 0.72432 |  | 0.0036785 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1035.2 | 2021/05/19 22:13:26 | 0
05/19/21 | 72 | 20210048.184 | 787 | 0.815 | 6390 | 0 |  |  |  | 2 | 14.5757 | 117.601 |  |  |  | 116.755 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1157.8 | 2021/05/19 22:13:26 | 0
05/19/21 | 72 | 20210048.184 | 787 | 0.815 | 3420 | 11.486 |  |  | Nitrogen | 3 | 11.602 | 82.242 |  |  |  | 81.644 | 0 |  | 0.73258 |  | 0.0037204 |  | 0.37067 | BGD 28: 40.9; BGD 29: 31.3; BGD 30: 1129.1 | 2021/05/19 22:13:26 | 0
05/19/21 | 72 | 20210048.184 | 787 | 0.815 |  |  | 845 | -32.886 | Carbon | 4 | 51.2212 | 22.913 |  |  | 22.559 |  | 0 | 1.1479 |  | 0.010812 |  | 1.0697 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.9 | 2021/05/19 22:13:26 | 95
05/19/21 | 72 | 20210048.184 | 787 | 0.815 |  |  | 5388 | -43.8 |  | 5 | 52.5111 | 100.82 |  |  | 99.29 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/19 22:13:26 | 95
05/19/21 | 72 | 20210048.184 | 787 | 0.815 |  |  | 5387 | -43.991 |  | 6 | 52.1721 | 101.019 |  |  | 99.486 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/19 22:13:26 | 95
05/19/21 | 73 | 20210048.185 | 789 | 0.805 | 6389 | 0.11 |  |  |  | 1 | 14.7521 | 117.531 |  |  |  | 116.686 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1037.3 | 2021/05/19 22:23:12 | 0
05/19/21 | 73 | 20210048.185 | 789 | 0.805 | 6393 | 0 |  |  |  | 2 | 14.7597 | 117.643 |  |  |  | 116.797 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1160.6 | 2021/05/19 22:23:12 | 0
05/19/21 | 73 | 20210048.185 | 789 | 0.805 | 3737 | 10.734 |  |  | Nitrogen | 3 | 12.4948 | 89.822 |  |  |  | 89.169 | 0 |  | 0.73204 |  | 0.0037177 |  | 0.37039 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1132.2 | 2021/05/19 22:23:12 | 0
05/19/21 | 73 | 20210048.185 | 789 | 0.805 |  |  | 798 | -30.796 | Carbon | 4 | 49.4472 | 21.561 |  |  | 21.228 |  | 0 | 1.1502 |  | 0.010836 |  | 1.072 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/19 22:23:12 | 95
05/19/21 | 73 | 20210048.185 | 789 | 0.805 |  |  | 5392 | -43.8 |  | 5 | 53.1121 | 100.85 |  |  | 99.319 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.0 | 2021/05/19 22:23:12 | 95
05/19/21 | 73 | 20210048.185 | 789 | 0.805 |  |  | 5387 | -44.018 |  | 6 | 52.9626 | 100.937 |  |  | 99.405 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.4 | 2021/05/19 22:23:12 | 95
05/19/21 | 74 | 20210048.186 | 791 | 0.848 | 6389 | 0.107 |  |  |  | 1 | 14.001 | 117.484 |  |  |  | 116.639 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1039.3 | 2021/05/19 22:32:57 | 0
05/19/21 | 74 | 20210048.186 | 791 | 0.848 | 6392 | 0 |  |  |  | 2 | 14.0067 | 117.573 |  |  |  | 116.727 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1162.4 | 2021/05/19 22:32:57 | 0
05/19/21 | 74 | 20210048.186 | 791 | 0.848 | 3610 | 10.47 |  |  | Nitrogen | 3 | 11.5861 | 86.819 |  |  |  | 86.189 | 0 |  | 0.73185 |  | 0.0037167 |  | 0.37029 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1133.8 | 2021/05/19 22:32:57 | 0
05/19/21 | 74 | 20210048.186 | 791 | 0.848 |  |  | 840 | -32.021 | Carbon | 4 | 48.974 | 22.76 |  |  | 22.409 |  | 0 | 1.1488 |  | 0.010822 |  | 1.0706 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/19 22:32:57 | 95
05/19/21 | 74 | 20210048.186 | 791 | 0.848 |  |  | 5384 | -43.8 |  | 5 | 50.6059 | 100.735 |  |  | 99.206 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/19 22:32:57 | 95
05/19/21 | 74 | 20210048.186 | 791 | 0.848 |  |  | 5383 | -44.015 |  | 6 | 50.2991 | 100.923 |  |  | 99.392 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/19 22:32:57 | 95
05/19/21 | 75 | 20210048.187 | 793 | 0.83 | 6385 | 0.127 |  |  |  | 1 | 14.3038 | 117.471 |  |  |  | 116.627 | 0 |  | 0.72434 |  | 0.0036787 |  | 0.36652 | BGD 28: 41.2; BGD 29: 31.5; BGD 30: 1041.0 | 2021/05/19 22:42:42 | 0
05/19/21 | 75 | 20210048.187 | 793 | 0.83 | 6379 | 0 |  |  |  | 2 | 14.3039 | 117.472 |  |  |  | 116.627 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1164.1 | 2021/05/19 22:42:42 | 0
05/19/21 | 75 | 20210048.187 | 793 | 0.83 | 3535 | 12.036 |  |  | Nitrogen | 3 | 11.6599 | 84.969 |  |  |  | 84.351 | 0 |  | 0.73296 |  | 0.0037225 |  | 0.37087 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 1135.5 | 2021/05/19 22:42:42 | 0
05/19/21 | 75 | 20210048.187 | 793 | 0.83 |  |  | 851 | -32.26 | Carbon | 4 | 50.6404 | 23.116 |  |  | 22.759 |  | 0 | 1.1486 |  | 0.010819 |  | 1.0704 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/19 22:42:42 | 95
05/19/21 | 75 | 20210048.187 | 793 | 0.83 |  |  | 5385 | -43.8 |  | 5 | 51.6442 | 100.771 |  |  | 99.241 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/19 22:42:42 | 95
05/19/21 | 75 | 20210048.187 | 793 | 0.83 |  |  | 5379 | -44.029 |  | 6 | 51.5678 | 100.817 |  |  | 99.287 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/19 22:42:42 | 95
05/19/21 | 76 | 20210048.188 | 795 | 0.845 | 6380 | 0.105 |  |  |  | 1 | 14.0403 | 117.323 |  |  |  | 116.479 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1041.8 | 2021/05/19 22:52:27 | 0
05/19/21 | 76 | 20210048.188 | 795 | 0.845 | 6381 | 0 |  |  |  | 2 | 14.0504 | 117.479 |  |  |  | 116.634 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1164.6 | 2021/05/19 22:52:27 | 0
05/19/21 | 76 | 20210048.188 | 795 | 0.845 | 3770 | 11.855 |  |  | Nitrogen | 3 | 11.9714 | 90.577 |  |  |  | 89.918 | 0 |  | 0.73284 |  | 0.0037218 |  | 0.3708 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1136.5 | 2021/05/19 22:52:27 | 0
05/19/21 | 76 | 20210048.188 | 795 | 0.845 |  |  | 867 | -32.324 | Carbon | 4 | 50.5011 | 23.576 |  |  | 23.212 |  | 0 | 1.1485 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/19 22:52:27 | 95
05/19/21 | 76 | 20210048.188 | 795 | 0.845 |  |  | 5376 | -43.8 |  | 5 | 50.9894 | 100.61 |  |  | 99.083 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 22:52:27 | 95
05/19/21 | 76 | 20210048.188 | 795 | 0.845 |  |  | 5383 | -44.019 |  | 6 | 50.5456 | 100.882 |  |  | 99.351 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 22:52:27 | 95
05/19/21 | 77 | 20210048.189 | 797 | 0.826 | 6385 | 0.121 |  |  |  | 1 | 14.3666 | 117.374 |  |  |  | 116.53 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36652 | BGD 28: 41.2; BGD 29: 31.6; BGD 30: 1043.8 | 2021/05/19 23:02:13 | 0
05/19/21 | 77 | 20210048.189 | 797 | 0.826 | 6383 | 0 |  |  |  | 2 | 14.3722 | 117.458 |  |  |  | 116.614 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.8; BGD 30: 1166.7 | 2021/05/19 23:02:13 | 0
05/19/21 | 77 | 20210048.189 | 797 | 0.826 | 3642 | 12.351 |  |  | Nitrogen | 3 | 11.9656 | 87.564 |  |  |  | 86.927 | 0 |  | 0.7332 |  | 0.0037236 |  | 0.37098 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1138.2 | 2021/05/19 23:02:13 | 0
05/19/21 | 77 | 20210048.189 | 797 | 0.826 |  |  | 845 | -32.25 | Carbon | 4 | 50.4986 | 22.889 |  |  | 22.536 |  | 0 | 1.1486 |  | 0.01082 |  | 1.0704 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 23:02:13 | 95
05/19/21 | 77 | 20210048.189 | 797 | 0.826 |  |  | 5385 | -43.8 |  | 5 | 52.1769 | 100.601 |  |  | 99.075 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/19 23:02:13 | 95
05/19/21 | 77 | 20210048.189 | 797 | 0.826 |  |  | 5381 | -44.006 |  | 6 | 51.8244 | 100.812 |  |  | 99.283 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 23:02:13 | 95
05/19/21 | 78 | 20210048.190 | 799 | 0.845 | 6370 | 0.111 |  |  |  | 1 | 14.0418 | 117.346 |  |  |  | 116.502 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1043.8 | 2021/05/19 23:11:58 | 0
05/19/21 | 78 | 20210048.190 | 799 | 0.845 | 6370 | 0 |  |  |  | 2 | 14.0414 | 117.34 |  |  |  | 116.496 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1167.0 | 2021/05/19 23:11:58 | 0
05/19/21 | 78 | 20210048.190 | 799 | 0.845 | 3824 | 10.523 |  |  | Nitrogen | 3 | 12.0836 | 91.833 |  |  |  | 91.166 | 0 |  | 0.73189 |  | 0.0037169 |  | 0.37031 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1138.2 | 2021/05/19 23:11:58 | 0
05/19/21 | 78 | 20210048.190 | 799 | 0.845 |  |  | 812 | -30.995 | Carbon | 4 | 47.7987 | 21.963 |  |  | 21.624 |  | 0 | 1.15 |  | 0.010834 |  | 1.0718 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 23:11:58 | 95
05/19/21 | 78 | 20210048.190 | 799 | 0.845 |  |  | 5381 | -43.8 |  | 5 | 51.0333 | 100.583 |  |  | 99.057 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 23:11:58 | 95
05/19/21 | 78 | 20210048.190 | 799 | 0.845 |  |  | 5378 | -44.024 |  | 6 | 50.6941 | 100.791 |  |  | 99.262 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.5 | 2021/05/19 23:11:58 | 95
05/19/21 | 79 | 20210048.191 | 801 | 0.848 | 6368 | 0.102 |  |  |  | 1 | 13.9811 | 117.176 |  |  |  | 116.333 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1045.3 | 2021/05/19 23:21:43 | 0
05/19/21 | 79 | 20210048.191 | 801 | 0.848 | 6377 | 0 |  |  |  | 2 | 13.9909 | 117.327 |  |  |  | 116.484 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1168.0 | 2021/05/19 23:21:43 | 0
05/19/21 | 79 | 20210048.191 | 801 | 0.848 | 3465 | 11.878 |  |  | Nitrogen | 3 | 11.2367 | 83.132 |  |  |  | 82.527 | 0 |  | 0.73286 |  | 0.0037219 |  | 0.37081 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1138.5 | 2021/05/19 23:21:43 | 0
05/19/21 | 79 | 20210048.191 | 801 | 0.848 |  |  | 908 | -33.069 | Carbon | 4 | 52.2135 | 24.75 |  |  | 24.368 |  | 0 | 1.1477 |  | 0.010811 |  | 1.0695 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 23:21:43 | 95
05/19/21 | 79 | 20210048.191 | 801 | 0.848 |  |  | 5382 | -43.8 |  | 5 | 50.8832 | 100.564 |  |  | 99.038 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 23:21:43 | 95
05/19/21 | 79 | 20210048.191 | 801 | 0.848 |  |  | 5375 | -44.011 |  | 6 | 50.508 | 100.795 |  |  | 99.266 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 23:21:43 | 95
05/19/21 | 80 | 20210048.192 | 803 | 0.833 | 6377 | 0.094 |  |  |  | 1 | 14.2386 | 117.263 |  |  |  | 116.42 | 0 |  | 0.72434 |  | 0.0036785 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1045.3 | 2021/05/19 23:31:28 | 0
05/19/21 | 80 | 20210048.192 | 803 | 0.833 | 6379 | 0 |  |  |  | 2 | 14.248 | 117.406 |  |  |  | 116.562 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1168.7 | 2021/05/19 23:31:28 | 0
05/19/21 | 80 | 20210048.192 | 803 | 0.833 | 3778 | 11.932 |  |  | Nitrogen | 3 | 12.172 | 90.886 |  |  |  | 90.225 | 0 |  | 0.73291 |  | 0.0037221 |  | 0.37083 | BGD 28: 40.9; BGD 29: 31.3; BGD 30: 1140.9 | 2021/05/19 23:31:28 | 0
05/19/21 | 80 | 20210048.192 | 803 | 0.833 |  |  | 824 | -31.12 | Carbon | 4 | 49.1308 | 22.336 |  |  | 21.991 |  | 0 | 1.1499 |  | 0.010832 |  | 1.0716 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 23:31:28 | 95
05/19/21 | 80 | 20210048.192 | 803 | 0.833 |  |  | 5388 | -43.8 |  | 5 | 51.7348 | 100.604 |  |  | 99.077 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 23:31:28 | 95
05/19/21 | 80 | 20210048.192 | 803 | 0.833 |  |  | 5373 | -44.01 |  | 6 | 51.4656 | 100.766 |  |  | 99.237 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 23:31:28 | 95
05/19/21 | 81 | 20210048.193 | 805 | 0.843 | 6369 | 0.11 |  |  |  | 1 | 14.0681 | 117.238 |  |  |  | 116.395 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1046.5 | 2021/05/19 23:41:13 | 0
05/19/21 | 81 | 20210048.193 | 805 | 0.843 | 6376 | 0 |  |  |  | 2 | 14.0764 | 117.366 |  |  |  | 116.522 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1169.4 | 2021/05/19 23:41:13 | 0
05/19/21 | 81 | 20210048.193 | 805 | 0.843 | 3974 | 12.232 |  |  | Nitrogen | 3 | 12.4327 | 95.509 |  |  |  | 94.814 | 0 |  | 0.73311 |  | 0.0037232 |  | 0.37094 | BGD 28: 41.0; BGD 29: 31.2; BGD 30: 1142.3 | 2021/05/19 23:41:13 | 0
05/19/21 | 81 | 20210048.193 | 805 | 0.843 |  |  | 845 | -31.869 | Carbon | 4 | 49.5151 | 22.91 |  |  | 22.556 |  | 0 | 1.149 |  | 0.010824 |  | 1.0708 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/19 23:41:13 | 95
05/19/21 | 81 | 20210048.193 | 805 | 0.843 |  |  | 5385 | -43.8 |  | 5 | 50.9474 | 100.71 |  |  | 99.181 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 23:41:13 | 95
05/19/21 | 81 | 20210048.193 | 805 | 0.843 |  |  | 5374 | -44.03 |  | 6 | 50.7754 | 100.815 |  |  | 99.285 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 23:41:13 | 95
05/19/21 | 82 | 20210048.194 | 807 | 0.818 | 6377 | 0.111 |  |  |  | 1 | 14.504 | 117.328 |  |  |  | 116.484 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1049.1 | 2021/05/19 23:50:59 | 0
05/19/21 | 82 | 20210048.194 | 807 | 0.818 | 6373 | 0 |  |  |  | 2 | 14.5061 | 117.358 |  |  |  | 116.514 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1172.0 | 2021/05/19 23:50:59 | 0
05/19/21 | 82 | 20210048.194 | 807 | 0.818 | 3464 | 12.681 |  |  | Nitrogen | 3 | 11.6588 | 83.231 |  |  |  | 82.625 | 0 |  | 0.73343 |  | 0.0037248 |  | 0.3711 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1143.2 | 2021/05/19 23:50:59 | 0
05/19/21 | 82 | 20210048.194 | 807 | 0.818 |  |  | 841 | -32.939 | Carbon | 4 | 50.8199 | 22.789 |  |  | 22.437 |  | 0 | 1.1478 |  | 0.010812 |  | 1.0696 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/19 23:50:59 | 95
05/19/21 | 82 | 20210048.194 | 807 | 0.818 |  |  | 5389 | -43.8 |  | 5 | 52.5411 | 100.688 |  |  | 99.16 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/19 23:50:59 | 95
05/19/21 | 82 | 20210048.194 | 807 | 0.818 |  |  | 5383 | -44.016 |  | 6 | 52.3587 | 100.796 |  |  | 99.267 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/19 23:50:59 | 95
05/20/21 | 83 | 20210048.195 | 809 | 0.803 | 6370 | 0.116 |  |  |  | 1 | 14.7788 | 117.383 |  |  |  | 116.539 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1048.4 | 2021/05/20 00:00:44 | 0
05/20/21 | 83 | 20210048.195 | 809 | 0.803 | 6382 | 0 |  |  |  | 2 | 14.7823 | 117.435 |  |  |  | 116.591 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.8; BGD 30: 1171.6 | 2021/05/20 00:00:44 | 0
05/20/21 | 83 | 20210048.195 | 809 | 0.803 | 3549 | 12.029 |  |  | Nitrogen | 3 | 12.117 | 85.621 |  |  |  | 84.998 | 0 |  | 0.73297 |  | 0.0037224 |  | 0.37086 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1142.0 | 2021/05/20 00:00:44 | 0
05/20/21 | 83 | 20210048.195 | 809 | 0.803 |  |  | 819 | -31.108 | Carbon | 4 | 50.6605 | 22.165 |  |  | 21.823 |  | 0 | 1.1499 |  | 0.010832 |  | 1.0716 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/20 00:00:44 | 95
05/20/21 | 83 | 20210048.195 | 809 | 0.803 |  |  | 5380 | -43.8 |  | 5 | 53.6207 | 100.631 |  |  | 99.104 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/20 00:00:44 | 95
05/20/21 | 83 | 20210048.195 | 809 | 0.803 |  |  | 5381 | -44.014 |  | 6 | 53.1006 | 100.933 |  |  | 99.402 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/20 00:00:44 | 95
05/20/21 | 84 | 20210048.196 | 811 | 0.843 | 6376 | 0.118 |  |  |  | 1 | 14.0802 | 117.425 |  |  |  | 116.581 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1049.2 | 2021/05/20 00:10:30 | 0
05/20/21 | 84 | 20210048.196 | 811 | 0.843 | 6384 | 0 |  |  |  | 2 | 14.0844 | 117.489 |  |  |  | 116.644 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1172.1 | 2021/05/20 00:10:30 | 0
05/20/21 | 84 | 20210048.196 | 811 | 0.843 | 3949 | 11.651 |  |  | Nitrogen | 3 | 12.3652 | 94.723 |  |  |  | 94.034 | 0 |  | 0.73269 |  | 0.0037211 |  | 0.37073 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1143.8 | 2021/05/20 00:10:30 | 0
05/20/21 | 84 | 20210048.196 | 811 | 0.843 |  |  | 859 | -31.101 | Carbon | 4 | 50.1683 | 23.302 |  |  | 22.942 |  | 0 | 1.1499 |  | 0.010833 |  | 1.0716 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/20 00:10:30 | 95
05/20/21 | 84 | 20210048.196 | 811 | 0.843 |  |  | 5380 | -43.8 |  | 5 | 51.0822 | 100.627 |  |  | 99.1 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/20 00:10:30 | 95
05/20/21 | 84 | 20210048.196 | 811 | 0.843 |  |  | 5371 | -44.013 |  | 6 | 50.853 | 100.768 |  |  | 99.239 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/20 00:10:30 | 95
05/20/21 | 85 | 20210048.197 | 813 | 0.831 | 6371 | 0.106 |  |  |  | 1 | 14.271 | 117.235 |  |  |  | 116.392 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1050.5 | 2021/05/20 00:20:15 | 0
05/20/21 | 85 | 20210048.197 | 813 | 0.831 | 6376 | 0 |  |  |  | 2 | 14.2804 | 117.378 |  |  |  | 116.534 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.2; BGD 29: 30.8; BGD 30: 1173.4 | 2021/05/20 00:20:15 | 0
05/20/21 | 85 | 20210048.197 | 813 | 0.831 | 3664 | 9.272 |  |  | Nitrogen | 3 | 11.9393 | 88.051 |  |  |  | 87.412 | 0 |  | 0.73097 |  | 0.0037123 |  | 0.36986 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1144.9 | 2021/05/20 00:20:15 | 0
05/20/21 | 85 | 20210048.197 | 813 | 0.831 |  |  | 829 | -30.328 | Carbon | 4 | 49.4885 | 22.475 |  |  | 22.128 |  | 0 | 1.1507 |  | 0.010841 |  | 1.0725 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/20 00:20:15 | 95
05/20/21 | 85 | 20210048.197 | 813 | 0.831 |  |  | 5382 | -43.8 |  | 5 | 51.9122 | 100.572 |  |  | 99.045 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/20 00:20:15 | 95
05/20/21 | 85 | 20210048.197 | 813 | 0.831 |  |  | 5367 | -44.008 |  | 6 | 51.5358 | 100.799 |  |  | 99.269 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/20 00:20:15 | 95
05/20/21 | 86 | 20210048.198 | 815 | 0.814 | 6366 | 0.122 |  |  |  | 1 | 14.5658 | 117.186 |  |  |  | 116.343 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36652 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1050.0 | 2021/05/20 00:30:01 | 0
05/20/21 | 86 | 20210048.198 | 815 | 0.814 | 6361 | 0 |  |  |  | 2 | 14.5697 | 117.245 |  |  |  | 116.402 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1173.3 | 2021/05/20 00:30:01 | 0
05/20/21 | 86 | 20210048.198 | 815 | 0.814 | 2726 | 11.121 |  |  | Nitrogen | 3 | 9.7687 | 65.375 |  |  |  | 64.9 | 0 |  | 0.73231 |  | 0.0037191 |  | 0.37053 | BGD 28: 40.9; BGD 29: 31.2; BGD 30: 1140.9 | 2021/05/20 00:30:01 | 0
05/20/21 | 86 | 20210048.198 | 815 | 0.814 |  |  | 940 | -34.145 | Carbon | 4 | 55.9403 | 25.697 |  |  | 25.301 |  | 0 | 1.1465 |  | 0.010799 |  | 1.0683 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/20 00:30:01 | 95
05/20/21 | 86 | 20210048.198 | 815 | 0.814 |  |  | 5377 | -43.8 |  | 5 | 53.1428 | 100.485 |  |  | 98.96 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/20 00:30:01 | 95
05/20/21 | 86 | 20210048.198 | 815 | 0.814 |  |  | 5386 | -44.02 |  | 6 | 52.5257 | 100.849 |  |  | 99.319 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.6; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/20 00:30:01 | 95
05/20/21 | 87 | 20210048.199 | 817 | 0.802 | 6369 | 0.092 |  |  |  | 1 | 14.7846 | 117.199 |  |  |  | 116.357 | 0 |  | 0.72434 |  | 0.0036785 |  | 0.36651 | BGD 28: 41.1; BGD 29: 31.5; BGD 30: 1046.7 | 2021/05/20 00:39:46 | 0
05/20/21 | 87 | 20210048.199 | 817 | 0.802 | 6364 | 0 |  |  |  | 2 | 14.7933 | 117.327 |  |  |  | 116.483 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.1; BGD 29: 30.7; BGD 30: 1170.2 | 2021/05/20 00:39:46 | 0
05/20/21 | 87 | 20210048.199 | 817 | 0.802 | 3589 | 12.201 |  |  | Nitrogen | 3 | 12.2149 | 86.456 |  |  |  | 85.827 | 0 |  | 0.73311 |  | 0.0037231 |  | 0.37093 | BGD 28: 41.0; BGD 29: 31.3; BGD 30: 1143.7 | 2021/05/20 00:39:46 | 0
05/20/21 | 87 | 20210048.199 | 817 | 0.802 |  |  | 801 | -32.236 | Carbon | 4 | 49.9108 | 21.714 |  |  | 21.379 |  | 0 | 1.1487 |  | 0.01082 |  | 1.0704 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 10.3 | 2021/05/20 00:39:46 | 95
05/20/21 | 87 | 20210048.199 | 817 | 0.802 |  |  | 5375 | -43.8 |  | 5 | 54.0378 | 100.426 |  |  | 98.902 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/20 00:39:46 | 95
05/20/21 | 87 | 20210048.199 | 817 | 0.802 |  |  | 5369 | -44.036 |  | 6 | 53.7164 | 100.614 |  |  | 99.088 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/20 00:39:46 | 95
05/20/21 | 88 | 20210048.200 | 819 | 0.825 | 6357 | 0.108 |  |  |  | 1 | 14.3682 | 117.136 |  |  |  | 116.294 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.0; BGD 29: 31.4; BGD 30: 1047.9 | 2021/05/20 00:49:31 | 0
05/20/21 | 88 | 20210048.200 | 819 | 0.825 | 6363 | 0 |  |  |  | 2 | 14.3704 | 117.169 |  |  |  | 116.326 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.7; BGD 30: 1171.6 | 2021/05/20 00:49:31 | 0
05/20/21 | 88 | 20210048.200 | 819 | 0.825 | 3726 | 11.63 |  |  | Nitrogen | 3 | 12.1665 | 89.548 |  |  |  | 88.897 | 0 |  | 0.73269 |  | 0.003721 |  | 0.37072 | BGD 28: 40.8; BGD 29: 31.2; BGD 30: 1142.6 | 2021/05/20 00:49:31 | 0
05/20/21 | 88 | 20210048.200 | 819 | 0.825 |  |  | 830 | -32.44 | Carbon | 4 | 49.8485 | 22.475 |  |  | 22.129 |  | 0 | 1.1484 |  | 0.010818 |  | 1.0702 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/20 00:49:31 | 95
05/20/21 | 88 | 20210048.200 | 819 | 0.825 |  |  | 5366 | -43.8 |  | 5 | 52.5478 | 100.417 |  |  | 98.893 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/20 00:49:31 | 95
05/20/21 | 88 | 20210048.200 | 819 | 0.825 |  |  | 5356 | -44.028 |  | 6 | 52.3516 | 100.535 |  |  | 99.009 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/20 00:49:31 | 95
05/20/21 | 89 | 20210048.201 | 821 | 0.831 | 6361 | 0.093 |  |  |  | 1 | 14.2531 | 116.963 |  |  |  | 116.122 | 0 |  | 0.72433 |  | 0.0036785 |  | 0.36651 | BGD 28: 41.0; BGD 29: 31.4; BGD 30: 1049.1 | 2021/05/20 00:59:17 | 0
05/20/21 | 89 | 20210048.201 | 821 | 0.831 | 6357 | 0 |  |  |  | 2 | 14.2613 | 117.087 |  |  |  | 116.245 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.7; BGD 30: 1172.3 | 2021/05/20 00:59:17 | 0
05/20/21 | 89 | 20210048.201 | 821 | 0.831 | 3450 | 11.06 |  |  | Nitrogen | 3 | 11.4412 | 82.874 |  |  |  | 82.272 | 0 |  | 0.73228 |  | 0.0037189 |  | 0.37051 | BGD 28: 40.8; BGD 29: 31.2; BGD 30: 1144.1 | 2021/05/20 00:59:17 | 0
05/20/21 | 89 | 20210048.201 | 821 | 0.831 |  |  | 887 | -31.805 | Carbon | 4 | 52.3415 | 24.173 |  |  | 23.8 |  | 0 | 1.1491 |  | 0.010825 |  | 1.0709 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/20 00:59:17 | 95
05/20/21 | 89 | 20210048.201 | 821 | 0.831 |  |  | 5371 | -43.8 |  | 5 | 52.3021 | 100.335 |  |  | 98.813 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/20 00:59:17 | 95
05/20/21 | 89 | 20210048.201 | 821 | 0.831 |  |  | 5363 | -44.005 |  | 6 | 51.7464 | 100.672 |  |  | 99.144 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/20 00:59:17 | 95
05/20/21 | 90 | 20210048.202 | 823 | 0.81 | 6358 | 0.108 |  |  |  | 1 | 14.6219 | 116.953 |  |  |  | 116.112 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 41.0; BGD 29: 31.4; BGD 30: 1048.9 | 2021/05/20 01:09:02 | 0
05/20/21 | 90 | 20210048.202 | 823 | 0.81 | 6360 | 0 |  |  |  | 2 | 14.6304 | 117.078 |  |  |  | 116.236 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 40.0; BGD 29: 30.6; BGD 30: 1171.9 | 2021/05/20 01:09:02 | 0
05/20/21 | 90 | 20210048.202 | 823 | 0.81 | 3555 | 13.059 |  |  | Nitrogen | 3 | 11.9968 | 85.465 |  |  |  | 84.842 | 0 |  | 0.73372 |  | 0.0037262 |  | 0.37124 | BGD 28: 40.8; BGD 29: 31.2; BGD 30: 1143.7 | 2021/05/20 01:09:02 | 0
05/20/21 | 90 | 20210048.202 | 823 | 0.81 |  |  | 819 | -32.344 | Carbon | 4 | 50.281 | 22.198 |  |  | 21.855 |  | 0 | 1.1485 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/20 01:09:02 | 95
05/20/21 | 90 | 20210048.202 | 823 | 0.81 |  |  | 5358 | -43.8 |  | 5 | 53.6295 | 100.352 |  |  | 98.829 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/20 01:09:02 | 95
05/20/21 | 90 | 20210048.202 | 823 | 0.81 |  |  | 5361 | -44.036 |  | 6 | 53.26 | 100.571 |  |  | 99.045 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/20 01:09:02 | 95
05/20/21 | 91 | 20210048.203 | 825 | 0.837 | 6353 | 0.131 |  |  |  | 1 | 14.1516 | 116.975 |  |  |  | 116.134 | 0 |  | 0.72434 |  | 0.0036787 |  | 0.36652 | BGD 28: 40.9; BGD 29: 31.4; BGD 30: 1049.3 | 2021/05/20 01:18:47 | 0
05/20/21 | 91 | 20210048.203 | 825 | 0.837 | 6355 | 0 |  |  |  | 2 | 14.1494 | 116.941 |  |  |  | 116.1 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.9; BGD 29: 30.6; BGD 30: 1172.3 | 2021/05/20 01:18:47 | 0
05/20/21 | 91 | 20210048.203 | 825 | 0.837 | 3640 | 13.358 |  |  | Nitrogen | 3 | 11.7821 | 87.284 |  |  |  | 86.648 | 0 |  | 0.73392 |  | 0.0037273 |  | 0.37135 | BGD 28: 40.8; BGD 29: 31.1; BGD 30: 1143.8 | 2021/05/20 01:18:47 | 0
05/20/21 | 91 | 20210048.203 | 825 | 0.837 |  |  | 848 | -32.148 | Carbon | 4 | 49.9702 | 22.969 |  |  | 22.615 |  | 0 | 1.1487 |  | 0.010821 |  | 1.0705 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.2 | 2021/05/20 01:18:47 | 95
05/20/21 | 91 | 20210048.203 | 825 | 0.837 |  |  | 5366 | -43.8 |  | 5 | 51.8905 | 100.358 |  |  | 98.835 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/20 01:18:47 | 95
05/20/21 | 91 | 20210048.203 | 825 | 0.837 |  |  | 5358 | -44.022 |  | 6 | 51.6459 | 100.507 |  |  | 98.982 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/20 01:18:47 | 95
05/20/21 | 92 | 20210048.204 | 827 | 0.836 | 6346 | 0.113 |  |  |  | 1 | 14.1593 | 116.834 |  |  |  | 115.994 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.9; BGD 29: 31.4; BGD 30: 1049.4 | 2021/05/20 01:28:33 | 0
05/20/21 | 92 | 20210048.204 | 827 | 0.836 | 6356 | 0 |  |  |  | 2 | 14.1663 | 116.94 |  |  |  | 116.099 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.9; BGD 29: 30.6; BGD 30: 1172.8 | 2021/05/20 01:28:33 | 0
05/20/21 | 92 | 20210048.204 | 827 | 0.836 | 3737 | 12.014 |  |  | Nitrogen | 3 | 12.0169 | 89.663 |  |  |  | 89.011 | 0 |  | 0.73296 |  | 0.0037224 |  | 0.37086 | BGD 28: 40.7; BGD 29: 31.1; BGD 30: 1144.0 | 2021/05/20 01:28:33 | 0
05/20/21 | 92 | 20210048.204 | 827 | 0.836 |  |  | 844 | -32.394 | Carbon | 4 | 49.8803 | 22.88 |  |  | 22.527 |  | 0 | 1.1485 |  | 0.010818 |  | 1.0702 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.3 | 2021/05/20 01:28:33 | 95
05/20/21 | 92 | 20210048.204 | 827 | 0.836 |  |  | 5369 | -43.8 |  | 5 | 52.1713 | 100.224 |  |  | 98.703 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.2 | 2021/05/20 01:28:33 | 95
05/20/21 | 92 | 20210048.204 | 827 | 0.836 |  |  | 5361 | -44.003 |  | 6 | 51.839 | 100.427 |  |  | 98.903 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/20 01:28:33 | 95
05/20/21 | 93 | 20210048.204 | Replicate H8 | 0.832 | 6347 | 0.103 |  |  |  | 1 | 14.2256 | 116.807 |  |  |  | 115.967 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.9; BGD 29: 31.4; BGD 30: 1050.4 | 2021/05/20 01:38:18 | 0
05/20/21 | 93 | 20210048.204 | Replicate H8 | 0.832 | 6348 | 0 |  |  |  | 2 | 14.2284 | 116.851 |  |  |  | 116.01 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.9; BGD 29: 30.6; BGD 30: 1173.5 | 2021/05/20 01:38:18 | 0
05/20/21 | 93 | 20210048.204 | Replicate H8 | 0.832 | 3676 | 12.002 |  |  | Nitrogen | 3 | 11.9453 | 88.268 |  |  |  | 87.626 | 0 |  | 0.73297 |  | 0.0037223 |  | 0.37085 | BGD 28: 40.7; BGD 29: 31.1; BGD 30: 1144.7 | 2021/05/20 01:38:18 | 0
05/20/21 | 93 | 20210048.204 | Replicate H8 | 0.832 |  |  | 842 | -32.393 | Carbon | 4 | 50.053 | 22.841 |  |  | 22.488 |  | 0 | 1.1484 |  | 0.010818 |  | 1.0702 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.3 | 2021/05/20 01:38:18 | 95
05/20/21 | 93 | 20210048.204 | Replicate H8 | 0.832 |  |  | 5364 | -43.8 |  | 5 | 52.4984 | 100.177 |  |  | 98.657 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.3 | 2021/05/20 01:38:18 | 95
05/20/21 | 93 | 20210048.204 | Replicate H8 | 0.832 |  |  | 5351 | -44.018 |  | 6 | 52.1872 | 100.367 |  |  | 98.844 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.7 | 2021/05/20 01:38:18 | 95
05/20/21 | 94 | 36-UWSIF-UT Glut 1-20210048.311 | 36-UWSIF-UT Glut 1- | 0.755 | 6342 | 0.112 |  |  |  | 1 | 15.6671 | 116.681 |  |  |  | 115.842 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.8; BGD 29: 31.3; BGD 30: 1050.4 | 2021/05/20 01:48:03 | 0
05/20/21 | 94 | 36-UWSIF-UT Glut 1-20210048.311 | 36-UWSIF-UT Glut 1- | 0.755 | 6341 | 0 |  |  |  | 2 | 15.6788 | 116.84 |  |  |  | 116 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 1173.7 | 2021/05/20 01:48:03 | 0
05/20/21 | 94 | 36-UWSIF-UT Glut 1-20210048.311 | 36-UWSIF-UT Glut 1- | 0.755 | 2214 | -2.888 |  |  | Nitrogen | 3 | 8.9735 | 53.66 |  |  |  | 53.275 | 0 |  | 0.72217 |  | 0.0036676 |  | 0.36542 | BGD 28: 40.6; BGD 29: 31.0; BGD 30: 1142.2 | 2021/05/20 01:48:03 | 0
05/20/21 | 94 | 36-UWSIF-UT Glut 1-20210048.311 | 36-UWSIF-UT Glut 1- | 0.755 |  |  | 632 | -29.902 | Carbon | 4 | 43.2751 | 16.936 |  |  | 16.675 |  | 0 | 1.1513 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 10.1 | 2021/05/20 01:48:03 | 95
05/20/21 | 94 | 36-UWSIF-UT Glut 1-20210048.311 | 36-UWSIF-UT Glut 1- | 0.755 |  |  | 5360 | -43.8 |  | 5 | 58.061 | 100.061 |  |  | 98.543 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/20 01:48:03 | 95
05/20/21 | 94 | 36-UWSIF-UT Glut 1-20210048.311 | 36-UWSIF-UT Glut 1- | 0.755 |  |  | 5349 | -44.022 |  | 6 | 57.7747 | 100.22 |  |  | 98.7 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.6 | 2021/05/20 01:48:03 | 95
05/20/21 | 95 | 36-UWSIF-UT Glut 1-20210048.312 | 36-UWSIF-UT Glut 1- | 0.756 | 6344 | 0.096 |  |  |  | 1 | 15.6509 | 116.742 |  |  |  | 115.903 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.6; BGD 29: 31.1; BGD 30: 1043.2 | 2021/05/20 01:57:49 | 0
05/20/21 | 95 | 36-UWSIF-UT Glut 1-20210048.312 | 36-UWSIF-UT Glut 1- | 0.756 | 6346 | 0 |  |  |  | 2 | 15.6549 | 116.796 |  |  |  | 115.956 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.6; BGD 29: 30.3; BGD 30: 1166.9 | 2021/05/20 01:57:49 | 0
05/20/21 | 95 | 36-UWSIF-UT Glut 1-20210048.312 | 36-UWSIF-UT Glut 1- | 0.756 | 2218 | -2.934 |  |  | Nitrogen | 3 | 8.9579 | 53.633 |  |  |  | 53.248 | 0 |  | 0.72214 |  | 0.0036674 |  | 0.3654 | BGD 28: 40.5; BGD 29: 30.8; BGD 30: 1136.2 | 2021/05/20 01:57:49 | 0
05/20/21 | 95 | 36-UWSIF-UT Glut 1-20210048.312 | 36-UWSIF-UT Glut 1- | 0.756 |  |  | 635 | -29.848 | Carbon | 4 | 43.2663 | 16.959 |  |  | 16.697 |  | 0 | 1.1513 |  | 0.010847 |  | 1.073 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/20 01:57:49 | 95
05/20/21 | 95 | 36-UWSIF-UT Glut 1-20210048.312 | 36-UWSIF-UT Glut 1- | 0.756 |  |  | 5363 | -43.8 |  | 5 | 57.9066 | 100.105 |  |  | 98.585 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.0 | 2021/05/20 01:57:49 | 95
05/20/21 | 95 | 36-UWSIF-UT Glut 1-20210048.312 | 36-UWSIF-UT Glut 1- | 0.756 |  |  | 5343 | -44.024 |  | 6 | 57.5551 | 100.3 |  |  | 98.778 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.5 | 2021/05/20 01:57:49 | 95
05/20/21 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 39-UWSIF-UW Glut 2- | 0.76 | 6338 | 0.117 |  |  |  | 1 | 15.5535 | 116.536 |  |  |  | 115.698 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.5; BGD 29: 31.0; BGD 30: 1039.4 | 2021/05/20 02:07:34 | 0
05/20/21 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 39-UWSIF-UW Glut 2- | 0.76 | 6344 | 0 |  |  |  | 2 | 15.571 | 116.776 |  |  |  | 115.936 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.6; BGD 29: 30.3; BGD 30: 1162.8 | 2021/05/20 02:07:34 | 0
05/20/21 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 39-UWSIF-UW Glut 2- | 0.76 | 2426 | 29.631 |  |  | Nitrogen | 3 | 9.6025 | 58.733 |  |  |  | 58.298 | 0 |  | 0.74572 |  | 0.0037872 |  | 0.37729 | BGD 28: 40.4; BGD 29: 30.8; BGD 30: 1133.7 | 2021/05/20 02:07:34 | 0
05/20/21 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 39-UWSIF-UW Glut 2- | 0.76 |  |  | 691 | 22.346 | Carbon | 4 | 46.4799 | 18.599 |  |  | 18.301 |  | 0 | 1.2089 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/20 02:07:34 | 95
05/20/21 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 39-UWSIF-UW Glut 2- | 0.76 |  |  | 5352 | -43.8 |  | 5 | 57.6868 | 100.057 |  |  | 98.539 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.1 | 2021/05/20 02:07:34 | 95
05/20/21 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 39-UWSIF-UW Glut 2- | 0.76 |  |  | 5346 | -44.002 |  | 6 | 57.3186 | 100.263 |  |  | 98.741 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.5 | 2021/05/20 02:07:34 | 95
05/20/21 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 39-UWSIF-UW Glut 2- | 0.706 | 6337 | 0.119 |  |  |  | 1 | 16.7492 | 116.613 |  |  |  | 115.774 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36652 | BGD 28: 40.5; BGD 29: 31.1; BGD 30: 1038.2 | 2021/05/20 02:17:19 | 0
05/20/21 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 39-UWSIF-UW Glut 2- | 0.706 | 6338 | 0 |  |  |  | 2 | 16.7533 | 116.665 |  |  |  | 115.826 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.6; BGD 29: 30.3; BGD 30: 1161.4 | 2021/05/20 02:17:19 | 0
05/20/21 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 39-UWSIF-UW Glut 2- | 0.706 | 2249 | 29.565 |  |  | Nitrogen | 3 | 9.7247 | 54.526 |  |  |  | 54.122 | 0 |  | 0.74567 |  | 0.0037869 |  | 0.37727 | BGD 28: 40.4; BGD 29: 30.9; BGD 30: 1130.7 | 2021/05/20 02:17:19 | 0
05/20/21 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 39-UWSIF-UW Glut 2- | 0.706 |  |  | 641 | 22.26 | Carbon | 4 | 46.8673 | 17.193 |  |  | 16.917 |  | 0 | 1.2088 |  | 0.011429 |  | 1.13 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 10.0 | 2021/05/20 02:17:19 | 95
05/20/21 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 39-UWSIF-UW Glut 2- | 0.706 |  |  | 5348 | -43.8 |  | 5 | 62.1256 | 100.043 |  |  | 98.525 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.0 | 2021/05/20 02:17:19 | 95
05/20/21 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 39-UWSIF-UW Glut 2- | 0.706 |  |  | 5348 | -44.043 |  | 6 | 61.6944 | 100.267 |  |  | 98.746 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.5 | 2021/05/20 02:17:19 | 95
05/20/21 | 98 | 01-UWSIF-Liver-20210048.37 | 01-UWSIF-Liver- | 0.73 | 6336 | 0.102 |  |  |  | 1 | 16.194 | 116.553 |  |  |  | 115.715 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.5; BGD 29: 31.0; BGD 30: 1035.3 | 2021/05/20 02:27:05 | 0
05/20/21 | 98 | 01-UWSIF-Liver-20210048.37 | 01-UWSIF-Liver- | 0.73 | 6337 | 0 |  |  |  | 2 | 16.1961 | 116.581 |  |  |  | 115.743 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.5; BGD 29: 30.2; BGD 30: 1158.7 | 2021/05/20 02:27:05 | 0
05/20/21 | 98 | 01-UWSIF-Liver-20210048.37 | 01-UWSIF-Liver- | 0.73 | 2946 | 8.508 |  |  | Nitrogen | 3 | 11.6216 | 71.061 |  |  |  | 70.545 | 0 |  | 0.73042 |  | 0.0037095 |  | 0.36958 | BGD 28: 40.3; BGD 29: 30.7; BGD 30: 1128.2 | 2021/05/20 02:27:05 | 0
05/20/21 | 98 | 01-UWSIF-Liver-20210048.37 | 01-UWSIF-Liver- | 0.73 |  |  | 751 | -19.579 | Carbon | 4 | 51.9096 | 20.272 |  |  | 19.956 |  | 0 | 1.1626 |  | 0.010961 |  | 1.0842 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.1 | 2021/05/20 02:27:05 | 95
05/20/21 | 98 | 01-UWSIF-Liver-20210048.37 | 01-UWSIF-Liver- | 0.73 |  |  | 5354 | -43.8 |  | 5 | 60.0898 | 100.04 |  |  | 98.521 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/20 02:27:05 | 95
05/20/21 | 98 | 01-UWSIF-Liver-20210048.37 | 01-UWSIF-Liver- | 0.73 |  |  | 5339 | -44.027 |  | 6 | 59.8657 | 100.16 |  |  | 98.64 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 9.6 | 2021/05/20 02:27:05 | 95
05/20/21 | 99 | 01-UWSIF-Liver-20210048.38 | 01-UWSIF-Liver- | 0.723 | 6329 | 0.11 |  |  |  | 1 | 16.3541 | 116.596 |  |  |  | 115.758 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.5; BGD 29: 31.1; BGD 30: 1037.0 | 2021/05/20 02:36:51 | 0
05/20/21 | 99 | 01-UWSIF-Liver-20210048.38 | 01-UWSIF-Liver- | 0.723 | 6334 | 0 |  |  |  | 2 | 16.3571 | 116.635 |  |  |  | 115.796 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.6; BGD 29: 30.3; BGD 30: 1160.7 | 2021/05/20 02:36:51 | 0
05/20/21 | 99 | 01-UWSIF-Liver-20210048.38 | 01-UWSIF-Liver- | 0.723 | 2929 | 8.43 |  |  | Nitrogen | 3 | 11.6816 | 70.645 |  |  |  | 70.133 | 0 |  | 0.73037 |  | 0.0037092 |  | 0.36955 | BGD 28: 40.4; BGD 29: 30.8; BGD 30: 1131.0 | 2021/05/20 02:36:51 | 0
05/20/21 | 99 | 01-UWSIF-Liver-20210048.38 | 01-UWSIF-Liver- | 0.723 |  |  | 745 | -19.498 | Carbon | 4 | 52.0494 | 20.098 |  |  | 19.785 |  | 0 | 1.1627 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 10.0 | 2021/05/20 02:36:51 | 95
05/20/21 | 99 | 01-UWSIF-Liver-20210048.38 | 01-UWSIF-Liver- | 0.723 |  |  | 5349 | -43.8 |  | 5 | 60.7294 | 100.009 |  |  | 98.491 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/20 02:36:51 | 95
05/20/21 | 99 | 01-UWSIF-Liver-20210048.38 | 01-UWSIF-Liver- | 0.723 |  |  | 5347 | -44.036 |  | 6 | 60.2657 | 100.255 |  |  | 98.734 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.5 | 2021/05/20 02:36:51 | 95

## Original 4
Date | Row | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Component | Peak Nr | Amt% | Area All | Comment | Preparation | Area 44 | Area 28 | Is Ref _ | rR 45CO2/44CO2 | rR 29N2/28N2 | R 13C/12C | R 15N/14N | AT% 13C/12C | AT% 15N/14N | BGD 28 | BGD 29 | BGD 30 | BGD 44 | BGD 45 | BGD 46 | Time Code | Sample Dilution
05/20/21 | 1 | 36-UWSIF-UT Glut 1-20210048.41 | 36-UWSIF-UT Glut 1- | 0.794 | 6232 | 0.119 |  |  |  | 1 | 14.7601 | 114.734 | JOB: 2021-0048 Tray 4 Nelson |  |  | 113.909 | 0 |  | 0.72424 |  | 0.0036786 |  | 0.36652 | BGD 28: 36.7; BGD 29: 27.4; BGD 30: 830.4 | 2021/05/20 14:40:22 | 0
05/20/21 | 1 | 36-UWSIF-UT Glut 1-20210048.41 | 36-UWSIF-UT Glut 1- | 0.794 | 6237 | 0 |  |  |  | 2 | 14.7625 | 114.767 | JOB: 2021-0048 Tray 4 Nelson |  |  | 113.942 | 1 |  | 0.72415 |  | 0.0036782 |  | 0.36647 | BGD 28: 36.0; BGD 29: 26.8; BGD 30: 946.8 | 2021/05/20 14:40:22 | 0
05/20/21 | 1 | 36-UWSIF-UT Glut 1-20210048.41 | 36-UWSIF-UT Glut 1- | 0.794 | 2344 | -2.749 |  |  | Nitrogen | 3 | 8.8548 | 56.116 | JOB: 2021-0048 Tray 4 Nelson |  |  | 55.714 | 0 |  | 0.72216 |  | 0.0036681 |  | 0.36547 | BGD 28: 37.1; BGD 29: 27.6; BGD 30: 916.4 | 2021/05/20 14:40:22 | 0
05/20/21 | 1 | 36-UWSIF-UT Glut 1-20210048.41 | 36-UWSIF-UT Glut 1- | 0.794 |  |  | 652 | -29.868 | Carbon | 4 | 42.3025 | 17.503 | JOB: 2021-0048 Tray 4 Nelson |  | 17.233 |  | 0 | 1.1506 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 8.6 | 2021/05/20 14:40:22 | 95
05/20/21 | 1 | 36-UWSIF-UT Glut 1-20210048.41 | 36-UWSIF-UT Glut 1- | 0.794 |  |  | 5291 | -43.8 |  | 5 | 57.6362 | 98.613 | JOB: 2021-0048 Tray 4 Nelson |  | 97.118 |  | 1 | 1.1339 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 6.5 | 2021/05/20 14:40:22 | 95
05/20/21 | 1 | 36-UWSIF-UT Glut 1-20210048.41 | 36-UWSIF-UT Glut 1- | 0.794 |  |  | 5286 | -44.159 |  | 6 | 56.8647 | 99.08 | JOB: 2021-0048 Tray 4 Nelson |  | 97.578 |  | 0 | 1.1335 |  | 0.010686 |  | 1.0574 |  | BGD 44: 1.3; BGD 45: 1.5; BGD 46: 7.8 | 2021/05/20 14:40:22 | 95
05/20/21 | 2 | 36-UWSIF-UT Glut 1-20210048.42 | 36-UWSIF-UT Glut 1- | 0.781 | 6237 | 0.123 |  |  |  | 1 | 15.0152 | 114.863 | ANALYST: gh |  |  | 114.037 | 0 |  | 0.72431 |  | 0.0036787 |  | 0.36652 | BGD 28: 38.4; BGD 29: 29.0; BGD 30: 873.9 | 2021/05/20 14:50:07 | 0
05/20/21 | 2 | 36-UWSIF-UT Glut 1-20210048.42 | 36-UWSIF-UT Glut 1- | 0.781 | 6231 | 0 |  |  |  | 2 | 15.0112 | 114.808 | ANALYST: gh |  |  | 113.983 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 37.5; BGD 29: 28.3; BGD 30: 987.8 | 2021/05/20 14:50:07 | 0
05/20/21 | 2 | 36-UWSIF-UT Glut 1-20210048.42 | 36-UWSIF-UT Glut 1- | 0.781 | 2239 | -2.874 |  |  | Nitrogen | 3 | 8.8444 | 54.927 | ANALYST: gh |  |  | 54.533 | 0 |  | 0.72214 |  | 0.0036676 |  | 0.36542 | BGD 28: 38.4; BGD 29: 28.9; BGD 30: 951.1 | 2021/05/20 14:50:07 | 0
05/20/21 | 2 | 36-UWSIF-UT Glut 1-20210048.42 | 36-UWSIF-UT Glut 1- | 0.781 |  |  | 638 | -29.831 | Carbon | 4 | 42.6062 | 17.309 | ANALYST: gh |  | 17.041 |  | 0 | 1.1509 |  | 0.010847 |  | 1.073 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 8.0 | 2021/05/20 14:50:07 | 95
05/20/21 | 2 | 36-UWSIF-UT Glut 1-20210048.42 | 36-UWSIF-UT Glut 1- | 0.781 |  |  | 5277 | -43.8 |  | 5 | 58.5508 | 98.64 | ANALYST: gh |  | 97.144 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 6.3 | 2021/05/20 14:50:07 | 95
05/20/21 | 2 | 36-UWSIF-UT Glut 1-20210048.42 | 36-UWSIF-UT Glut 1- | 0.781 |  |  | 5275 | -44.048 |  | 6 | 58.1484 | 98.88 | ANALYST: gh |  | 97.38 |  | 0 | 1.134 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.3; BGD 45: 1.5; BGD 46: 7.6 | 2021/05/20 14:50:07 | 95
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 | 6215 | 0.099 |  |  |  | 1 | 9.52 | 114.548 | SAMPLE COUNT: 0306 |  |  | 113.725 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.0; BGD 29: 29.5; BGD 30: 873.4 | 2021/05/20 14:59:51 | 0
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 | 6233 | 0 |  |  |  | 2 | 9.52 | 114.664 | SAMPLE COUNT: 0306 |  |  | 113.84 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.0; BGD 29: 28.7; BGD 30: 988.5 | 2021/05/20 14:59:51 | 0
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 | 1339 | -2.815 |  |  | Nitrogen | 3 | 9.52 | 33.177 | SAMPLE COUNT: 0306 |  |  | 32.94 | 0 |  | 0.72219 |  | 0.0036678 |  | 0.36544 | BGD 28: 38.9; BGD 29: 29.3; BGD 30: 951.4 | 2021/05/20 14:59:51 | 0
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 |  |  | 382 | -29.917 | Carbon | 4 | 40.81 | 10.228 | SAMPLE COUNT: 0306 |  | 10.07 |  | 0 | 1.1509 |  | 0.010846 |  | 1.0729 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/20 14:59:51 | 95
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 |  |  | 5271 | -43.8 |  | 5 | 40.81 | 98.559 | SAMPLE COUNT: 0306 |  | 97.064 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.5; BGD 45: 0.6; BGD 46: 6.0 | 2021/05/20 14:59:51 | 95
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 |  |  | 5261 | -44.083 |  | 6 | 40.81 | 98.704 | SAMPLE COUNT: 0306 |  | 97.207 |  | 0 | 1.134 |  | 0.010687 |  | 1.0574 |  | BGD 44: 1.2; BGD 45: 1.5; BGD 46: 7.5 | 2021/05/20 14:59:51 | 95
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 | 6223 | 0.129 |  |  |  | 1 | 9.52 | 114.571 | HE PSI: 1250 |  |  | 113.747 | 0 |  | 0.7243 |  | 0.0036787 |  | 0.36652 | BGD 28: 38.9; BGD 29: 29.4; BGD 30: 871.7 | 2021/05/20 15:09:36 | 0
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 | 6218 | 0 |  |  |  | 2 | 9.52 | 114.54 | HE PSI: 1250 |  |  | 113.716 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.0; BGD 29: 28.7; BGD 30: 987.3 | 2021/05/20 15:09:36 | 0
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 | 2291 | -2.753 |  |  | Nitrogen | 3 | 9.52 | 56.272 | HE PSI: 1250 |  |  | 55.868 | 0 |  | 0.72222 |  | 0.0036681 |  | 0.36547 | BGD 28: 38.9; BGD 29: 29.3; BGD 30: 953.9 | 2021/05/20 15:09:36 | 0
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 |  |  | 650 | -29.767 | Carbon | 4 | 40.81 | 17.694 | HE PSI: 1250 |  | 17.421 |  | 0 | 1.151 |  | 0.010847 |  | 1.0731 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.9 | 2021/05/20 15:09:36 | 95
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 |  |  | 5273 | -43.8 |  | 5 | 40.81 | 98.537 | HE PSI: 1250 |  | 97.042 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.7; BGD 46: 6.2 | 2021/05/20 15:09:36 | 95
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 |  |  | 5262 | -44.044 |  | 6 | 40.81 | 98.654 | HE PSI: 1250 |  | 97.157 |  | 0 | 1.1341 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.3; BGD 45: 1.5; BGD 46: 7.5 | 2021/05/20 15:09:36 | 95
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 | 6216 | 0.113 |  |  |  | 1 | 9.52 | 114.466 | O2 PSI: 350 |  |  | 113.643 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.3; BGD 29: 29.7; BGD 30: 879.7 | 2021/05/20 15:19:21 | 0
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 | 6222 | 0 |  |  |  | 2 | 9.52 | 114.568 | O2 PSI: 350 |  |  | 113.744 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.3; BGD 29: 29.0; BGD 30: 995.5 | 2021/05/20 15:19:21 | 0
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 | 4769 | -2.876 |  |  | Nitrogen | 3 | 9.52 | 115.459 | O2 PSI: 350 |  |  | 114.631 | 0 |  | 0.72214 |  | 0.0036676 |  | 0.36542 | BGD 28: 39.2; BGD 29: 29.6; BGD 30: 964.5 | 2021/05/20 15:19:21 | 0
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 |  |  | 1284 | -30.028 | Carbon | 4 | 40.81 | 36.757 | O2 PSI: 350 |  | 36.19 |  | 0 | 1.1505 |  | 0.010844 |  | 1.0728 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 8.2 | 2021/05/20 15:19:21 | 95
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 |  |  | 5265 | -43.8 |  | 5 | 40.81 | 98.434 | O2 PSI: 350 |  | 96.941 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 6.8 | 2021/05/20 15:19:21 | 95
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 |  |  | 5260 | -44.017 |  | 6 | 40.81 | 98.621 | O2 PSI: 350 |  | 97.125 |  | 0 | 1.134 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.0 | 2021/05/20 15:19:21 | 95
05/20/21 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 39-UWSIF-UW Glut 2- | 0.717 | 6210 | 0.114 |  |  |  | 1 | 17.2195 | 114.428 | MASS 28: 36 |  |  | 113.605 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 899.1 | 2021/05/20 15:29:06 | 0
05/20/21 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 39-UWSIF-UW Glut 2- | 0.717 | 6207 | 0 |  |  |  | 2 | 17.2124 | 114.35 | MASS 28: 36 |  |  | 113.527 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.7; BGD 30: 1014.3 | 2021/05/20 15:29:06 | 0
05/20/21 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 39-UWSIF-UW Glut 2- | 0.717 | 2213 | 29.706 |  |  | Nitrogen | 3 | 9.928 | 54.486 | MASS 28: 36 |  |  | 54.083 | 0 |  | 0.74573 |  | 0.0037875 |  | 0.37732 | BGD 28: 39.7; BGD 29: 30.1; BGD 30: 979.8 | 2021/05/20 15:29:06 | 0
05/20/21 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 39-UWSIF-UW Glut 2- | 0.717 |  |  | 629 | 22.375 | Carbon | 4 | 47.7168 | 17.15 | MASS 28: 36 |  | 16.876 |  | 0 | 1.2085 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/20 15:29:06 | 95
05/20/21 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 39-UWSIF-UW Glut 2- | 0.717 |  |  | 5257 | -43.8 |  | 5 | 62.1956 | 98.362 | MASS 28: 36 |  | 96.87 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.6; BGD 45: 0.8; BGD 46: 6.4 | 2021/05/20 15:29:06 | 95
05/20/21 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 39-UWSIF-UW Glut 2- | 0.717 |  |  | 5252 | -44.042 |  | 6 | 61.9489 | 98.487 | MASS 28: 36 |  | 96.993 |  | 0 | 1.134 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.3; BGD 45: 1.6; BGD 46: 7.7 | 2021/05/20 15:29:06 | 95
05/20/21 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 39-UWSIF-UW Glut 2- | 0.807 | 6200 | 0.119 |  |  |  | 1 | 15.2772 | 114.155 | MASS 29: 27 |  |  | 113.334 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36652 | BGD 28: 39.2; BGD 29: 29.8; BGD 30: 898.4 | 2021/05/20 15:38:51 | 0
05/20/21 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 39-UWSIF-UW Glut 2- | 0.807 | 6208 | 0 |  |  |  | 2 | 15.2875 | 114.284 | MASS 29: 27 |  |  | 113.462 | 1 |  | 0.7242 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.4; BGD 29: 29.2; BGD 30: 1014.0 | 2021/05/20 15:38:51 | 0
05/20/21 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 39-UWSIF-UW Glut 2- | 0.807 | 2510 | 29.772 |  |  | Nitrogen | 3 | 9.764 | 61.582 | MASS 29: 27 |  |  | 61.126 | 0 |  | 0.74576 |  | 0.0037877 |  | 0.37734 | BGD 28: 39.4; BGD 29: 29.8; BGD 30: 981.8 | 2021/05/20 15:38:51 | 0
05/20/21 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 39-UWSIF-UW Glut 2- | 0.807 |  |  | 714 | 22.314 | Carbon | 4 | 47.1096 | 19.492 | MASS 29: 27 |  | 19.18 |  | 0 | 1.2084 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 8.1 | 2021/05/20 15:38:51 | 95
05/20/21 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 39-UWSIF-UW Glut 2- | 0.807 |  |  | 5259 | -43.8 |  | 5 | 55.4419 | 98.258 | MASS 29: 27 |  | 96.767 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.4 | 2021/05/20 15:38:51 | 95
05/20/21 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 39-UWSIF-UW Glut 2- | 0.807 |  |  | 5252 | -44.037 |  | 6 | 55.0821 | 98.463 | MASS 29: 27 |  | 96.969 |  | 0 | 1.134 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 7.8 | 2021/05/20 15:38:51 | 95
05/20/21 | 8 | 01-UWSIF-Liver-20210048.41 | 01-UWSIF-Liver- | 0.751 | 6210 | 0.132 |  |  |  | 1 | 16.4196 | 114.192 | MASS 30: 600 |  |  | 113.371 | 0 |  | 0.72431 |  | 0.0036787 |  | 0.36652 | BGD 28: 39.5; BGD 29: 30.0; BGD 30: 903.7 | 2021/05/20 15:48:36 | 0
05/20/21 | 8 | 01-UWSIF-Liver-20210048.41 | 01-UWSIF-Liver- | 0.751 | 6204 | 0 |  |  |  | 2 | 16.4256 | 114.261 | MASS 30: 600 |  |  | 113.44 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.4; BGD 29: 29.2; BGD 30: 1019.5 | 2021/05/20 15:48:36 | 0
05/20/21 | 8 | 01-UWSIF-Liver-20210048.41 | 01-UWSIF-Liver- | 0.751 | 2961 | 8.514 |  |  | Nitrogen | 3 | 11.9606 | 72.565 | MASS 30: 600 |  |  | 72.039 | 0 |  | 0.73038 |  | 0.0037095 |  | 0.36958 | BGD 28: 39.4; BGD 29: 29.9; BGD 30: 988.8 | 2021/05/20 15:48:36 | 0
05/20/21 | 8 | 01-UWSIF-Liver-20210048.41 | 01-UWSIF-Liver- | 0.751 |  |  | 750 | -19.534 | Carbon | 4 | 52.8711 | 20.574 | MASS 30: 600 |  | 20.254 |  | 0 | 1.1622 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 8.2 | 2021/05/20 15:48:36 | 95
05/20/21 | 8 | 01-UWSIF-Liver-20210048.41 | 01-UWSIF-Liver- | 0.751 |  |  | 5253 | -43.8 |  | 5 | 59.5889 | 98.251 | MASS 30: 600 |  | 96.76 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 6.6 | 2021/05/20 15:48:36 | 95
05/20/21 | 8 | 01-UWSIF-Liver-20210048.41 | 01-UWSIF-Liver- | 0.751 |  |  | 5248 | -44.031 |  | 6 | 59.3073 | 98.401 | MASS 30: 600 |  | 96.908 |  | 0 | 1.134 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 7.8 | 2021/05/20 15:48:36 | 95
05/20/21 | 9 | 01-UWSIF-Liver-20210048.42 | 01-UWSIF-Liver- | 0.809 | 6212 | 0.109 |  |  |  | 1 | 15.2423 | 114.191 |  |  |  | 113.37 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.6; BGD 29: 30.1; BGD 30: 911.5 | 2021/05/20 15:58:21 | 0
05/20/21 | 9 | 01-UWSIF-Liver-20210048.42 | 01-UWSIF-Liver- | 0.809 | 6207 | 0 |  |  |  | 2 | 15.2508 | 114.297 |  |  |  | 113.475 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.6; BGD 29: 29.4; BGD 30: 1027.9 | 2021/05/20 15:58:21 | 0
05/20/21 | 9 | 01-UWSIF-Liver-20210048.42 | 01-UWSIF-Liver- | 0.809 | 3194 | 8.584 |  |  | Nitrogen | 3 | 11.7566 | 78.183 |  |  |  | 77.616 | 0 |  | 0.73043 |  | 0.0037098 |  | 0.36961 | BGD 28: 39.5; BGD 29: 29.9; BGD 30: 996.8 | 2021/05/20 15:58:21 | 0
05/20/21 | 9 | 01-UWSIF-Liver-20210048.42 | 01-UWSIF-Liver- | 0.809 |  |  | 808 | -19.538 | Carbon | 4 | 52.1498 | 22.221 |  |  | 21.875 |  | 0 | 1.1622 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.4 | 2021/05/20 15:58:21 | 95
05/20/21 | 9 | 01-UWSIF-Liver-20210048.42 | 01-UWSIF-Liver- | 0.809 |  |  | 5255 | -43.8 |  | 5 | 55.4023 | 98.202 |  |  | 96.712 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.6 | 2021/05/20 15:58:21 | 95
05/20/21 | 9 | 01-UWSIF-Liver-20210048.42 | 01-UWSIF-Liver- | 0.809 |  |  | 5248 | -44.021 |  | 6 | 55.0474 | 98.405 |  |  | 96.912 |  | 0 | 1.134 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 7.9 | 2021/05/20 15:58:21 | 95
05/20/21 | 10 | 20210048.205 | 829 | 0.844 | 6209 | 0.098 |  |  |  | 1 | 14.6203 | 114.322 |  |  |  | 113.5 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.7; BGD 29: 30.3; BGD 30: 920.1 | 2021/05/20 16:08:06 | 0
05/20/21 | 10 | 20210048.205 | 829 | 0.844 | 6208 | 0 |  |  |  | 2 | 14.6186 | 114.3 |  |  |  | 113.478 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.7; BGD 29: 29.5; BGD 30: 1036.7 | 2021/05/20 16:08:06 | 0
05/20/21 | 10 | 20210048.205 | 829 | 0.844 | 3445 | 11.884 |  |  | Nitrogen | 3 | 11.8895 | 84.016 |  |  |  | 83.405 | 0 |  | 0.73282 |  | 0.0037219 |  | 0.37081 | BGD 28: 39.6; BGD 29: 30.1; BGD 30: 1006.5 | 2021/05/20 16:08:06 | 0
05/20/21 | 10 | 20210048.205 | 829 | 0.844 |  |  | 836 | -31.834 | Carbon | 4 | 51.3996 | 23.036 |  |  | 22.681 |  | 0 | 1.1486 |  | 0.010824 |  | 1.0708 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.4 | 2021/05/20 16:08:06 | 95
05/20/21 | 10 | 20210048.205 | 829 | 0.844 |  |  | 5254 | -43.8 |  | 5 | 52.9768 | 98.278 |  |  | 96.787 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/20 16:08:06 | 95
05/20/21 | 10 | 20210048.205 | 829 | 0.844 |  |  | 5250 | -44.01 |  | 6 | 52.5232 | 98.549 |  |  | 97.054 |  | 0 | 1.134 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.1 | 2021/05/20 16:08:06 | 95
05/20/21 | 11 | 20210048.206 | 831 | 0.807 | 6214 | 0.099 |  |  |  | 1 | 15.2955 | 114.384 |  |  |  | 113.561 | 0 |  | 0.72429 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 928.6 | 2021/05/20 16:17:51 | 0
05/20/21 | 11 | 20210048.206 | 831 | 0.807 | 6211 | 0 |  |  |  | 2 | 15.2963 | 114.394 |  |  |  | 113.571 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.6; BGD 30: 1045.4 | 2021/05/20 16:17:51 | 0
05/20/21 | 11 | 20210048.206 | 831 | 0.807 | 2736 | 11.556 |  |  | Nitrogen | 3 | 10.4266 | 66.794 |  |  |  | 66.308 | 0 |  | 0.73259 |  | 0.0037207 |  | 0.37069 | BGD 28: 39.7; BGD 29: 30.2; BGD 30: 1013.8 | 2021/05/20 16:17:51 | 0
05/20/21 | 11 | 20210048.206 | 831 | 0.807 |  |  | 880 | -34.433 | Carbon | 4 | 56.0251 | 24.325 |  |  | 23.95 |  | 0 | 1.1458 |  | 0.010795 |  | 1.068 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 8.3 | 2021/05/20 16:17:51 | 95
05/20/21 | 11 | 20210048.206 | 831 | 0.807 |  |  | 5261 | -43.8 |  | 5 | 55.3839 | 98.291 |  |  | 96.8 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.7 | 2021/05/20 16:17:51 | 95
05/20/21 | 11 | 20210048.206 | 831 | 0.807 |  |  | 5259 | -44 |  | 6 | 55.0554 | 98.478 |  |  | 96.984 |  | 0 | 1.1341 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.1 | 2021/05/20 16:17:51 | 95
05/20/21 | 12 | 20210048.206 | Replicate A11 | 0.815 | 6210 | 0.106 |  |  |  | 1 | 15.1454 | 114.384 |  |  |  | 113.561 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.9; BGD 29: 30.4; BGD 30: 933.1 | 2021/05/20 16:27:36 | 0
05/20/21 | 12 | 20210048.206 | Replicate A11 | 0.815 | 6225 | 0 |  |  |  | 2 | 15.1643 | 114.622 |  |  |  | 113.798 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.6; BGD 30: 1050.4 | 2021/05/20 16:27:36 | 0
05/20/21 | 12 | 20210048.206 | Replicate A11 | 0.815 | 2690 | 11.582 |  |  | Nitrogen | 3 | 10.1533 | 65.416 |  |  |  | 64.94 | 0 |  | 0.73262 |  | 0.0037208 |  | 0.3707 | BGD 28: 39.8; BGD 29: 30.2; BGD 30: 1019.4 | 2021/05/20 16:27:36 | 0
05/20/21 | 12 | 20210048.206 | Replicate A11 | 0.815 |  |  | 867 | -34.482 | Carbon | 4 | 54.7604 | 23.91 |  |  | 23.542 |  | 0 | 1.1458 |  | 0.010795 |  | 1.0679 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.4 | 2021/05/20 16:27:36 | 95
05/20/21 | 12 | 20210048.206 | Replicate A11 | 0.815 |  |  | 5262 | -43.8 |  | 5 | 54.7423 | 98.347 |  |  | 96.855 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.8 | 2021/05/20 16:27:36 | 95
05/20/21 | 12 | 20210048.206 | Replicate A11 | 0.815 |  |  | 5259 | -44.007 |  | 6 | 54.2289 | 98.642 |  |  | 97.146 |  | 0 | 1.1341 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.0 | 2021/05/20 16:27:36 | 95
05/20/21 | 13 | 20210048.207 | 833 | 0.839 | 6226 | 0.117 |  |  |  | 1 | 14.7325 | 114.649 |  |  |  | 113.824 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 938.0 | 2021/05/20 16:37:21 | 0
05/20/21 | 13 | 20210048.207 | 833 | 0.839 | 6232 | 0 |  |  |  | 2 | 14.736 | 114.695 |  |  |  | 113.87 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1056.3 | 2021/05/20 16:37:21 | 0
05/20/21 | 13 | 20210048.207 | 833 | 0.839 | 3641 | 11.888 |  |  | Nitrogen | 3 | 12.4568 | 88.876 |  |  |  | 88.229 | 0 |  | 0.73283 |  | 0.0037219 |  | 0.37081 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1027.5 | 2021/05/20 16:37:21 | 0
05/20/21 | 13 | 20210048.207 | 833 | 0.839 |  |  | 837 | -31.448 | Carbon | 4 | 51.6665 | 23.014 |  |  | 22.658 |  | 0 | 1.1492 |  | 0.010829 |  | 1.0713 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.6 | 2021/05/20 16:37:21 | 95
05/20/21 | 13 | 20210048.207 | 833 | 0.839 |  |  | 5258 | -43.8 |  | 5 | 53.1193 | 98.381 |  |  | 96.889 |  | 1 | 1.1343 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/20 16:37:21 | 95
05/20/21 | 13 | 20210048.207 | 833 | 0.839 |  |  | 5247 | -44.02 |  | 6 | 53.0671 | 98.412 |  |  | 96.919 |  | 0 | 1.1341 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.1 | 2021/05/20 16:37:21 | 95
05/20/21 | 14 | 20210048.208 | 835 | 0.828 | 6225 | 0.127 |  |  |  | 1 | 14.9175 | 114.51 |  |  |  | 113.687 | 0 |  | 0.72431 |  | 0.0036787 |  | 0.36652 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 947.0 | 2021/05/20 16:47:06 | 0
05/20/21 | 14 | 20210048.208 | 835 | 0.828 | 6215 | 0 |  |  |  | 2 | 14.9083 | 114.393 |  |  |  | 113.57 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.7; BGD 30: 1065.0 | 2021/05/20 16:47:06 | 0
05/20/21 | 14 | 20210048.208 | 835 | 0.828 | 3312 | 12.12 |  |  | Nitrogen | 3 | 11.7691 | 80.751 |  |  |  | 80.164 | 0 |  | 0.733 |  | 0.0037228 |  | 0.3709 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1035.4 | 2021/05/20 16:47:06 | 0
05/20/21 | 14 | 20210048.208 | 835 | 0.828 |  |  | 833 | -32.786 | Carbon | 4 | 52.2186 | 22.937 |  |  | 22.583 |  | 0 | 1.1477 |  | 0.010814 |  | 1.0698 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.6 | 2021/05/20 16:47:06 | 95
05/20/21 | 14 | 20210048.208 | 835 | 0.828 |  |  | 5265 | -43.8 |  | 5 | 53.7442 | 98.428 |  |  | 96.935 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 6.9 | 2021/05/20 16:47:06 | 95
05/20/21 | 14 | 20210048.208 | 835 | 0.828 |  |  | 5265 | -44.003 |  | 6 | 53.3503 | 98.658 |  |  | 97.161 |  | 0 | 1.1341 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.2 | 2021/05/20 16:47:06 | 95
05/20/21 | 15 | 20210048.209 | 837 | 0.831 | 6227 | 0.109 |  |  |  | 1 | 14.8696 | 114.588 |  |  |  | 113.764 | 0 |  | 0.7243 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.1; BGD 29: 30.6; BGD 30: 952.8 | 2021/05/20 16:56:51 | 0
05/20/21 | 15 | 20210048.209 | 837 | 0.831 | 6232 | 0 |  |  |  | 2 | 14.8773 | 114.687 |  |  |  | 113.863 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1070.8 | 2021/05/20 16:56:51 | 0
05/20/21 | 15 | 20210048.209 | 837 | 0.831 | 3319 | 12.522 |  |  | Nitrogen | 3 | 11.7199 | 80.689 |  |  |  | 80.102 | 0 |  | 0.73329 |  | 0.0037243 |  | 0.37104 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1041.4 | 2021/05/20 16:56:51 | 0
05/20/21 | 15 | 20210048.209 | 837 | 0.831 |  |  | 849 | -33.494 | Carbon | 4 | 52.7087 | 23.328 |  |  | 22.968 |  | 0 | 1.147 |  | 0.010806 |  | 1.069 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/20 16:56:51 | 95
05/20/21 | 15 | 20210048.209 | 837 | 0.831 |  |  | 5264 | -43.8 |  | 5 | 53.575 | 98.414 |  |  | 96.921 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.0 | 2021/05/20 16:56:51 | 95
05/20/21 | 15 | 20210048.209 | 837 | 0.831 |  |  | 5263 | -44.013 |  | 6 | 52.8393 | 98.843 |  |  | 97.344 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.3 | 2021/05/20 16:56:51 | 95
05/20/21 | 16 | 20210048.210 | 839 | 0.805 | 6239 | 0.155 |  |  |  | 1 | 15.3556 | 114.66 |  |  |  | 113.835 | 0 |  | 0.72432 |  | 0.0036788 |  | 0.36653 | BGD 28: 40.1; BGD 29: 30.6; BGD 30: 957.7 | 2021/05/20 17:06:36 | 0
05/20/21 | 16 | 20210048.210 | 839 | 0.805 | 6234 | 0 |  |  |  | 2 | 15.3607 | 114.724 |  |  |  | 113.899 | 1 |  | 0.72421 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.9; BGD 30: 1076.2 | 2021/05/20 17:06:36 | 0
05/20/21 | 16 | 20210048.210 | 839 | 0.805 | 3182 | 11.818 |  |  | Nitrogen | 3 | 11.723 | 77.38 |  |  |  | 76.818 | 0 |  | 0.73277 |  | 0.0037217 |  | 0.37079 | BGD 28: 40.0; BGD 29: 30.4; BGD 30: 1045.6 | 2021/05/20 17:06:36 | 0
05/20/21 | 16 | 20210048.210 | 839 | 0.805 |  |  | 874 | -32.492 | Carbon | 4 | 55.6107 | 24.007 |  |  | 23.637 |  | 0 | 1.1481 |  | 0.010817 |  | 1.0701 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.8 | 2021/05/20 17:06:36 | 95
05/20/21 | 16 | 20210048.210 | 839 | 0.805 |  |  | 5260 | -43.8 |  | 5 | 55.4597 | 98.326 |  |  | 96.834 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.0 | 2021/05/20 17:06:36 | 95
05/20/21 | 16 | 20210048.210 | 839 | 0.805 |  |  | 5247 | -44.027 |  | 6 | 55.3466 | 98.39 |  |  | 96.898 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/20 17:06:36 | 95
05/20/21 | 17 | 20210048.211 | 841 | 0.825 | 6215 | 0.104 |  |  |  | 1 | 14.9621 | 114.388 |  |  |  | 113.565 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.2; BGD 29: 30.6; BGD 30: 960.9 | 2021/05/20 17:16:21 | 0
05/20/21 | 17 | 20210048.211 | 841 | 0.825 | 6208 | 0 |  |  |  | 2 | 14.9643 | 114.415 |  |  |  | 113.593 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.9; BGD 30: 1079.6 | 2021/05/20 17:16:21 | 0
05/20/21 | 17 | 20210048.211 | 841 | 0.825 | 3319 | 11.895 |  |  | Nitrogen | 3 | 11.8191 | 80.817 |  |  |  | 80.23 | 0 |  | 0.73285 |  | 0.003722 |  | 0.37082 | BGD 28: 40.0; BGD 29: 30.4; BGD 30: 1049.3 | 2021/05/20 17:16:21 | 0
05/20/21 | 17 | 20210048.211 | 841 | 0.825 |  |  | 845 | -31.71 | Carbon | 4 | 52.9805 | 23.264 |  |  | 22.905 |  | 0 | 1.149 |  | 0.010826 |  | 1.071 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.8 | 2021/05/20 17:16:21 | 95
05/20/21 | 17 | 20210048.211 | 841 | 0.825 |  |  | 5247 | -43.8 |  | 5 | 54.4711 | 98.118 |  |  | 96.629 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.0 | 2021/05/20 17:16:21 | 95
05/20/21 | 17 | 20210048.211 | 841 | 0.825 |  |  | 5238 | -44.014 |  | 6 | 54.2545 | 98.245 |  |  | 96.754 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/20 17:16:21 | 95
05/20/21 | 18 | 20210048.212 | 843 | 0.822 | 6204 | 0.113 |  |  |  | 1 | 15.0012 | 114.19 |  |  |  | 113.368 | 0 |  | 0.72432 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.2; BGD 29: 30.7; BGD 30: 964.7 | 2021/05/20 17:26:06 | 0
05/20/21 | 18 | 20210048.212 | 843 | 0.822 | 6201 | 0 |  |  |  | 2 | 15.0049 | 114.237 |  |  |  | 113.415 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.2; BGD 29: 29.9; BGD 30: 1083.2 | 2021/05/20 17:26:06 | 0
05/20/21 | 18 | 20210048.212 | 843 | 0.822 | 3437 | 11.733 |  |  | Nitrogen | 3 | 12.1399 | 83.381 |  |  |  | 82.775 | 0 |  | 0.73273 |  | 0.0037214 |  | 0.37076 | BGD 28: 40.0; BGD 29: 30.4; BGD 30: 1053.4 | 2021/05/20 17:26:06 | 0
05/20/21 | 18 | 20210048.212 | 843 | 0.822 |  |  | 814 | -32.307 | Carbon | 4 | 51.424 | 22.276 |  |  | 21.932 |  | 0 | 1.1484 |  | 0.010819 |  | 1.0703 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/20 17:26:06 | 95
05/20/21 | 18 | 20210048.212 | 843 | 0.822 |  |  | 5237 | -43.8 |  | 5 | 54.9925 | 97.928 |  |  | 96.442 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/20 17:26:06 | 95
05/20/21 | 18 | 20210048.212 | 843 | 0.822 |  |  | 5229 | -44.022 |  | 6 | 54.6995 | 98.1 |  |  | 96.612 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.4 | 2021/05/20 17:26:06 | 95
05/20/21 | 19 | 20210048.213 | 845 | 0.802 | 6187 | 0.116 |  |  |  | 1 | 15.3505 | 113.883 |  |  |  | 113.064 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.2; BGD 29: 30.7; BGD 30: 967.5 | 2021/05/20 17:35:51 | 0
05/20/21 | 19 | 20210048.213 | 845 | 0.802 | 6190 | 0 |  |  |  | 2 | 15.3542 | 113.929 |  |  |  | 113.11 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.9; BGD 30: 1086.2 | 2021/05/20 17:35:51 | 0
05/20/21 | 19 | 20210048.213 | 845 | 0.802 | 3635 | 10.58 |  |  | Nitrogen | 3 | 12.933 | 87.938 |  |  |  | 87.299 | 0 |  | 0.73191 |  | 0.0037171 |  | 0.37034 | BGD 28: 40.0; BGD 29: 30.4; BGD 30: 1057.6 | 2021/05/20 17:35:51 | 0
05/20/21 | 19 | 20210048.213 | 845 | 0.802 |  |  | 755 | -31.063 | Carbon | 4 | 49.3584 | 20.496 |  |  | 20.179 |  | 0 | 1.1497 |  | 0.010833 |  | 1.0717 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/20 17:35:51 | 95
05/20/21 | 19 | 20210048.213 | 845 | 0.802 |  |  | 5225 | -43.8 |  | 5 | 56.6888 | 97.741 |  |  | 96.258 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/20 17:35:51 | 95
05/20/21 | 19 | 20210048.213 | 845 | 0.802 |  |  | 5219 | -44.016 |  | 6 | 56.3851 | 97.916 |  |  | 96.431 |  | 0 | 1.1341 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.4 | 2021/05/20 17:35:51 | 95
05/20/21 | 20 | 20210048.214 | 002A | 0.827 | 6179 | 0.11 |  |  |  | 1 | 14.8679 | 113.647 |  |  |  | 112.83 | 0 |  | 0.72431 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.1; BGD 29: 30.6; BGD 30: 971.5 | 2021/05/20 17:45:36 | 0
05/20/21 | 20 | 20210048.214 | 002A | 0.827 | 6183 | 0 |  |  |  | 2 | 14.8865 | 113.884 |  |  |  | 113.065 | 1 |  | 0.72423 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.9; BGD 30: 1090.2 | 2021/05/20 17:45:36 | 0
05/20/21 | 20 | 20210048.214 | 002A | 0.827 | 2421 | 5.741 |  |  | Nitrogen | 3 | 9.2197 | 59.166 |  |  |  | 58.738 | 0 |  | 0.72839 |  | 0.0036993 |  | 0.36857 | BGD 28: 40.0; BGD 29: 30.4; BGD 30: 1058.3 | 2021/05/20 17:45:36 | 0
05/20/21 | 20 | 20210048.214 | 002A | 0.827 |  |  | 713 | -35.04 | Carbon | 4 | 45.8158 | 19.411 |  |  | 19.112 |  | 0 | 1.1453 |  | 0.010788 |  | 1.0673 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/20 17:45:36 | 95
05/20/21 | 20 | 20210048.214 | 002A | 0.827 |  |  | 5227 | -43.8 |  | 5 | 54.9196 | 97.774 |  |  | 96.291 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.1 | 2021/05/20 17:45:36 | 95
05/20/21 | 20 | 20210048.214 | 002A | 0.827 |  |  | 5219 | -44.017 |  | 6 | 54.6108 | 97.957 |  |  | 96.471 |  | 0 | 1.1341 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.3 | 2021/05/20 17:45:36 | 95
05/20/21 | 21 | 20210048.215 | 002B | 0.844 | 6181 | 0.126 |  |  |  | 1 | 14.5907 | 113.937 |  |  |  | 113.117 | 0 |  | 0.72431 |  | 0.0036787 |  | 0.36652 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 971.0 | 2021/05/20 17:55:21 | 0
05/20/21 | 21 | 20210048.215 | 002B | 0.844 | 6192 | 0 |  |  |  | 2 | 14.5922 | 113.956 |  |  |  | 113.136 | 1 |  | 0.72422 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.8; BGD 30: 1090.6 | 2021/05/20 17:55:21 | 0
05/20/21 | 21 | 20210048.215 | 002B | 0.844 | 2514 | 5.722 |  |  | Nitrogen | 3 | 9.2896 | 61.209 |  |  |  | 60.766 | 0 |  | 0.72836 |  | 0.0036992 |  | 0.36856 | BGD 28: 40.0; BGD 29: 30.4; BGD 30: 1058.3 | 2021/05/20 17:55:21 | 0
05/20/21 | 21 | 20210048.215 | 002B | 0.844 |  |  | 733 | -34.858 | Carbon | 4 | 45.9288 | 19.967 |  |  | 19.659 |  | 0 | 1.1455 |  | 0.01079 |  | 1.0675 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/20 17:55:21 | 95
05/20/21 | 21 | 20210048.215 | 002B | 0.844 |  |  | 5231 | -43.8 |  | 5 | 53.6483 | 97.874 |  |  | 96.389 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.1 | 2021/05/20 17:55:21 | 95
05/20/21 | 21 | 20210048.215 | 002B | 0.844 |  |  | 5229 | -44.006 |  | 6 | 53.4173 | 98.014 |  |  | 96.527 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.4 | 2021/05/20 17:55:21 | 95
05/20/21 | 22 | 20210048.216 | 002C | 0.812 | 6188 | 0.116 |  |  |  | 1 | 15.1722 | 114.017 |  |  |  | 113.197 | 0 |  | 0.72433 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 972.9 | 2021/05/20 18:05:06 | 0
05/20/21 | 22 | 20210048.216 | 002C | 0.812 | 6191 | 0 |  |  |  | 2 | 15.1661 | 113.941 |  |  |  | 113.122 | 1 |  | 0.72424 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.7; BGD 30: 1092.2 | 2021/05/20 18:05:06 | 0
05/20/21 | 22 | 20210048.216 | 002C | 0.812 | 2470 | 5.615 |  |  | Nitrogen | 3 | 9.5175 | 60.142 |  |  |  | 59.707 | 0 |  | 0.72831 |  | 0.0036989 |  | 0.36852 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1060.8 | 2021/05/20 18:05:06 | 0
05/20/21 | 22 | 20210048.216 | 002C | 0.812 |  |  | 714 | -35.111 | Carbon | 4 | 46.6696 | 19.415 |  |  | 19.116 |  | 0 | 1.1453 |  | 0.010788 |  | 1.0673 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/20 18:05:06 | 95
05/20/21 | 22 | 20210048.216 | 002C | 0.812 |  |  | 5238 | -43.8 |  | 5 | 55.7577 | 97.877 |  |  | 96.392 |  | 1 | 1.1344 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.1 | 2021/05/20 18:05:06 | 95
05/20/21 | 22 | 20210048.216 | 002C | 0.812 |  |  | 5229 | -44.013 |  | 6 | 55.4623 | 98.049 |  |  | 96.561 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.4 | 2021/05/20 18:05:06 | 95
05/20/21 | 23 | 20210048.217 | 004A | 0.83 | 6186 | 0.093 |  |  |  | 1 | 14.8364 | 113.931 |  |  |  | 113.111 | 0 |  | 0.72432 |  | 0.0036785 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 974.0 | 2021/05/20 18:14:51 | 0
05/20/21 | 23 | 20210048.217 | 004A | 0.83 | 6192 | 0 |  |  |  | 2 | 14.8447 | 114.038 |  |  |  | 113.218 | 1 |  | 0.72425 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.7; BGD 30: 1094.1 | 2021/05/20 18:14:51 | 0
05/20/21 | 23 | 20210048.217 | 004A | 0.83 | 2824 | 8.532 |  |  | Nitrogen | 3 | 10.3565 | 68.611 |  |  |  | 68.114 | 0 |  | 0.73043 |  | 0.0037096 |  | 0.36959 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1062.7 | 2021/05/20 18:14:51 | 0
05/20/21 | 23 | 20210048.217 | 004A | 0.83 |  |  | 822 | -35.118 | Carbon | 4 | 51.389 | 22.536 |  |  | 22.189 |  | 0 | 1.1453 |  | 0.010788 |  | 1.0672 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.8 | 2021/05/20 18:14:51 | 95
05/20/21 | 23 | 20210048.217 | 004A | 0.83 |  |  | 5227 | -43.8 |  | 5 | 54.628 | 97.83 |  |  | 96.345 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/20 18:14:51 | 95
05/20/21 | 23 | 20210048.217 | 004A | 0.83 |  |  | 5224 | -44.017 |  | 6 | 54.2773 | 98.038 |  |  | 96.551 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.5 | 2021/05/20 18:14:51 | 95
05/20/21 | 24 | 20210048.218 | 004B | 0.821 | 6188 | 0.112 |  |  |  | 1 | 15.0074 | 114.036 |  |  |  | 113.216 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.1; BGD 29: 30.6; BGD 30: 978.2 | 2021/05/20 18:24:36 | 0
05/20/21 | 24 | 20210048.218 | 004B | 0.821 | 6199 | 0 |  |  |  | 2 | 15.014 | 114.12 |  |  |  | 113.3 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.8; BGD 30: 1098.3 | 2021/05/20 18:24:36 | 0
05/20/21 | 24 | 20210048.218 | 004B | 0.821 | 2407 | 8.536 |  |  | Nitrogen | 3 | 9.189 | 58.411 |  |  |  | 57.988 | 0 |  | 0.73044 |  | 0.0037096 |  | 0.36959 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1065.0 | 2021/05/20 18:24:36 | 0
05/20/21 | 24 | 20210048.218 | 004B | 0.821 |  |  | 860 | -36.132 | Carbon | 4 | 53.8271 | 23.601 |  |  | 23.238 |  | 0 | 1.1442 |  | 0.010776 |  | 1.0661 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 8.9 | 2021/05/20 18:24:36 | 95
05/20/21 | 24 | 20210048.218 | 004B | 0.821 |  |  | 5239 | -43.8 |  | 5 | 55.0777 | 97.918 |  |  | 96.432 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/20 18:24:36 | 95
05/20/21 | 24 | 20210048.218 | 004B | 0.821 |  |  | 5223 | -44.018 |  | 6 | 54.8539 | 98.049 |  |  | 96.561 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.5 | 2021/05/20 18:24:36 | 95
05/20/21 | 25 | 20210048.219 | 004C | 0.818 | 6197 | 0.113 |  |  |  | 1 | 15.0636 | 114.052 |  |  |  | 113.232 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.1; BGD 29: 30.6; BGD 30: 979.2 | 2021/05/20 18:34:21 | 0
05/20/21 | 25 | 20210048.219 | 004C | 0.818 | 6198 | 0 |  |  |  | 2 | 15.0657 | 114.078 |  |  |  | 113.258 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.8; BGD 30: 1099.8 | 2021/05/20 18:34:21 | 0
05/20/21 | 25 | 20210048.219 | 004C | 0.818 | 2046 | 7.962 |  |  | Nitrogen | 3 | 8.0478 | 49.698 |  |  |  | 49.338 | 0 |  | 0.73002 |  | 0.0037075 |  | 0.36938 | BGD 28: 40.0; BGD 29: 30.4; BGD 30: 1065.2 | 2021/05/20 18:34:21 | 0
05/20/21 | 25 | 20210048.219 | 004C | 0.818 |  |  | 875 | -37.046 | Carbon | 4 | 54.8196 | 24.061 |  |  | 23.691 |  | 0 | 1.1432 |  | 0.010766 |  | 1.0651 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.8 | 2021/05/20 18:34:21 | 95
05/20/21 | 25 | 20210048.219 | 004C | 0.818 |  |  | 5225 | -43.8 |  | 5 | 55.3839 | 97.857 |  |  | 96.372 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/20 18:34:21 | 95
05/20/21 | 25 | 20210048.219 | 004C | 0.818 |  |  | 5227 | -44 |  | 6 | 55.108 | 98.018 |  |  | 96.531 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.5 | 2021/05/20 18:34:21 | 95
05/20/21 | 26 | 20210048.220 | 006A | 0.843 | 6191 | 0.101 |  |  |  | 1 | 14.6129 | 114 |  |  |  | 113.18 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.6; BGD 30: 978.4 | 2021/05/20 18:44:06 | 0
05/20/21 | 26 | 20210048.220 | 006A | 0.843 | 6189 | 0 |  |  |  | 2 | 14.6093 | 113.952 |  |  |  | 113.133 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.8; BGD 30: 1099.0 | 2021/05/20 18:44:06 | 0
05/20/21 | 26 | 20210048.220 | 006A | 0.843 | 2515 | 5.232 |  |  | Nitrogen | 3 | 9.2946 | 61.16 |  |  |  | 60.718 | 0 |  | 0.72806 |  | 0.0036974 |  | 0.36838 | BGD 28: 39.9; BGD 29: 30.4; BGD 30: 1067.1 | 2021/05/20 18:44:06 | 0
05/20/21 | 26 | 20210048.220 | 006A | 0.843 |  |  | 770 | -34.657 | Carbon | 4 | 47.9081 | 21.019 |  |  | 20.695 |  | 0 | 1.1458 |  | 0.010793 |  | 1.0677 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.0 | 2021/05/20 18:44:06 | 95
05/20/21 | 26 | 20210048.220 | 006A | 0.843 |  |  | 5233 | -43.8 |  | 5 | 53.7654 | 97.842 |  |  | 96.357 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/20 18:44:06 | 95
05/20/21 | 26 | 20210048.220 | 006A | 0.843 |  |  | 5229 | -44.025 |  | 6 | 53.4215 | 98.049 |  |  | 96.562 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.5 | 2021/05/20 18:44:06 | 95
05/20/21 | 27 | 20210048.221 | 006B | 0.82 | 6197 | 0.083 |  |  |  | 1 | 15.0284 | 114.071 |  |  |  | 113.25 | 0 |  | 0.72434 |  | 0.0036785 |  | 0.3665 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 980.0 | 2021/05/20 18:53:52 | 0
05/20/21 | 27 | 20210048.221 | 006B | 0.82 | 6200 | 0 |  |  |  | 2 | 15.0335 | 114.136 |  |  |  | 113.315 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1100.7 | 2021/05/20 18:53:52 | 0
05/20/21 | 27 | 20210048.221 | 006B | 0.82 | 2457 | 5.346 |  |  | Nitrogen | 3 | 9.3772 | 59.775 |  |  |  | 59.343 | 0 |  | 0.72815 |  | 0.0036979 |  | 0.36842 | BGD 28: 39.8; BGD 29: 30.2; BGD 30: 1068.9 | 2021/05/20 18:53:52 | 0
05/20/21 | 27 | 20210048.221 | 006B | 0.82 |  |  | 745 | -34.753 | Carbon | 4 | 47.948 | 20.322 |  |  | 20.009 |  | 0 | 1.1457 |  | 0.010792 |  | 1.0676 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 8.9 | 2021/05/20 18:53:52 | 95
05/20/21 | 27 | 20210048.221 | 006B | 0.82 |  |  | 5237 | -43.8 |  | 5 | 55.1497 | 97.915 |  |  | 96.429 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.1 | 2021/05/20 18:53:52 | 95
05/20/21 | 27 | 20210048.221 | 006B | 0.82 |  |  | 5233 | -44.016 |  | 6 | 54.7022 | 98.177 |  |  | 96.687 |  | 0 | 1.1342 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.5 | 2021/05/20 18:53:52 | 95
05/20/21 | 28 | 20210048.222 | 006C | 0.841 | 6190 | 0.095 |  |  |  | 1 | 14.6457 | 113.974 |  |  |  | 113.155 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 980.2 | 2021/05/20 19:03:37 | 0
05/20/21 | 28 | 20210048.222 | 006C | 0.841 | 6195 | 0 |  |  |  | 2 | 14.6517 | 114.052 |  |  |  | 113.232 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.7; BGD 30: 1101.1 | 2021/05/20 19:03:37 | 0
05/20/21 | 28 | 20210048.222 | 006C | 0.841 | 2457 | 5.173 |  |  | Nitrogen | 3 | 9.1563 | 59.88 |  |  |  | 59.447 | 0 |  | 0.72803 |  | 0.0036972 |  | 0.36836 | BGD 28: 39.7; BGD 29: 30.2; BGD 30: 1069.0 | 2021/05/20 19:03:37 | 0
05/20/21 | 28 | 20210048.222 | 006C | 0.841 |  |  | 753 | -34.91 | Carbon | 4 | 47.1502 | 20.54 |  |  | 20.224 |  | 0 | 1.1456 |  | 0.01079 |  | 1.0675 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/20 19:03:37 | 95
05/20/21 | 28 | 20210048.222 | 006C | 0.841 |  |  | 5236 | -43.8 |  | 5 | 53.8057 | 97.895 |  |  | 96.409 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.2 | 2021/05/20 19:03:37 | 95
05/20/21 | 28 | 20210048.222 | 006C | 0.841 |  |  | 5236 | -44.022 |  | 6 | 53.476 | 98.093 |  |  | 96.605 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.5 | 2021/05/20 19:03:37 | 95
05/20/21 | 29 | 20210048.223 | 008A | 0.834 | 6190 | 0.11 |  |  |  | 1 | 14.7613 | 113.88 |  |  |  | 113.061 | 0 |  | 0.72436 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.9; BGD 29: 30.4; BGD 30: 980.8 | 2021/05/20 19:13:22 | 0
05/20/21 | 29 | 20210048.223 | 008A | 0.834 | 6188 | 0 |  |  |  | 2 | 14.7696 | 113.987 |  |  |  | 113.167 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.7; BGD 30: 1101.4 | 2021/05/20 19:13:22 | 0
05/20/21 | 29 | 20210048.223 | 008A | 0.834 | 2864 | 8.267 |  |  | Nitrogen | 3 | 10.3812 | 69.238 |  |  |  | 68.736 | 0 |  | 0.73027 |  | 0.0037086 |  | 0.36949 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1071.0 | 2021/05/20 19:13:22 | 0
05/20/21 | 29 | 20210048.223 | 008A | 0.834 |  |  | 826 | -34.75 | Carbon | 4 | 51.155 | 22.543 |  |  | 22.196 |  | 0 | 1.1458 |  | 0.010792 |  | 1.0676 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.0 | 2021/05/20 19:13:22 | 95
05/20/21 | 29 | 20210048.223 | 008A | 0.834 |  |  | 5233 | -43.8 |  | 5 | 54.2561 | 97.895 |  |  | 96.41 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/20 19:13:22 | 95
05/20/21 | 29 | 20210048.223 | 008A | 0.834 |  |  | 5239 | -44.001 |  | 6 | 53.8461 | 98.14 |  |  | 96.651 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.6 | 2021/05/20 19:13:22 | 95
05/20/21 | 30 | 20210048.224 | 008B | 0.85 | 6188 | 0.11 |  |  |  | 1 | 14.485 | 113.901 |  |  |  | 113.082 | 0 |  | 0.72436 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 983.7 | 2021/05/20 19:23:07 | 0
05/20/21 | 30 | 20210048.224 | 008B | 0.85 | 6191 | 0 |  |  |  | 2 | 14.4878 | 113.937 |  |  |  | 113.118 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1104.4 | 2021/05/20 19:23:07 | 0
05/20/21 | 30 | 20210048.224 | 008B | 0.85 | 2741 | 7.812 |  |  | Nitrogen | 3 | 9.849 | 66.371 |  |  |  | 65.89 | 0 |  | 0.72994 |  | 0.0037069 |  | 0.36932 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1072.4 | 2021/05/20 19:23:07 | 0
05/20/21 | 30 | 20210048.224 | 008B | 0.85 |  |  | 830 | -34.941 | Carbon | 4 | 50.4234 | 22.678 |  |  | 22.328 |  | 0 | 1.1456 |  | 0.01079 |  | 1.0674 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/20 19:23:07 | 95
05/20/21 | 30 | 20210048.224 | 008B | 0.85 |  |  | 5228 | -43.8 |  | 5 | 53.4297 | 97.777 |  |  | 96.293 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/20 19:23:07 | 95
05/20/21 | 30 | 20210048.224 | 008B | 0.85 |  |  | 5227 | -44.038 |  | 6 | 53.0466 | 98.01 |  |  | 96.523 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.6 | 2021/05/20 19:23:07 | 95
05/20/21 | 31 | 20210048.225 | 008C | 0.81 | 6187 | 0.09 |  |  |  | 1 | 15.1924 | 113.802 |  |  |  | 112.984 | 0 |  | 0.72435 |  | 0.0036785 |  | 0.36651 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 984.8 | 2021/05/20 19:32:52 | 0
05/20/21 | 31 | 20210048.225 | 008C | 0.81 | 6184 | 0 |  |  |  | 2 | 15.195 | 113.835 |  |  |  | 113.016 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.7; BGD 30: 1105.3 | 2021/05/20 19:32:52 | 0
05/20/21 | 31 | 20210048.225 | 008C | 0.81 | 2672 | 7.824 |  |  | Nitrogen | 3 | 10.1309 | 64.74 |  |  |  | 64.271 | 0 |  | 0.72996 |  | 0.003707 |  | 0.36933 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1074.0 | 2021/05/20 19:32:52 | 0
05/20/21 | 31 | 20210048.225 | 008C | 0.81 |  |  | 781 | -34.542 | Carbon | 4 | 50.2958 | 21.251 |  |  | 20.924 |  | 0 | 1.1461 |  | 0.010794 |  | 1.0679 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.0 | 2021/05/20 19:32:52 | 95
05/20/21 | 31 | 20210048.225 | 008C | 0.81 |  |  | 5226 | -43.8 |  | 5 | 56.2345 | 97.68 |  |  | 96.197 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/20 19:32:52 | 95
05/20/21 | 31 | 20210048.225 | 008C | 0.81 |  |  | 5212 | -44.025 |  | 6 | 56.0908 | 97.763 |  |  | 96.28 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.6 | 2021/05/20 19:32:52 | 95
05/20/21 | 32 | 20210048.226 | 010A | 0.849 | 6176 | 0.112 |  |  |  | 1 | 14.4929 | 113.78 |  |  |  | 112.962 | 0 |  | 0.72436 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 986.0 | 2021/05/20 19:42:37 | 0
05/20/21 | 32 | 20210048.226 | 010A | 0.849 | 6174 | 0 |  |  |  | 2 | 14.4833 | 113.656 |  |  |  | 112.839 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.7; BGD 30: 1106.3 | 2021/05/20 19:42:37 | 0
05/20/21 | 32 | 20210048.226 | 010A | 0.849 | 2429 | 5.923 |  |  | Nitrogen | 3 | 8.9754 | 59.123 |  |  |  | 58.696 | 0 |  | 0.72857 |  | 0.0037 |  | 0.36863 | BGD 28: 39.8; BGD 29: 30.2; BGD 30: 1075.1 | 2021/05/20 19:42:37 | 0
05/20/21 | 32 | 20210048.226 | 010A | 0.849 |  |  | 780 | -35.541 | Carbon | 4 | 48.2253 | 21.386 |  |  | 21.057 |  | 0 | 1.1449 |  | 0.010783 |  | 1.0668 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/20 19:42:37 | 95
05/20/21 | 32 | 20210048.226 | 010A | 0.849 |  |  | 5216 | -43.8 |  | 5 | 53.8587 | 97.553 |  |  | 96.072 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/20 19:42:37 | 95
05/20/21 | 32 | 20210048.226 | 010A | 0.849 |  |  | 5210 | -44.023 |  | 6 | 53.6086 | 97.706 |  |  | 96.223 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.6 | 2021/05/20 19:42:37 | 95
05/20/21 | 33 | 20210048.226 | Replicate C8 | 0.843 | 6170 | 0.125 |  |  |  | 1 | 14.5723 | 113.473 |  |  |  | 112.657 | 0 |  | 0.72435 |  | 0.0036787 |  | 0.36652 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 985.8 | 2021/05/20 19:52:22 | 0
05/20/21 | 33 | 20210048.226 | Replicate C8 | 0.843 | 6160 | 0 |  |  |  | 2 | 14.5771 | 113.535 |  |  |  | 112.719 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.7; BGD 30: 1106.2 | 2021/05/20 19:52:22 | 0
05/20/21 | 33 | 20210048.226 | Replicate C8 | 0.843 | 2416 | 5.851 |  |  | Nitrogen | 3 | 9.0132 | 58.917 |  |  |  | 58.49 | 0 |  | 0.7285 |  | 0.0036997 |  | 0.36861 | BGD 28: 39.7; BGD 29: 30.2; BGD 30: 1074.1 | 2021/05/20 19:52:22 | 0
05/20/21 | 33 | 20210048.226 | Replicate C8 | 0.843 |  |  | 782 | -35.578 | Carbon | 4 | 48.6302 | 21.421 |  |  | 21.091 |  | 0 | 1.1449 |  | 0.010782 |  | 1.0667 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/20 19:52:22 | 95
05/20/21 | 33 | 20210048.226 | Replicate C8 | 0.843 |  |  | 5209 | -43.8 |  | 5 | 54.4064 | 97.452 |  |  | 95.973 |  | 1 | 1.1345 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/20 19:52:22 | 95
05/20/21 | 33 | 20210048.226 | Replicate C8 | 0.843 |  |  | 5205 | -43.995 |  | 6 | 54.1405 | 97.614 |  |  | 96.133 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/20 19:52:22 | 95
05/20/21 | 34 | 36-UWSIF-UT Glut 1-20210048.46 | 36-UWSIF-UT Glut 1- | 0.73 | 6152 | 0.081 |  |  |  | 1 | 16.818 | 113.361 |  |  |  | 112.546 | 0 |  | 0.72435 |  | 0.0036785 |  | 0.3665 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 986.9 | 2021/05/20 20:02:07 | 0
05/20/21 | 34 | 36-UWSIF-UT Glut 1-20210048.46 | 36-UWSIF-UT Glut 1- | 0.73 | 6167 | 0 |  |  |  | 2 | 16.829 | 113.484 |  |  |  | 112.668 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.7; BGD 30: 1106.9 | 2021/05/20 20:02:07 | 0
05/20/21 | 34 | 36-UWSIF-UT Glut 1-20210048.46 | 36-UWSIF-UT Glut 1- | 0.73 | 2106 | -2.824 |  |  | Nitrogen | 3 | 9.2838 | 51.416 |  |  |  | 51.047 | 0 |  | 0.72225 |  | 0.0036678 |  | 0.36544 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1073.8 | 2021/05/20 20:02:07 | 0
05/20/21 | 34 | 36-UWSIF-UT Glut 1-20210048.46 | 36-UWSIF-UT Glut 1- | 0.73 |  |  | 598 | -29.87 | Carbon | 4 | 44.297 | 16.04 |  |  | 15.792 |  | 0 | 1.1513 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.0 | 2021/05/20 20:02:07 | 95
05/20/21 | 34 | 36-UWSIF-UT Glut 1-20210048.46 | 36-UWSIF-UT Glut 1- | 0.73 |  |  | 5207 | -43.8 |  | 5 | 62.9896 | 97.367 |  |  | 95.889 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 7.2 | 2021/05/20 20:02:07 | 95
05/20/21 | 34 | 36-UWSIF-UT Glut 1-20210048.46 | 36-UWSIF-UT Glut 1- | 0.73 |  |  | 5202 | -44.033 |  | 6 | 62.6291 | 97.558 |  |  | 96.077 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.6 | 2021/05/20 20:02:07 | 95
05/20/21 | 35 | 36-UWSIF-UT Glut 1-20210048.47 | 36-UWSIF-UT Glut 1- | 0.798 | 6160 | 0.121 |  |  |  | 1 | 15.3925 | 113.454 |  |  |  | 112.638 | 0 |  | 0.72437 |  | 0.0036786 |  | 0.36652 | BGD 28: 39.7; BGD 29: 30.3; BGD 30: 985.4 | 2021/05/20 20:11:52 | 0
05/20/21 | 35 | 36-UWSIF-UT Glut 1-20210048.47 | 36-UWSIF-UT Glut 1- | 0.798 | 6155 | 0 |  |  |  | 2 | 15.3776 | 113.272 |  |  |  | 112.458 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.8; BGD 29: 29.6; BGD 30: 1106.3 | 2021/05/20 20:11:52 | 0
05/20/21 | 35 | 36-UWSIF-UT Glut 1-20210048.47 | 36-UWSIF-UT Glut 1- | 0.798 | 2298 | -2.736 |  |  | Nitrogen | 3 | 9.1106 | 55.874 |  |  |  | 55.474 | 0 |  | 0.7223 |  | 0.0036681 |  | 0.36547 | BGD 28: 39.7; BGD 29: 30.1; BGD 30: 1072.9 | 2021/05/20 20:11:52 | 0
05/20/21 | 35 | 36-UWSIF-UT Glut 1-20210048.47 | 36-UWSIF-UT Glut 1- | 0.798 |  |  | 651 | -29.903 | Carbon | 4 | 43.7627 | 17.578 |  |  | 17.306 |  | 0 | 1.1512 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/20 20:11:52 | 95
05/20/21 | 35 | 36-UWSIF-UT Glut 1-20210048.47 | 36-UWSIF-UT Glut 1- | 0.798 |  |  | 5211 | -43.8 |  | 5 | 57.5966 | 97.382 |  |  | 95.904 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/20 20:11:52 | 95
05/20/21 | 35 | 36-UWSIF-UT Glut 1-20210048.47 | 36-UWSIF-UT Glut 1- | 0.798 |  |  | 5201 | -44.037 |  | 6 | 57.3577 | 97.52 |  |  | 96.04 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.7 | 2021/05/20 20:11:52 | 95
05/20/21 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 39-UWSIF-UW Glut 2- | 0.835 | 6147 | 0.119 |  |  |  | 1 | 14.7047 | 113.381 |  |  |  | 112.565 | 0 |  | 0.72437 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 986.9 | 2021/05/20 20:21:38 | 0
05/20/21 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 39-UWSIF-UW Glut 2- | 0.835 | 6155 | 0 |  |  |  | 2 | 14.7023 | 113.35 |  |  |  | 112.534 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.8; BGD 29: 29.6; BGD 30: 1107.4 | 2021/05/20 20:21:38 | 0
05/20/21 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 39-UWSIF-UW Glut 2- | 0.835 | 2618 | 29.682 |  |  | Nitrogen | 3 | 9.675 | 63.499 |  |  |  | 63.029 | 0 |  | 0.74578 |  | 0.0037874 |  | 0.37731 | BGD 28: 39.7; BGD 29: 30.2; BGD 30: 1077.9 | 2021/05/20 20:21:38 | 0
05/20/21 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 39-UWSIF-UW Glut 2- | 0.835 |  |  | 737 | 22.128 | Carbon | 4 | 46.5429 | 20.03 |  |  | 19.71 |  | 0 | 1.2086 |  | 0.011428 |  | 1.1298 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.1 | 2021/05/20 20:21:38 | 95
05/20/21 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 39-UWSIF-UW Glut 2- | 0.835 |  |  | 5206 | -43.8 |  | 5 | 55.1166 | 97.338 |  |  | 95.861 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/20 20:21:38 | 95
05/20/21 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 39-UWSIF-UW Glut 2- | 0.835 |  |  | 5203 | -43.999 |  | 6 | 54.8968 | 97.471 |  |  | 95.992 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.7 | 2021/05/20 20:21:38 | 95
05/20/21 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 39-UWSIF-UW Glut 2- | 0.722 | 6147 | 0.105 |  |  |  | 1 | 16.9859 | 113.157 |  |  |  | 112.343 | 0 |  | 0.72437 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.8; BGD 29: 30.5; BGD 30: 990.5 | 2021/05/20 20:31:23 | 0
05/20/21 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 39-UWSIF-UW Glut 2- | 0.722 | 6154 | 0 |  |  |  | 2 | 16.9973 | 113.283 |  |  |  | 112.469 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.8; BGD 29: 29.6; BGD 30: 1111.0 | 2021/05/20 20:31:23 | 0
05/20/21 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 39-UWSIF-UW Glut 2- | 0.722 | 2243 | 29.67 |  |  | Nitrogen | 3 | 9.9123 | 54.835 |  |  |  | 54.429 | 0 |  | 0.74578 |  | 0.0037873 |  | 0.3773 | BGD 28: 39.7; BGD 29: 30.2; BGD 30: 1078.6 | 2021/05/20 20:31:23 | 0
05/20/21 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 39-UWSIF-UW Glut 2- | 0.722 |  |  | 636 | 22.192 | Carbon | 4 | 47.5412 | 17.217 |  |  | 16.942 |  | 0 | 1.2087 |  | 0.011428 |  | 1.1299 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.2 | 2021/05/20 20:31:23 | 95
05/20/21 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 39-UWSIF-UW Glut 2- | 0.722 |  |  | 5201 | -43.8 |  | 5 | 63.6859 | 97.368 |  |  | 95.89 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.7; BGD 45: 0.9; BGD 46: 7.3 | 2021/05/20 20:31:23 | 95
05/20/21 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 39-UWSIF-UW Glut 2- | 0.722 |  |  | 5197 | -44.004 |  | 6 | 63.4033 | 97.516 |  |  | 96.036 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.6; BGD 46: 8.7 | 2021/05/20 20:31:23 | 95
05/20/21 | 38 | 01-UWSIF-Liver-20210048.43 | 01-UWSIF-Liver- | 0.783 | 6152 | 0.119 |  |  |  | 1 | 15.6716 | 113.265 |  |  |  | 112.45 | 0 |  | 0.72438 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.8; BGD 29: 30.4; BGD 30: 991.1 | 2021/05/20 20:41:08 | 0
05/20/21 | 38 | 01-UWSIF-Liver-20210048.43 | 01-UWSIF-Liver- | 0.783 | 6148 | 0 |  |  |  | 2 | 15.6753 | 113.309 |  |  |  | 112.495 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.8; BGD 29: 29.6; BGD 30: 1112.3 | 2021/05/20 20:41:08 | 0
05/20/21 | 38 | 01-UWSIF-Liver-20210048.43 | 01-UWSIF-Liver- | 0.783 | 3130 | 8.554 |  |  | Nitrogen | 3 | 11.8462 | 75.649 |  |  |  | 75.1 | 0 |  | 0.73049 |  | 0.0037097 |  | 0.3696 | BGD 28: 39.7; BGD 29: 30.2; BGD 30: 1081.4 | 2021/05/20 20:41:08 | 0
05/20/21 | 38 | 01-UWSIF-Liver-20210048.43 | 01-UWSIF-Liver- | 0.783 |  |  | 790 | -19.547 | Carbon | 4 | 52.4915 | 21.491 |  |  | 21.156 |  | 0 | 1.1626 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.2 | 2021/05/20 20:41:08 | 95
05/20/21 | 38 | 01-UWSIF-Liver-20210048.43 | 01-UWSIF-Liver- | 0.783 |  |  | 5206 | -43.8 |  | 5 | 58.7889 | 97.331 |  |  | 95.854 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/20 20:41:08 | 95
05/20/21 | 38 | 01-UWSIF-Liver-20210048.43 | 01-UWSIF-Liver- | 0.783 |  |  | 5197 | -44.005 |  | 6 | 58.5448 | 97.47 |  |  | 95.991 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.8 | 2021/05/20 20:41:08 | 95
05/20/21 | 39 | 01-UWSIF-Liver-20210048.44 | 01-UWSIF-Liver- | 0.727 | 6155 | 0.111 |  |  |  | 1 | 16.8717 | 113.186 |  |  |  | 112.372 | 0 |  | 0.72437 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 997.3 | 2021/05/20 20:50:53 | 0
05/20/21 | 39 | 01-UWSIF-Liver-20210048.44 | 01-UWSIF-Liver- | 0.727 | 6158 | 0 |  |  |  | 2 | 16.8886 | 113.374 |  |  |  | 112.559 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.8; BGD 30: 1117.9 | 2021/05/20 20:50:53 | 0
05/20/21 | 39 | 01-UWSIF-Liver-20210048.44 | 01-UWSIF-Liver- | 0.727 | 2885 | 8.59 |  |  | Nitrogen | 3 | 12.005 | 69.945 |  |  |  | 69.438 | 0 |  | 0.73052 |  | 0.0037098 |  | 0.36961 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1086.3 | 2021/05/20 20:50:53 | 0
05/20/21 | 39 | 01-UWSIF-Liver-20210048.44 | 01-UWSIF-Liver- | 0.727 |  |  | 731 | -19.51 | Carbon | 4 | 53.0244 | 19.829 |  |  | 19.52 |  | 0 | 1.1627 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/20 20:50:53 | 95
05/20/21 | 39 | 01-UWSIF-Liver-20210048.44 | 01-UWSIF-Liver- | 0.727 |  |  | 5205 | -43.8 |  | 5 | 63.344 | 97.317 |  |  | 95.84 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/20 20:50:53 | 95
05/20/21 | 39 | 01-UWSIF-Liver-20210048.44 | 01-UWSIF-Liver- | 0.727 |  |  | 5196 | -44.027 |  | 6 | 63.1104 | 97.44 |  |  | 95.962 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 8.8 | 2021/05/20 20:50:53 | 95
05/20/21 | 40 | 20210048.227 | 010B | 0.84 | 6155 | 0.116 |  |  |  | 1 | 14.6127 | 113.323 |  |  |  | 112.508 | 0 |  | 0.72437 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 999.6 | 2021/05/20 21:00:38 | 0
05/20/21 | 40 | 20210048.227 | 010B | 0.84 | 6146 | 0 |  |  |  | 2 | 14.6104 | 113.293 |  |  |  | 112.479 | 1 |  | 0.72429 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.7; BGD 30: 1121.0 | 2021/05/20 21:00:38 | 0
05/20/21 | 40 | 20210048.227 | 010B | 0.84 | 2527 | 5.928 |  |  | Nitrogen | 3 | 9.3615 | 61.43 |  |  |  | 60.986 | 0 |  | 0.72858 |  | 0.0037 |  | 0.36864 | BGD 28: 39.8; BGD 29: 30.2; BGD 30: 1088.0 | 2021/05/20 21:00:38 | 0
05/20/21 | 40 | 20210048.227 | 010B | 0.84 |  |  | 789 | -35.398 | Carbon | 4 | 49.095 | 21.583 |  |  | 21.251 |  | 0 | 1.1451 |  | 0.010784 |  | 1.0669 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.3 | 2021/05/20 21:00:38 | 95
05/20/21 | 40 | 20210048.227 | 010B | 0.84 |  |  | 5200 | -43.8 |  | 5 | 54.8467 | 97.302 |  |  | 95.826 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.6 | 2021/05/20 21:00:38 | 95
05/20/21 | 40 | 20210048.227 | 010B | 0.84 |  |  | 5194 | -44 |  | 6 | 54.5805 | 97.465 |  |  | 95.986 |  | 0 | 1.1343 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.9 | 2021/05/20 21:00:38 | 95
05/20/21 | 41 | 20210048.228 | 010C | 0.814 | 6160 | 0.108 |  |  |  | 1 | 15.0843 | 113.383 |  |  |  | 112.568 | 0 |  | 0.72439 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 1000.7 | 2021/05/20 21:10:23 | 0
05/20/21 | 41 | 20210048.228 | 010C | 0.814 | 6162 | 0 |  |  |  | 2 | 15.0946 | 113.512 |  |  |  | 112.696 | 1 |  | 0.72431 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1122.2 | 2021/05/20 21:10:23 | 0
05/20/21 | 41 | 20210048.228 | 010C | 0.814 | 2384 | 5.892 |  |  | Nitrogen | 3 | 9.2057 | 57.938 |  |  |  | 57.519 | 0 |  | 0.72857 |  | 0.0036999 |  | 0.36862 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1088.4 | 2021/05/20 21:10:23 | 0
05/20/21 | 41 | 20210048.228 | 010C | 0.814 |  |  | 748 | -35.264 | Carbon | 4 | 48.4117 | 20.38 |  |  | 20.067 |  | 0 | 1.1453 |  | 0.010786 |  | 1.0671 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.4 | 2021/05/20 21:10:23 | 95
05/20/21 | 41 | 20210048.228 | 010C | 0.814 |  |  | 5206 | -43.8 |  | 5 | 56.6074 | 97.297 |  |  | 95.82 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.5 | 2021/05/20 21:10:23 | 95
05/20/21 | 41 | 20210048.228 | 010C | 0.814 |  |  | 5193 | -44.025 |  | 6 | 56.2911 | 97.484 |  |  | 96.005 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.9 | 2021/05/20 21:10:23 | 95
05/20/21 | 42 | 20210048.229 | 012A | 0.827 | 6157 | 0.123 |  |  |  | 1 | 14.8511 | 113.434 |  |  |  | 112.618 | 0 |  | 0.7244 |  | 0.0036787 |  | 0.36652 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 1001.2 | 2021/05/20 21:20:08 | 0
05/20/21 | 42 | 20210048.229 | 012A | 0.827 | 6158 | 0 |  |  |  | 2 | 14.854 | 113.47 |  |  |  | 112.654 | 1 |  | 0.72431 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1122.8 | 2021/05/20 21:20:08 | 0
05/20/21 | 42 | 20210048.229 | 012A | 0.827 | 2830 | 8.857 |  |  | Nitrogen | 3 | 10.3706 | 68.416 |  |  |  | 67.919 | 0 |  | 0.73072 |  | 0.0037108 |  | 0.36971 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1091.4 | 2021/05/20 21:20:08 | 0
05/20/21 | 42 | 20210048.229 | 012A | 0.827 |  |  | 800 | -34.735 | Carbon | 4 | 50.2319 | 21.785 |  |  | 21.449 |  | 0 | 1.1459 |  | 0.010792 |  | 1.0677 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.4 | 2021/05/20 21:20:08 | 95
05/20/21 | 42 | 20210048.229 | 012A | 0.827 |  |  | 5204 | -43.8 |  | 5 | 55.6841 | 97.317 |  |  | 95.84 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.6 | 2021/05/20 21:20:08 | 95
05/20/21 | 42 | 20210048.229 | 012A | 0.827 |  |  | 5192 | -44.009 |  | 6 | 55.5034 | 97.426 |  |  | 95.947 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.9 | 2021/05/20 21:20:08 | 95
05/20/21 | 43 | 20210048.230 | 012B | 0.842 | 6160 | 0.106 |  |  |  | 1 | 14.5817 | 113.372 |  |  |  | 112.556 | 0 |  | 0.72439 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 1003.9 | 2021/05/20 21:29:53 | 0
05/20/21 | 43 | 20210048.230 | 012B | 0.842 | 6159 | 0 |  |  |  | 2 | 14.5875 | 113.445 |  |  |  | 112.63 | 1 |  | 0.72431 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1125.6 | 2021/05/20 21:29:53 | 0
05/20/21 | 43 | 20210048.230 | 012B | 0.842 | 2717 | 8.145 |  |  | Nitrogen | 3 | 9.8958 | 65.981 |  |  |  | 65.503 | 0 |  | 0.73021 |  | 0.0037082 |  | 0.36945 | BGD 28: 39.9; BGD 29: 30.4; BGD 30: 1094.2 | 2021/05/20 21:29:53 | 0
05/20/21 | 43 | 20210048.230 | 012B | 0.842 |  |  | 782 | -34.36 | Carbon | 4 | 48.6133 | 21.379 |  |  | 21.049 |  | 0 | 1.1464 |  | 0.010796 |  | 1.0681 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.4 | 2021/05/20 21:29:53 | 95
05/20/21 | 43 | 20210048.230 | 012B | 0.842 |  |  | 5200 | -43.8 |  | 5 | 54.7926 | 97.255 |  |  | 95.779 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.6 | 2021/05/20 21:29:53 | 95
05/20/21 | 43 | 20210048.230 | 012B | 0.842 |  |  | 5203 | -44.022 |  | 6 | 54.1224 | 97.665 |  |  | 96.183 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 8.9 | 2021/05/20 21:29:53 | 95
05/20/21 | 44 | 20210048.231 | 012C | 0.811 | 6158 | 0.127 |  |  |  | 1 | 15.1386 | 113.366 |  |  |  | 112.55 | 0 |  | 0.72441 |  | 0.0036787 |  | 0.36652 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 1006.1 | 2021/05/20 21:39:38 | 0
05/20/21 | 44 | 20210048.231 | 012C | 0.811 | 6158 | 0 |  |  |  | 2 | 15.1474 | 113.474 |  |  |  | 112.658 | 1 |  | 0.72432 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1127.6 | 2021/05/20 21:39:38 | 0
05/20/21 | 44 | 20210048.231 | 012C | 0.811 | 2829 | 8.507 |  |  | Nitrogen | 3 | 10.5701 | 68.374 |  |  |  | 67.878 | 0 |  | 0.73048 |  | 0.0037095 |  | 0.36958 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1095.9 | 2021/05/20 21:39:38 | 0
05/20/21 | 44 | 20210048.231 | 012C | 0.811 |  |  | 787 | -34.482 | Carbon | 4 | 50.5449 | 21.418 |  |  | 21.088 |  | 0 | 1.1462 |  | 0.010795 |  | 1.0679 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.4 | 2021/05/20 21:39:38 | 95
05/20/21 | 44 | 20210048.231 | 012C | 0.811 |  |  | 5205 | -43.8 |  | 5 | 56.8114 | 97.3 |  |  | 95.823 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.6 | 2021/05/20 21:39:38 | 95
05/20/21 | 44 | 20210048.231 | 012C | 0.811 |  |  | 5194 | -43.996 |  | 6 | 56.5934 | 97.429 |  |  | 95.95 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.0 | 2021/05/20 21:39:38 | 95
05/20/21 | 45 | 20210048.232 | 014A | 0.836 | 6155 | 0.132 |  |  |  | 1 | 14.6782 | 113.267 |  |  |  | 112.453 | 0 |  | 0.7244 |  | 0.0036787 |  | 0.36652 | BGD 28: 40.0; BGD 29: 30.6; BGD 30: 1007.2 | 2021/05/20 21:49:24 | 0
05/20/21 | 45 | 20210048.232 | 014A | 0.836 | 6159 | 0 |  |  |  | 2 | 14.6906 | 113.426 |  |  |  | 112.61 | 1 |  | 0.72431 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1129.0 | 2021/05/20 21:49:24 | 0
05/20/21 | 45 | 20210048.232 | 014A | 0.836 | 2905 | 8.546 |  |  | Nitrogen | 3 | 10.4993 | 70.453 |  |  |  | 69.942 | 0 |  | 0.7305 |  | 0.0037096 |  | 0.36959 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1097.6 | 2021/05/20 21:49:24 | 0
05/20/21 | 45 | 20210048.232 | 014A | 0.836 |  |  | 759 | -34.031 | Carbon | 4 | 47.7122 | 20.692 |  |  | 20.373 |  | 0 | 1.1467 |  | 0.0108 |  | 1.0684 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.5 | 2021/05/20 21:49:24 | 95
05/20/21 | 45 | 20210048.232 | 014A | 0.836 |  |  | 5206 | -43.8 |  | 5 | 55.034 | 97.348 |  |  | 95.87 |  | 1 | 1.1346 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.6 | 2021/05/20 21:49:24 | 95
05/20/21 | 45 | 20210048.232 | 014A | 0.836 |  |  | 5197 | -44.005 |  | 6 | 54.5754 | 97.626 |  |  | 96.144 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.0 | 2021/05/20 21:49:24 | 95
05/20/21 | 46 | 20210048.233 | 014B | 0.843 | 6152 | 0.108 |  |  |  | 1 | 14.5609 | 113.326 |  |  |  | 112.511 | 0 |  | 0.7244 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 1009.2 | 2021/05/20 21:59:09 | 0
05/20/21 | 46 | 20210048.233 | 014B | 0.843 | 6160 | 0 |  |  |  | 2 | 14.5652 | 113.382 |  |  |  | 112.566 | 1 |  | 0.72433 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.7; BGD 30: 1131.0 | 2021/05/20 21:59:09 | 0
05/20/21 | 46 | 20210048.233 | 014B | 0.843 | 2628 | 8.387 |  |  | Nitrogen | 3 | 9.5809 | 63.48 |  |  |  | 63.02 | 0 |  | 0.7304 |  | 0.003709 |  | 0.36953 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1099.0 | 2021/05/20 21:59:09 | 0
05/20/21 | 46 | 20210048.233 | 014B | 0.843 |  |  | 842 | -35.024 | Carbon | 4 | 51.3822 | 22.991 |  |  | 22.637 |  | 0 | 1.1456 |  | 0.010789 |  | 1.0673 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.5 | 2021/05/20 21:59:09 | 95
05/20/21 | 46 | 20210048.233 | 014B | 0.843 |  |  | 5211 | -43.8 |  | 5 | 54.5202 | 97.383 |  |  | 95.905 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.7 | 2021/05/20 21:59:09 | 95
05/20/21 | 46 | 20210048.233 | 014B | 0.843 |  |  | 5205 | -44.009 |  | 6 | 54.4098 | 97.45 |  |  | 95.972 |  | 0 | 1.1344 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/20 21:59:09 | 95
05/20/21 | 47 | 20210048.234 | 014C | 0.815 | 6146 | 0.12 |  |  |  | 1 | 15.0464 | 113.142 |  |  |  | 112.328 | 0 |  | 0.72441 |  | 0.0036786 |  | 0.36652 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 1009.5 | 2021/05/20 22:08:54 | 0
05/20/21 | 47 | 20210048.234 | 014C | 0.815 | 6151 | 0 |  |  |  | 2 | 15.067 | 113.399 |  |  |  | 112.584 | 1 |  | 0.72433 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.7; BGD 30: 1130.8 | 2021/05/20 22:08:54 | 0
05/20/21 | 47 | 20210048.234 | 014C | 0.815 | 2754 | 8.203 |  |  | Nitrogen | 3 | 10.3002 | 66.599 |  |  |  | 66.116 | 0 |  | 0.73027 |  | 0.0037084 |  | 0.36947 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1099.6 | 2021/05/20 22:08:54 | 0
05/20/21 | 47 | 20210048.234 | 014C | 0.815 |  |  | 749 | -34.167 | Carbon | 4 | 48.2239 | 20.313 |  |  | 20 |  | 0 | 1.1466 |  | 0.010798 |  | 1.0683 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.5 | 2021/05/20 22:08:54 | 95
05/20/21 | 47 | 20210048.234 | 014C | 0.815 |  |  | 5199 | -43.8 |  | 5 | 56.6294 | 97.243 |  |  | 95.767 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/20 22:08:54 | 95
05/20/21 | 47 | 20210048.234 | 014C | 0.815 |  |  | 5190 | -43.983 |  | 6 | 56.2757 | 97.452 |  |  | 95.973 |  | 0 | 1.1345 |  | 0.010689 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.0 | 2021/05/20 22:08:54 | 95
05/20/21 | 48 | 20210048.235 | 016A | 0.817 | 6141 | 0.086 |  |  |  | 1 | 15.0139 | 113.196 |  |  |  | 112.382 | 0 |  | 0.72432 |  | 0.0036785 |  | 0.3665 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 1010.2 | 2021/05/20 22:18:39 | 0
05/20/21 | 48 | 20210048.235 | 016A | 0.817 | 6153 | 0 |  |  |  | 2 | 15.024 | 113.323 |  |  |  | 112.508 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.7; BGD 30: 1132.0 | 2021/05/20 22:18:39 | 0
05/20/21 | 48 | 20210048.235 | 016A | 0.817 | 2626 | 8.424 |  |  | Nitrogen | 3 | 9.9199 | 63.75 |  |  |  | 63.288 | 0 |  | 0.73036 |  | 0.0037092 |  | 0.36955 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1100.3 | 2021/05/20 22:18:39 | 0
05/20/21 | 48 | 20210048.235 | 016A | 0.817 |  |  | 760 | -34.537 | Carbon | 4 | 48.8264 | 20.695 |  |  | 20.376 |  | 0 | 1.1463 |  | 0.010794 |  | 1.0679 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.5 | 2021/05/20 22:18:39 | 95
05/20/21 | 48 | 20210048.235 | 016A | 0.817 |  |  | 5198 | -43.8 |  | 5 | 56.493 | 97.241 |  |  | 95.765 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/20 22:18:39 | 95
05/20/21 | 48 | 20210048.235 | 016A | 0.817 |  |  | 5194 | -44.009 |  | 6 | 56.0632 | 97.496 |  |  | 96.017 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/20 22:18:39 | 95
05/20/21 | 49 | 20210048.236 | 016B | 0.835 | 6143 | 0.121 |  |  |  | 1 | 14.7042 | 113.374 |  |  |  | 112.559 | 0 |  | 0.72443 |  | 0.0036786 |  | 0.36652 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 1010.7 | 2021/05/20 22:28:24 | 0
05/20/21 | 49 | 20210048.236 | 016B | 0.835 | 6158 | 0 |  |  |  | 2 | 14.7022 | 113.349 |  |  |  | 112.534 | 1 |  | 0.72434 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.7; BGD 30: 1133.1 | 2021/05/20 22:28:24 | 0
05/20/21 | 49 | 20210048.236 | 016B | 0.835 | 2846 | 8.912 |  |  | Nitrogen | 3 | 10.3365 | 68.965 |  |  |  | 68.465 | 0 |  | 0.73079 |  | 0.003711 |  | 0.36973 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1101.4 | 2021/05/20 22:28:24 | 0
05/20/21 | 49 | 20210048.236 | 016B | 0.835 |  |  | 761 | -34.07 | Carbon | 4 | 47.7433 | 20.678 |  |  | 20.359 |  | 0 | 1.1467 |  | 0.010799 |  | 1.0684 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/20 22:28:24 | 95
05/20/21 | 49 | 20210048.236 | 016B | 0.835 |  |  | 5201 | -43.8 |  | 5 | 55.2018 | 97.286 |  |  | 95.809 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/20 22:28:24 | 95
05/20/21 | 49 | 20210048.236 | 016B | 0.835 |  |  | 5198 | -44.005 |  | 6 | 54.8935 | 97.473 |  |  | 95.994 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.0 | 2021/05/20 22:28:24 | 95
05/20/21 | 50 | 20210048.237 | 016C | 0.839 | 6151 | 0.115 |  |  |  | 1 | 14.634 | 113.373 |  |  |  | 112.558 | 0 |  | 0.72434 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 1012.3 | 2021/05/20 22:38:09 | 0
05/20/21 | 50 | 20210048.237 | 016C | 0.839 | 6164 | 0 |  |  |  | 2 | 14.6355 | 113.392 |  |  |  | 112.577 | 1 |  | 0.72426 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1134.8 | 2021/05/20 22:38:09 | 0
05/20/21 | 50 | 20210048.237 | 016C | 0.839 | 2824 | 8.845 |  |  | Nitrogen | 3 | 10.2063 | 68.281 |  |  |  | 67.785 | 0 |  | 0.73066 |  | 0.0037107 |  | 0.3697 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1102.5 | 2021/05/20 22:38:09 | 0
05/20/21 | 50 | 20210048.237 | 016C | 0.839 |  |  | 813 | -34.686 | Carbon | 4 | 50.187 | 22.165 |  |  | 21.823 |  | 0 | 1.1461 |  | 0.010792 |  | 1.0677 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/20 22:38:09 | 95
05/20/21 | 50 | 20210048.237 | 016C | 0.839 |  |  | 5202 | -43.8 |  | 5 | 54.8881 | 97.317 |  |  | 95.84 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/20 22:38:09 | 95
05/20/21 | 50 | 20210048.237 | 016C | 0.839 |  |  | 5198 | -44.023 |  | 6 | 54.6666 | 97.452 |  |  | 95.973 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/20 22:38:09 | 95
05/20/21 | 51 | 20210048.238 | 018A | 0.819 | 6152 | 0.114 |  |  |  | 1 | 14.9966 | 113.439 |  |  |  | 112.624 | 0 |  | 0.72443 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.6; BGD 30: 1013.4 | 2021/05/20 22:47:55 | 0
05/20/21 | 51 | 20210048.238 | 018A | 0.819 | 6159 | 0 |  |  |  | 2 | 14.9961 | 113.433 |  |  |  | 112.617 | 1 |  | 0.72435 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.9; BGD 30: 1136.1 | 2021/05/20 22:47:55 | 0
05/20/21 | 51 | 20210048.238 | 018A | 0.819 | 2830 | 8.899 |  |  | Nitrogen | 3 | 10.4675 | 68.379 |  |  |  | 67.883 | 0 |  | 0.73079 |  | 0.0037109 |  | 0.36972 | BGD 28: 39.9; BGD 29: 30.4; BGD 30: 1104.1 | 2021/05/20 22:47:55 | 0
05/20/21 | 51 | 20210048.238 | 018A | 0.819 |  |  | 795 | -34.758 | Carbon | 4 | 50.3797 | 21.597 |  |  | 21.264 |  | 0 | 1.146 |  | 0.010792 |  | 1.0676 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/20 22:47:55 | 95
05/20/21 | 51 | 20210048.238 | 018A | 0.819 |  |  | 5202 | -43.8 |  | 5 | 56.3217 | 97.261 |  |  | 95.785 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/20 22:47:55 | 95
05/20/21 | 51 | 20210048.238 | 018A | 0.819 |  |  | 5196 | -44.01 |  | 6 | 56.0871 | 97.401 |  |  | 95.923 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.2 | 2021/05/20 22:47:55 | 95
05/20/21 | 52 | 20210048.239 | 018B | 0.841 | 6148 | 0.102 |  |  |  | 1 | 14.5788 | 113.11 |  |  |  | 112.297 | 0 |  | 0.72425 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 1013.4 | 2021/05/20 22:57:40 | 0
05/20/21 | 52 | 20210048.239 | 018B | 0.841 | 6156 | 0 |  |  |  | 2 | 14.5915 | 113.274 |  |  |  | 112.46 | 1 |  | 0.72417 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1135.3 | 2021/05/20 22:57:40 | 0
05/20/21 | 52 | 20210048.239 | 018B | 0.841 | 2613 | 9.105 |  |  | Nitrogen | 3 | 9.5548 | 63.082 |  |  |  | 62.625 | 0 |  | 0.73077 |  | 0.0037117 |  | 0.3698 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1102.3 | 2021/05/20 22:57:40 | 0
05/20/21 | 52 | 20210048.239 | 018B | 0.841 |  |  | 844 | -35.373 | Carbon | 4 | 51.5674 | 23.027 |  |  | 22.673 |  | 0 | 1.1453 |  | 0.010785 |  | 1.067 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 9.6 | 2021/05/20 22:57:40 | 95
05/20/21 | 52 | 20210048.239 | 018B | 0.841 |  |  | 5204 | -43.8 |  | 5 | 54.7738 | 97.307 |  |  | 95.83 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.8 | 2021/05/20 22:57:40 | 95
05/20/21 | 52 | 20210048.239 | 018B | 0.841 |  |  | 5195 | -44.004 |  | 6 | 54.512 | 97.467 |  |  | 95.988 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.2 | 2021/05/20 22:57:40 | 95
05/20/21 | 53 | 20210048.240 | 018C | 0.848 | 6159 | 0.121 |  |  |  | 1 | 14.4747 | 113.321 |  |  |  | 112.506 | 0 |  | 0.72444 |  | 0.0036786 |  | 0.36652 | BGD 28: 40.1; BGD 29: 30.6; BGD 30: 1014.0 | 2021/05/20 23:07:25 | 0
05/20/21 | 53 | 20210048.240 | 018C | 0.848 | 6143 | 0 |  |  |  | 2 | 14.4741 | 113.313 |  |  |  | 112.498 | 1 |  | 0.72436 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.9; BGD 30: 1136.2 | 2021/05/20 23:07:25 | 0
05/20/21 | 53 | 20210048.240 | 018C | 0.848 | 2728 | 9.701 |  |  | Nitrogen | 3 | 9.7989 | 65.756 |  |  |  | 65.279 | 0 |  | 0.73138 |  | 0.0037139 |  | 0.37001 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1103.1 | 2021/05/20 23:07:25 | 0
05/20/21 | 53 | 20210048.240 | 018C | 0.848 |  |  | 851 | -35.227 | Carbon | 4 | 51.4324 | 23.198 |  |  | 22.841 |  | 0 | 1.1455 |  | 0.010786 |  | 1.0671 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/20 23:07:25 | 95
05/20/21 | 53 | 20210048.240 | 018C | 0.848 |  |  | 5203 | -43.8 |  | 5 | 54.2917 | 97.325 |  |  | 95.848 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 7.9 | 2021/05/20 23:07:25 | 95
05/20/21 | 53 | 20210048.240 | 018C | 0.848 |  |  | 5199 | -43.994 |  | 6 | 53.9014 | 97.565 |  |  | 96.085 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.2 | 2021/05/20 23:07:25 | 95
05/20/21 | 54 | 36-UWSIF-UT Glut 1-20210048.411 | 36-UWSIF-UT Glut 1- | 0.851 | 6157 | 0.112 |  |  |  | 1 | 14.417 | 113.234 |  |  |  | 112.42 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.1; BGD 29: 30.6; BGD 30: 1014.3 | 2021/05/20 23:17:10 | 0
05/20/21 | 54 | 36-UWSIF-UT Glut 1-20210048.411 | 36-UWSIF-UT Glut 1- | 0.851 | 6153 | 0 |  |  |  | 2 | 14.4253 | 113.342 |  |  |  | 112.527 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.1; BGD 29: 29.8; BGD 30: 1136.1 | 2021/05/20 23:17:10 | 0
05/20/21 | 54 | 36-UWSIF-UT Glut 1-20210048.411 | 36-UWSIF-UT Glut 1- | 0.851 | 2452 | -2.819 |  |  | Nitrogen | 3 | 8.9988 | 59.479 |  |  |  | 59.053 | 0 |  | 0.72223 |  | 0.0036678 |  | 0.36544 | BGD 28: 39.9; BGD 29: 30.4; BGD 30: 1104.3 | 2021/05/20 23:17:10 | 0
05/20/21 | 54 | 36-UWSIF-UT Glut 1-20210048.411 | 36-UWSIF-UT Glut 1- | 0.851 |  |  | 696 | -29.886 | Carbon | 4 | 43.4201 | 18.822 |  |  | 18.531 |  | 0 | 1.1514 |  | 0.010846 |  | 1.073 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/20 23:17:10 | 95
05/20/21 | 54 | 36-UWSIF-UT Glut 1-20210048.411 | 36-UWSIF-UT Glut 1- | 0.851 |  |  | 5205 | -43.8 |  | 5 | 54.0743 | 97.341 |  |  | 95.864 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/20 23:17:10 | 95
05/20/21 | 54 | 36-UWSIF-UT Glut 1-20210048.411 | 36-UWSIF-UT Glut 1- | 0.851 |  |  | 5198 | -44.009 |  | 6 | 53.9841 | 97.397 |  |  | 95.919 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/20 23:17:10 | 95
05/20/21 | 55 | 36-UWSIF-UT Glut 1-20210048.412 | 36-UWSIF-UT Glut 1- | 0.83 | 6152 | 0.111 |  |  |  | 1 | 14.7928 | 113.374 |  |  |  | 112.559 | 0 |  | 0.72444 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.9; BGD 29: 30.5; BGD 30: 1012.6 | 2021/05/20 23:26:56 | 0
05/20/21 | 55 | 36-UWSIF-UT Glut 1-20210048.412 | 36-UWSIF-UT Glut 1- | 0.83 | 6162 | 0 |  |  |  | 2 | 14.7924 | 113.37 |  |  |  | 112.554 | 1 |  | 0.72436 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1134.9 | 2021/05/20 23:26:56 | 0
05/20/21 | 55 | 36-UWSIF-UT Glut 1-20210048.412 | 36-UWSIF-UT Glut 1- | 0.83 | 2390 | -2.798 |  |  | Nitrogen | 3 | 9.0661 | 58.231 |  |  |  | 57.813 | 0 |  | 0.72233 |  | 0.0036679 |  | 0.36545 | BGD 28: 39.8; BGD 29: 30.3; BGD 30: 1103.1 | 2021/05/20 23:26:56 | 0
05/20/21 | 55 | 36-UWSIF-UT Glut 1-20210048.412 | 36-UWSIF-UT Glut 1- | 0.83 |  |  | 678 | -29.966 | Carbon | 4 | 43.6059 | 18.353 |  |  | 18.069 |  | 0 | 1.1513 |  | 0.010845 |  | 1.0729 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/20 23:26:56 | 95
05/20/21 | 55 | 36-UWSIF-UT Glut 1-20210048.412 | 36-UWSIF-UT Glut 1- | 0.83 |  |  | 5200 | -43.8 |  | 5 | 55.5578 | 97.272 |  |  | 95.795 |  | 1 | 1.1347 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/20 23:26:56 | 95
05/20/21 | 55 | 36-UWSIF-UT Glut 1-20210048.412 | 36-UWSIF-UT Glut 1- | 0.83 |  |  | 5194 | -44.009 |  | 6 | 55.3918 | 97.372 |  |  | 95.894 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 9.0 | 2021/05/20 23:26:56 | 95
05/20/21 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 39-UWSIF-UW Glut 2- | 0.779 | 6154 | 0.113 |  |  |  | 1 | 15.7442 | 113.172 |  |  |  | 112.358 | 0 |  | 0.72444 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.8; BGD 29: 30.4; BGD 30: 1011.0 | 2021/05/20 23:36:41 | 0
05/20/21 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 39-UWSIF-UW Glut 2- | 0.779 | 6154 | 0 |  |  |  | 2 | 15.7564 | 113.316 |  |  |  | 112.501 | 1 |  | 0.72436 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.7; BGD 30: 1133.0 | 2021/05/20 23:36:41 | 0
05/20/21 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 39-UWSIF-UW Glut 2- | 0.779 | 2415 | 29.663 |  |  | Nitrogen | 3 | 9.7665 | 59.01 |  |  |  | 58.574 | 0 |  | 0.74585 |  | 0.0037873 |  | 0.3773 | BGD 28: 39.8; BGD 29: 30.2; BGD 30: 1100.7 | 2021/05/20 23:36:41 | 0
05/20/21 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 39-UWSIF-UW Glut 2- | 0.779 |  |  | 685 | 22.381 | Carbon | 4 | 47.0646 | 18.644 |  |  | 18.345 |  | 0 | 1.2091 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.9; BGD 45: 1.0; BGD 46: 9.5 | 2021/05/20 23:36:41 | 95
05/20/21 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 39-UWSIF-UW Glut 2- | 0.779 |  |  | 5197 | -43.8 |  | 5 | 59.2867 | 97.22 |  |  | 95.744 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/20 23:36:41 | 95
05/20/21 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 39-UWSIF-UW Glut 2- | 0.779 |  |  | 5193 | -44.008 |  | 6 | 58.9119 | 97.432 |  |  | 95.954 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/20 23:36:41 | 95
05/20/21 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 39-UWSIF-UW Glut 2- | 0.855 | 6151 | 0.11 |  |  |  | 1 | 14.3426 | 113.143 |  |  |  | 112.329 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 39.8; BGD 29: 30.4; BGD 30: 1010.4 | 2021/05/20 23:46:26 | 0
05/20/21 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 39-UWSIF-UW Glut 2- | 0.855 | 6157 | 0 |  |  |  | 2 | 14.3493 | 113.231 |  |  |  | 112.416 | 1 |  | 0.72427 |  | 0.0036782 |  | 0.36647 | BGD 28: 38.9; BGD 29: 29.7; BGD 30: 1132.6 | 2021/05/20 23:46:26 | 0
05/20/21 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 39-UWSIF-UW Glut 2- | 0.855 | 2675 | 29.64 |  |  | Nitrogen | 3 | 9.6168 | 64.9 |  |  |  | 64.42 | 0 |  | 0.74574 |  | 0.0037872 |  | 0.37729 | BGD 28: 39.7; BGD 29: 30.2; BGD 30: 1102.1 | 2021/05/20 23:46:26 | 0
05/20/21 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 39-UWSIF-UW Glut 2- | 0.855 |  |  | 758 | 22.39 | Carbon | 4 | 46.4967 | 20.606 |  |  | 20.276 |  | 0 | 1.2091 |  | 0.01143 |  | 1.1301 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/20 23:46:26 | 95
05/20/21 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 39-UWSIF-UW Glut 2- | 0.855 |  |  | 5198 | -43.8 |  | 5 | 53.7385 | 97.393 |  |  | 95.915 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.7 | 2021/05/20 23:46:26 | 95
05/20/21 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 39-UWSIF-UW Glut 2- | 0.855 |  |  | 5199 | -44.024 |  | 6 | 53.6833 | 97.427 |  |  | 95.949 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/20 23:46:26 | 95
05/20/21 | 58 | 01-UWSIF-Liver-20210048.47 | 01-UWSIF-Liver- | 0.724 | 6154 | 0.125 |  |  |  | 1 | 16.9568 | 113.355 |  |  |  | 112.54 | 0 |  | 0.72445 |  | 0.0036787 |  | 0.36652 | BGD 28: 40.0; BGD 29: 30.6; BGD 30: 1012.0 | 2021/05/20 23:56:11 | 0
05/20/21 | 58 | 01-UWSIF-Liver-20210048.47 | 01-UWSIF-Liver- | 0.724 | 6151 | 0 |  |  |  | 2 | 16.9525 | 113.308 |  |  |  | 112.493 | 1 |  | 0.72436 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1134.2 | 2021/05/20 23:56:11 | 0
05/20/21 | 58 | 01-UWSIF-Liver-20210048.47 | 01-UWSIF-Liver- | 0.724 | 2878 | 8.55 |  |  | Nitrogen | 3 | 12.0165 | 69.663 |  |  |  | 69.158 | 0 |  | 0.73055 |  | 0.0037096 |  | 0.36959 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1102.7 | 2021/05/20 23:56:11 | 0
05/20/21 | 58 | 01-UWSIF-Liver-20210048.47 | 01-UWSIF-Liver- | 0.724 |  |  | 730 | -19.537 | Carbon | 4 | 53.14 | 19.781 |  |  | 19.473 |  | 0 | 1.1629 |  | 0.010962 |  | 1.0843 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.6 | 2021/05/20 23:56:11 | 95
05/20/21 | 58 | 01-UWSIF-Liver-20210048.47 | 01-UWSIF-Liver- | 0.724 |  |  | 5207 | -43.8 |  | 5 | 63.6281 | 97.305 |  |  | 95.829 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/20 23:56:11 | 95
05/20/21 | 58 | 01-UWSIF-Liver-20210048.47 | 01-UWSIF-Liver- | 0.724 |  |  | 5191 | -44.008 |  | 6 | 63.4434 | 97.403 |  |  | 95.925 |  | 0 | 1.1345 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.5; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/20 23:56:11 | 95
05/21/21 | 59 | 01-UWSIF-Liver-20210048.48 | 01-UWSIF-Liver- | 0.723 | 6150 | 0.1 |  |  |  | 1 | 16.9646 | 113.181 |  |  |  | 112.367 | 0 |  | 0.72435 |  | 0.0036786 |  | 0.36651 | BGD 28: 40.0; BGD 29: 30.5; BGD 30: 1013.3 | 2021/05/21 00:05:58 | 0
05/21/21 | 59 | 01-UWSIF-Liver-20210048.48 | 01-UWSIF-Liver- | 0.723 | 6152 | 0 |  |  |  | 2 | 16.9771 | 113.319 |  |  |  | 112.505 | 1 |  | 0.72428 |  | 0.0036782 |  | 0.36647 | BGD 28: 39.0; BGD 29: 29.8; BGD 30: 1135.5 | 2021/05/21 00:05:58 | 0
05/21/21 | 59 | 01-UWSIF-Liver-20210048.48 | 01-UWSIF-Liver- | 0.723 | 2858 | 8.573 |  |  | Nitrogen | 3 | 11.9686 | 69.191 |  |  |  | 68.689 | 0 |  | 0.73048 |  | 0.0037097 |  | 0.3696 | BGD 28: 39.9; BGD 29: 30.3; BGD 30: 1103.6 | 2021/05/21 00:05:58 | 0
05/21/21 | 59 | 01-UWSIF-Liver-20210048.48 | 01-UWSIF-Liver- | 0.723 |  |  | 726 | -19.47 | Carbon | 4 | 52.9501 | 19.66 |  |  | 19.354 |  | 0 | 1.1629 |  | 0.010963 |  | 1.0844 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 9.7 | 2021/05/21 00:05:58 | 95
05/21/21 | 59 | 01-UWSIF-Liver-20210048.48 | 01-UWSIF-Liver- | 0.723 |  |  | 5194 | -43.8 |  | 5 | 63.8359 | 97.242 |  |  | 95.766 |  | 1 | 1.1348 |  | 0.010691 |  | 1.0577 |  | BGD 44: 0.8; BGD 45: 0.9; BGD 46: 7.8 | 2021/05/21 00:05:58 | 95
05/21/21 | 59 | 01-UWSIF-Liver-20210048.48 | 01-UWSIF-Liver- | 0.723 |  |  | 5194 | -43.993 |  | 6 | 63.5107 | 97.413 |  |  | 95.935 |  | 0 | 1.1346 |  | 0.010688 |  | 1.0575 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 9.1 | 2021/05/21 00:05:58 | 95

## Sorted 1
Date | Row | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | %N | %C
05/17/21 | 1 | 36-UWSIF-UT Glut 1-20210048.11 | 36-UWSIF-UT Glut 1- | 0.735 | 2403 | -2.762 | 663 | -29.828 | 9.662 | 41.5389
05/17/21 | 2 | 36-UWSIF-UT Glut 1-20210048.12 | 36-UWSIF-UT Glut 1- | 0.813 | 2512 | -2.76 | 695 | -29.841 | 9.2644 | 39.8903
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 | 1327 | -2.852 | 370 | -29.59 | 9.6168 | 41.3622
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 | 3291 | -2.836 | 899 | -29.901 | 9.4625 | 40.513
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 | 10284 | -2.9 | 2395 | -29.976 | 9.5207 | 40.833
05/17/21 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 39-UWSIF-UW Glut 2- | 0.728 | 2501 | 29.606 | 690 | 22.358 | 10.3379 | 44.1575
05/17/21 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 39-UWSIF-UW Glut 2- | 0.747 | 2580 | 29.738 | 713 | 22.251 | 10.3375 | 44.2951
05/17/21 | 8 | 01-UWSIF-Liver-20210048.11 | 01-UWSIF-Liver- | 0.713 | 3148 | 8.575 | 776 | -19.369 | 13.137 | 50.4859
05/17/21 | 9 | 01-UWSIF-Liver-20210048.12 | 01-UWSIF-Liver- | 0.715 | 3137 | 8.605 | 775 | -19.477 | 13.0696 | 50.263
05/17/21 | 10 | 20210048.001 | 501 | 0.823 | 3916 | 13.59 | 850 | -32.506 | 14.0808 | 47.9424
05/17/21 | 11 | 20210048.002 | 503 | 0.805 | 3833 | 12.77 | 858 | -33.315 | 14.082 | 49.4339
05/17/21 | 12 | 20210048.002 | Replicate A11 | 0.834 | 4021 | 12.805 | 893 | -33.258 | 14.2457 | 49.695
05/17/21 | 13 | 20210048.003 | 505 | 0.841 | 3989 | 10.204 | 892 | -32.879 | 13.9989 | 49.2138
05/17/21 | 14 | 20210048.004 | 507 | 0.846 | 3816 | 12.526 | 914 | -33.262 | 13.3339 | 50.086
05/17/21 | 15 | 20210048.005 | 509 | 0.83 | 4001 | 12.168 | 873 | -33.502 | 14.2681 | 48.839
05/17/21 | 16 | 20210048.006 | 511 | 0.847 | 4174 | 8.944 | 879 | -32.542 | 14.5578 | 48.1578
05/17/21 | 17 | 20210048.007 | 513 | 0.832 | 3350 | 10.151 | 941 | -33.499 | 11.8983 | 52.4939
05/17/21 | 18 | 20210048.008 | 515 | 0.839 | 4075 | 12.337 | 862 | -32.778 | 14.3434 | 47.5313
05/17/21 | 19 | 20210048.009 | 517 | 0.84 | 3905 | 12.444 | 903 | -33.224 | 13.7 | 49.7421
05/17/21 | 20 | 20210048.010 | 519 | 0.826 | 3839 | 8.949 | 876 | -33.595 | 13.7089 | 49.1462
05/17/21 | 21 | 20210048.011 | 521 | 0.832 | 3794 | 12.899 | 896 | -33.506 | 13.4193 | 49.7891
05/17/21 | 22 | 20210048.012 | 523 | 0.841 | 3968 | 12.704 | 896 | -33.052 | 13.87 | 49.208
05/17/21 | 23 | 20210048.013 | 525 | 0.817 | 4102 | 12.084 | 834 | -31.685 | 14.7831 | 47.2054
05/17/21 | 24 | 20210048.014 | 527 | 0.842 | 3703 | 13.28 | 905 | -33.68 | 12.9262 | 49.6798
05/17/21 | 25 | 20210048.015 | 529 | 0.806 | 3956 | 9.147 | 853 | -31.415 | 14.3938 | 48.7743
05/17/21 | 26 | 20210048.016 | 531 | 0.818 | 4029 | 13.058 | 870 | -32.751 | 14.3933 | 49.0016
05/17/21 | 27 | 20210048.017 | 533 | 0.85 | 3912 | 12.487 | 945 | -30.835 | 13.431 | 51.2161
05/17/21 | 28 | 20210048.018 | 535 | 0.811 | 4090 | 9.468 | 859 | -32.941 | 14.7386 | 48.7378
05/17/21 | 29 | 20210048.019 | 537 | 0.828 | 4124 | 10.733 | 871 | -32.353 | 14.552 | 48.3745
05/17/21 | 30 | 20210048.020 | 541 | 0.844 | 3940 | 11.865 | 908 | -29.54 | 13.6453 | 49.5063
05/17/21 | 31 | 20210048.021 | 543 | 0.826 | 4097 | 9.986 | 866 | -31.269 | 14.5024 | 48.2931
05/17/21 | 32 | 20210048.022 | 545 | 0.831 | 3794 | 8.84 | 905 | -33.152 | 13.3568 | 50.0971
05/17/21 | 33 | 20210048.022 | Replicate C8 | 0.829 | 3797 | 8.84 | 915 | -33.132 | 13.3526 | 50.7997
05/17/21 | 34 | 36-UWSIF-UT Glut 1-20210048.16 | 36-UWSIF-UT Glut 1- | 0.842 | 2708 | -2.848 | 750 | -29.822 | 9.4947 | 40.9385
05/17/21 | 35 | 36-UWSIF-UT Glut 1-20210048.17 | 36-UWSIF-UT Glut 1- | 0.766 | 2437 | -2.777 | 677 | -29.878 | 9.451 | 40.6972
05/17/21 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 39-UWSIF-UW Glut 2- | 0.788 | 2722 | 29.711 | 754 | 22.276 | 10.2473 | 44.1684
05/17/21 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 39-UWSIF-UW Glut 2- | 0.782 | 2713 | 29.779 | 752 | 22.428 | 10.2817 | 44.2436
05/17/21 | 38 | 01-UWSIF-Liver-20210048.13 | 01-UWSIF-Liver- | 0.776 | 3440 | 8.569 | 852 | -19.454 | 12.9628 | 50.4109
05/17/21 | 39 | 01-UWSIF-Liver-20210048.14 | 01-UWSIF-Liver- | 0.752 | 3351 | 8.527 | 832 | -19.433 | 13.0363 | 50.7242
05/17/21 | 40 | 20210048.023 | 547 | 0.832 | 4182 | 11.99 | 886 | -30.766 | 14.6859 | 49.0475
05/17/21 | 41 | 20210048.024 | 549 | 0.821 | 4132 | 12.51 | 822 | -31.256 | 14.7625 | 46.1032
05/17/21 | 42 | 20210048.025 | 551 | 0.83 | 3591 | 13.091 | 945 | -34.044 | 12.6304 | 52.3387
05/17/21 | 43 | 20210048.026 | 553 | 0.826 | 4149 | 12.83 | 873 | -32.839 | 14.6545 | 48.5703
05/17/21 | 44 | 20210048.027 | 555 | 0.827 | 3924 | 12.524 | 910 | -33.221 | 13.8575 | 50.5809
05/17/21 | 45 | 20210048.028 | 557 | 0.82 | 3922 | 12.513 | 902 | -33.315 | 13.9678 | 50.5731
05/17/21 | 46 | 20210048.029 | 559 | 0.84 | 4188 | 12.549 | 895 | -31.919 | 14.5706 | 49.0964
05/17/21 | 47 | 20210048.030 | 561 | 0.843 | 4307 | 9.411 | 872 | -32.66 | 14.9226 | 47.5889
05/17/21 | 48 | 20210048.031 | 563 | 0.844 | 4286 | 12.97 | 884 | -31.669 | 14.8483 | 48.2539
05/17/21 | 49 | 20210048.032 | 565 | 0.847 | 4325 | 9.562 | 895 | -32.757 | 14.9346 | 48.8057
05/17/21 | 50 | 20210048.033 | 567 | 0.82 | 3845 | 13.209 | 907 | -32.658 | 13.7048 | 50.9533
05/17/21 | 51 | 20210048.034 | 569 | 0.837 | 4231 | 9.589 | 851 | -32.431 | 14.7879 | 46.8087
05/17/21 | 52 | 20210048.035 | 571 | 0.848 | 4289 | 9.372 | 871 | -32.726 | 14.7386 | 47.2049
05/17/21 | 53 | 20210048.036 | 573 | 0.802 | 3945 | 12.325 | 855 | -31.863 | 14.3844 | 49.0499
05/17/21 | 54 | 20210048.037 | 575 | 0.832 | 3966 | 10.273 | 885 | -32.934 | 13.9285 | 48.9967
05/17/21 | 55 | 20210048.038 | 577 | 0.826 | 4164 | 10.112 | 859 | -32.359 | 14.704 | 47.7977
05/17/21 | 56 | 20210048.039 | 579 | 0.836 | 3060 | 12.481 | 1009 | -32.499 | 10.6961 | 55.5989
05/17/21 | 57 | 20210048.040 | 581 | 0.836 | 3989 | 11.349 | 889 | -32.768 | 13.9399 | 48.9235
05/17/21 | 58 | 20210048.041 | 583 | 0.836 | 3378 | 12.947 | 988 | -34.764 | 11.773 | 54.319
05/17/21 | 59 | 20210048.042 | 585 | 0.839 | 4148 | 10.235 | 871 | -32.395 | 14.4545 | 47.8951
05/17/21 | 60 | 20210048.042 | Replicate E11 | 0.832 | 4116 | 10.245 | 866 | -32.312 | 14.4219 | 47.8063
05/17/21 | 61 | 20210048.043 | 587 | 0.805 | 2431 | 12.831 | 994 | -32.999 | 8.8494 | 56.7517
05/18/21 | 62 | 20210048.044 | 589 | 0.818 | 3803 | 12.376 | 901 | -33.411 | 13.5258 | 50.5618
05/18/21 | 63 | 20210048.045 | 591 | 0.812 | 3604 | 13.26 | 877 | -33.542 | 12.9334 | 49.6396
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.18 | 36-UWSIF-UT Glut 1- | 0.783 | 2484 | -2.853 | 694 | -29.77 | 9.3773 | 40.6804
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.19 | 36-UWSIF-UT Glut 1- | 0.794 | 2519 | -2.841 | 703 | -29.887 | 9.3719 | 40.624
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 39-UWSIF-UW Glut 2- | 0.735 | 2520 | 29.733 | 703 | 22.538 | 10.1607 | 44.033
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 39-UWSIF-UW Glut 2- | 0.722 | 2487 | 29.732 | 694 | 22.268 | 10.2194 | 44.2906
05/18/21 | 68 | 01-UWSIF-Liver-20210048.15 | 01-UWSIF-Liver- | 0.728 | 3215 | 8.523 | 803 | -19.338 | 12.9184 | 50.5363
05/18/21 | 69 | 01-UWSIF-Liver-20210048.16 | 01-UWSIF-Liver- | 0.738 | 3269 | 8.516 | 816 | -19.393 | 12.9316 | 50.6414
05/18/21 | 70 | 20210048.046 | 593 | 0.822 | 4152 | 12.73 | 872 | -31.499 | 14.7289 | 48.7632
05/18/21 | 71 | 20210048.047 | 595 | 0.826 | 3983 | 9.002 | 894 | -33.458 | 14.0177 | 49.6868
05/18/21 | 72 | 20210048.048 | 597 | 0.849 | 4276 | 9.143 | 886 | -31.891 | 14.6869 | 48.0233
05/18/21 | 73 | 20210048.049 | 599 | 0.807 | 3983 | 10.379 | 853 | -32.602 | 14.3735 | 48.5247
05/18/21 | 74 | 20210048.050 | 601 | 0.833 | 4005 | 12.882 | 890 | -32.978 | 14.0076 | 49.0704
05/18/21 | 75 | 20210048.051 | 603 | 0.822 | 3921 | 10.957 | 880 | -32.799 | 13.9009 | 49.1306
05/18/21 | 76 | 20210048.052 | 605 | 0.85 | 4231 | 10.06 | 893 | -32.559 | 14.4587 | 48.2495
05/18/21 | 77 | 20210048.053 | 607 | 0.818 | 4068 | 12.396 | 848 | -31.584 | 14.4838 | 47.6062
05/18/21 | 78 | 20210048.054 | 609 | 0.833 | 3642 | 12.405 | 965 | -33.279 | 12.7149 | 53.3003
05/18/21 | 79 | 20210048.055 | 611 | 0.826 | 3246 | 12.698 | 960 | -33.18 | 11.4487 | 53.3883
05/18/21 | 80 | 20210048.056 | 613 | 0.801 | 4030 | 9.36 | 838 | -32.673 | 14.6584 | 47.9841
05/18/21 | 81 | 20210048.057 | 615 | 0.801 | 3439 | 9.342 | 890 | -34.063 | 12.4959 | 50.9964
05/18/21 | 82 | 20210048.058 | 617 | 0.848 | 3887 | 10.602 | 920 | -33.13 | 13.3194 | 49.8357
05/18/21 | 83 | 20210048.059 | 619 | 0.839 | 4057 | 13.709 | 863 | -32.012 | 14.0637 | 47.241
05/18/21 | 84 | 20210048.060 | 623 | 0.817 | 4097 | 12.631 | 865 | -32.421 | 14.5562 | 48.5349
05/18/21 | 85 | 20210048.061 | 625 | 0.813 | 4005 | 12.791 | 855 | -32.506 | 14.3776 | 48.2924
05/18/21 | 86 | 20210048.062 | 627 | 0.816 | 3791 | 12.504 | 861 | -33.077 | 13.5573 | 48.4113
05/18/21 | 87 | 20210048.063 | 629 | 0.843 | 4210 | 10.196 | 873 | -32.288 | 14.5186 | 47.5388
05/18/21 | 88 | 20210048.064 | 631 | 0.825 | 3648 | 12.644 | 901 | -32.41 | 12.8806 | 50.1564
05/18/21 | 89 | 20210048.065 | 635 | 0.821 | 4069 | 11.307 | 878 | -32.128 | 14.4124 | 49.1291
05/18/21 | 90 | 20210048.066 | 637 | 0.836 | 4090 | 12.918 | 888 | -32.717 | 14.2609 | 48.8186
05/18/21 | 91 | 20210048.067 | 639 | 0.83 | 3947 | 13.182 | 875 | -32.597 | 13.8572 | 48.4888
05/18/21 | 92 | 20210048.068 | 641 | 0.848 | 4252 | 10.997 | 870 | -32.064 | 14.6138 | 47.157
05/18/21 | 93 | 20210048.068 | Replicate H8 | 0.828 | 4218 | 11.025 | 860 | -32.029 | 14.7989 | 47.6548
05/18/21 | 94 | 36-UWSIF-UT Glut 1-20210048.111 | 36-UWSIF-UT Glut 1- | 0.754 | 2364 | -2.858 | 664 | -29.871 | 9.2965 | 40.3414
05/18/21 | 95 | 36-UWSIF-UT Glut 1-20210048.112 | 36-UWSIF-UT Glut 1- | 0.787 | 2474 | -2.84 | 692 | -29.818 | 9.289 | 40.3623
05/18/21 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 39-UWSIF-UW Glut 2- | 0.703 | 2394 | 29.779 | 670 | 22.266 | 10.1099 | 43.7854
05/18/21 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 39-UWSIF-UW Glut 2- | 0.83 | 2829 | 29.692 | 787 | 22.35 | 10.0426 | 43.6353
05/18/21 | 98 | 01-UWSIF-Liver-20210048.17 | 01-UWSIF-Liver- | 0.753 | 3279 | 8.488 | 818 | -19.536 | 12.7475 | 49.8914
05/18/21 | 99 | 01-UWSIF-Liver-20210048.18 | 01-UWSIF-Liver- | 0.735 | 3199 | 8.539 | 801 | -19.438 | 12.7605 | 50.002
Comment
JOB: 2021-0048 Tray 1 Nelson
ANALYST: GH
SAMPLE COUNT: 0009
HE PSI: 2200
O2 PSI: 500
MASS 28: 42
MASS 29: 32
MASS 30: 748

## Sorted 2
Date | Row | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | %N | %C
05/18/21 | 1 | 36-UWSIF-UT Glut 1-20210048.21 | 36-UWSIF-UT Glut 1- | 0.769 | 2461 | -2.774 | 685 | -29.792 | 9.5601 | 41.0465
05/18/21 | 2 | 36-UWSIF-UT Glut 1-20210048.22 | 36-UWSIF-UT Glut 1- | 0.824 | 2532 | -2.754 | 710 | -29.974 | 9.4545 | 40.4581
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 | 1441 | -2.877 | 407 | -29.845 | 9.4495 | 40.5235
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 | 2250 | -2.812 | 634 | -29.899 | 9.5877 | 41.0583
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 | 4838 | -2.865 | 1304 | -30.108 | 9.5163 | 40.7823
05/18/21 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 39-UWSIF-UW Glut 2- | 0.742 | 2482 | 29.77 | 698 | 22.086 | 10.2938 | 44.2125
05/18/21 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 39-UWSIF-UW Glut 2- | 0.736 | 2330 | 29.795 | 657 | 22.324 | 9.7602 | 41.9235
05/18/21 | 8 | 01-UWSIF-Liver-20210048.21 | 01-UWSIF-Liver- | 0.723 | 3081 | 8.603 | 775 | -19.377 | 13.059 | 50.3669
05/18/21 | 9 | 01-UWSIF-Liver-20210048.22 | 01-UWSIF-Liver- | 0.717 | 3071 | 8.584 | 773 | -19.431 | 13.0774 | 50.4634
05/18/21 | 10 | 20210048.069 | 643 | 0.806 | 3894 | 12.548 | 844 | -31.649 | 14.7932 | 49.2795
05/18/21 | 11 | 20210048.070 | 645 | 0.814 | 4036 | 11.909 | 846 | -31.644 | 15.1732 | 48.9549
05/18/21 | 12 | 20210048.070 | Replicate A11 | 0.84 | 4158 | 11.929 | 872 | -31.616 | 15.1423 | 48.9229
05/18/21 | 13 | 20210048.071 | 651 | 0.824 | 4064 | 8.812 | 838 | -31.711 | 15.0998 | 47.8608
05/18/21 | 14 | 20210048.072 | 653 | 0.845 | 4071 | 8.639 | 881 | -32.292 | 14.6697 | 48.9291
05/18/21 | 15 | 20210048.073 | 655 | 0.84 | 4065 | 12.569 | 888 | -31.397 | 14.7152 | 49.6143
05/18/21 | 16 | 20210048.074 | 657 | 0.837 | 4012 | 8.67 | 873 | -31.821 | 14.6388 | 49.0862
05/18/21 | 17 | 20210048.075 | 659 | 0.838 | 3833 | 11.905 | 903 | -29.306 | 13.9163 | 50.7763
05/18/21 | 18 | 20210048.076 | 661 | 0.846 | 4009 | 12.152 | 862 | -30.974 | 14.4542 | 47.9373
05/18/21 | 19 | 20210048.077 | 663 | 0.818 | 4046 | 12.894 | 839 | -32.338 | 15.1015 | 48.2594
05/18/21 | 20 | 20210048.078 | 665 | 0.824 | 3053 | 12.812 | 878 | -34.457 | 11.3278 | 50.2376
05/18/21 | 21 | 20210048.079 | 667 | 0.802 | 3890 | 10.027 | 819 | -32.76 | 14.7922 | 47.7844
05/18/21 | 22 | 20210048.080 | 669 | 0.836 | 4143 | 12.39 | 856 | -31.62 | 15.0876 | 47.9493
05/18/21 | 23 | 20210048.081 | 671 | 0.828 | 4104 | 9.875 | 848 | -32.656 | 15.131 | 47.909
05/18/21 | 24 | 20210048.082 | 673 | 0.816 | 3965 | 12.267 | 828 | -32.448 | 14.8596 | 47.5687
05/18/21 | 25 | 20210048.083 | 703 | 0.832 | 3511 | 10.044 | 901 | -33.758 | 12.8267 | 50.8828
05/18/21 | 26 | 20210048.084 | 705 | 0.801 | 3498 | 13.419 | 856 | -32.737 | 13.3644 | 50.2173
05/18/21 | 27 | 20210048.085 | 709 | 0.835 | 4087 | 8.927 | 858 | -32.089 | 14.928 | 48.1847
05/18/21 | 28 | 20210048.086 | 713 | 0.829 | 3576 | 12.342 | 885 | -31.08 | 13.1582 | 50.1952
05/18/21 | 29 | 20210048.087 | 715 | 0.81 | 3136 | 10.601 | 935 | -36.106 | 11.7339 | 54.2136
05/18/21 | 30 | 20210048.088 | 719 | 0.829 | 3940 | 12.448 | 859 | -32.646 | 14.3995 | 48.3685
05/18/21 | 31 | 20210048.089 | 721 | 0.814 | 3828 | 12.965 | 845 | -32.968 | 14.3432 | 48.6707
05/18/21 | 32 | 20210048.090 | 723 | 0.825 | 4061 | 9.013 | 841 | -32.878 | 14.9333 | 47.6555
05/18/21 | 33 | 20210048.090 | Replicate C8 | 0.81 | 4024 | 9.098 | 831 | -32.931 | 15.1184 | 47.9481
05/18/21 | 34 | 36-UWSIF-UT Glut 1-20210048.26 | 36-UWSIF-UT Glut 1- | 0.731 | 2255 | -2.832 | 627 | -29.752 | 9.4356 | 39.866
05/18/21 | 35 | 36-UWSIF-UT Glut 1-20210048.27 | 36-UWSIF-UT Glut 1- | 0.829 | 2572 | -2.801 | 714 | -29.84 | 9.432 | 39.9942
05/18/21 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 39-UWSIF-UW Glut 2- | 0.769 | 2594 | 29.79 | 720 | 22.208 | 10.2865 | 43.6591
05/18/21 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 39-UWSIF-UW Glut 2- | 0.841 | 2808 | 29.667 | 779 | 22.32 | 10.193 | 43.3863
05/18/21 | 38 | 01-UWSIF-Liver-20210048.23 | 01-UWSIF-Liver- | 0.736 | 3138 | 8.547 | 785 | -19.604 | 12.9558 | 49.8053
05/18/21 | 39 | 01-UWSIF-Liver-20210048.24 | 01-UWSIF-Liver- | 0.731 | 3123 | 8.548 | 781 | -19.591 | 12.9834 | 49.8134
05/18/21 | 40 | 20210048.091 | 725 | 0.827 | 3783 | 12.448 | 881 | -33.413 | 13.9065 | 49.9696
05/18/21 | 41 | 20210048.092 | 727 | 0.807 | 3914 | 10.218 | 837 | -32.867 | 14.6505 | 48.413
05/18/21 | 42 | 20210048.093 | 729 | 0.85 | 4022 | 13.663 | 894 | -32.187 | 14.3059 | 49.1498
05/18/21 | 43 | 20210048.094 | 731 | 0.823 | 3563 | 12.792 | 908 | -33.942 | 13.0812 | 51.6628
05/18/21 | 44 | 20210048.095 | 733 | 0.837 | 4184 | 12.792 | 854 | -31.889 | 15.1021 | 47.654
05/18/21 | 45 | 20210048.096 | 735 | 0.85 | 3998 | 9.97 | 893 | -33.027 | 14.1989 | 49.1359
05/18/21 | 46 | 20210048.097 | 739 | 0.826 | 3972 | 11.513 | 835 | -32.072 | 14.522 | 47.1583
05/18/21 | 47 | 20210048.098 | 741 | 0.825 | 3878 | 13.087 | 851 | -31.878 | 14.2719 | 48.2899
05/18/21 | 48 | 20210048.099 | 747 | 0.849 | 4041 | 10.437 | 892 | -32.791 | 14.3789 | 49.1898
05/18/21 | 49 | 20210048.100 | 847 | 0.815 | 2397 | 11.908 | 964 | -35.331 | 8.8645 | 55.6604
05/18/21 | 50 | 20210048.101 | 849 | 0.833 | 3440 | 12.408 | 922 | -34.271 | 12.4706 | 51.908
05/18/21 | 51 | 20210048.102 | 851 | 0.82 | 3904 | 10.429 | 818 | -30.963 | 14.3991 | 46.5687
05/18/21 | 52 | 20210048.103 | 853 | 0.833 | 3722 | 11.684 | 887 | -32.471 | 13.4686 | 49.7576
05/18/21 | 53 | 20210048.104 | 855 | 0.825 | 3690 | 12.721 | 872 | -31.982 | 13.534 | 49.44
05/18/21 | 54 | 20210048.105 | 857 | 0.848 | 3649 | 12.047 | 921 | -32.47 | 13.0178 | 51.0336
05/18/21 | 55 | 20210048.106 | 859 | 0.848 | 3771 | 12.024 | 894 | -32.933 | 13.4805 | 49.4608
05/18/21 | 56 | 20210048.107 | 861 | 0.806 | 3701 | 12.5 | 847 | -32.174 | 13.8992 | 49.1831
05/18/21 | 57 | 20210048.108 | 863 | 0.812 | 3878 | 9.885 | 826 | -28.967 | 14.4835 | 47.5784
05/18/21 | 58 | 20210048.109 | 865 | 0.827 | 3753 | 12.073 | 874 | -31.502 | 13.7352 | 49.596
05/18/21 | 59 | 20210048.110 | 867 | 0.816 | 2236 | 13.475 | 1032 | -35.787 | 8.254 | 59.7406
05/18/21 | 60 | 20210048.110 | Replicate E11 | 0.845 | 2727 | 13.506 | 1032 | -34.933 | 9.7175 | 57.5684
05/18/21 | 61 | 20210048.111 | 869 | 0.834 | [6 empty] | Sample not in tray when loaded into carousel
05/18/21 | 62 | 20210048.112 | 871 | 0.825 | 3917 | 11.976 | 843 | -31.703 | 14.389 | 47.8622
05/18/21 | 63 | 20210048.113 | 873 | 0.839 | 4048 | 11.291 | 850 | -31.78 | 14.638 | 47.4881
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.28 | 36-UWSIF-UT Glut 1- | 0.827 | 2538 | -2.827 | 717 | -29.749 | 9.3064 | 40.3644
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.29 | 36-UWSIF-UT Glut 1- | 0.791 | 2411 | -2.848 | 680 | -29.812 | 9.3123 | 40.2144
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 39-UWSIF-UW Glut 2- | 0.716 | 2387 | 29.695 | 675 | 22.277 | 10.1644 | 44.0236
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 39-UWSIF-UW Glut 2- | 0.798 | 2665 | 29.723 | 753 | 22.299 | 10.1586 | 44.1332
05/18/21 | 68 | 01-UWSIF-Liver-20210048.25 | 01-UWSIF-Liver- | 0.741 | 3140 | 8.538 | 795 | -19.359 | 12.8421 | 50.1619
05/18/21 | 69 | 01-UWSIF-Liver-20210048.26 | 01-UWSIF-Liver- | 0.767 | 3266 | 8.529 | 823 | -19.484 | 12.9159 | 50.2422
05/18/21 | 70 | 20210048.114 | 875 | 0.826 | 3949 | 9.294 | 845 | -31.478 | 14.5257 | 48.1374
05/18/21 | 71 | 20210048.115 | 877 | 0.832 | 3197 | 12.098 | 947 | -35.144 | 11.6242 | 53.5877
05/18/21 | 72 | 20210048.116 | 879 | 0.824 | 4014 | 12.092 | 836 | -31.314 | 14.7611 | 47.49
05/18/21 | 73 | 20210048.117 | 881 | 0.845 | 4095 | 11.592 | 857 | -31.727 | 14.7341 | 47.6085
05/18/21 | 74 | 20210048.118 | 883 | 0.841 | 3957 | 12.331 | 848 | -31.341 | 14.3076 | 47.2728
05/18/21 | 75 | 20210048.119 | 885 | 0.807 | 3250 | 13.858 | 902 | -32.93 | 12.2251 | 52.5804
05/18/21 | 76 | 20210048.120 | 887 | 0.849 | 4153 | 11.666 | 857 | -30.451 | 14.8365 | 47.3762
05/18/21 | 77 | 20210048.121 | 889 | 0.811 | 3910 | 11.857 | 819 | -30.548 | 14.6396 | 47.2728
05/18/21 | 78 | 20210048.122 | 891 | 0.801 | 3644 | 12.94 | 820 | -31.594 | 13.8194 | 47.9241
05/18/21 | 79 | 20210048.123 | 893 | 0.813 | 3699 | 11.211 | 813 | -32.459 | 13.8323 | 46.8097
05/18/21 | 80 | 20210048.124 | 895 | 0.812 | 3650 | 12.26 | 871 | -32.612 | 13.6862 | 50.3854
05/18/21 | 81 | 20210048.125 | 897 | 0.841 | 3892 | 11.603 | 848 | -32.28 | 14.041 | 47.2587
05/19/21 | 82 | 20210048.126 | 899 | 0.81 | 3525 | 9.629 | 851 | -32.003 | 13.1978 | 49.2717
05/19/21 | 83 | 20210048.127 | 901 | 0.811 | 3097 | 11.444 | 907 | -34.162 | 11.5576 | 52.5355
05/19/21 | 84 | 20210048.128 | 903 | 0.832 | 3624 | 11.885 | 897 | -33.744 | 13.1859 | 50.5094
05/19/21 | 85 | 20210048.129 | 905 | 0.8 | 3140 | 11.456 | 898 | -34.205 | 11.8749 | 52.6395
05/19/21 | 86 | 20210048.130 | 907 | 0.846 | 3482 | 11.942 | 945 | -34.24 | 12.4451 | 52.6011
05/19/21 | 87 | 20210048.131 | 909 | 0.807 | 3815 | 11.722 | 831 | -32.495 | 14.3419 | 48.2331
05/19/21 | 88 | 20210048.132 | 911 | 0.844 | 3993 | 12.154 | 866 | -32.566 | 14.3359 | 48.0809
05/19/21 | 89 | 20210048.133 | 913 | 0.815 | 3823 | 11.86 | 850 | -32.299 | 14.2056 | 48.8886
05/19/21 | 90 | 20210048.134 | 915 | 0.838 | 3783 | 11.953 | 887 | -33.38 | 13.6658 | 49.6924
05/19/21 | 91 | 20210048.135 | 917 | 0.822 | 3981 | 12.022 | 833 | -32.137 | 14.6687 | 47.4433
05/19/21 | 92 | 20210048.136 | 919 | 0.817 | 3644 | 11.517 | 856 | -33.092 | 13.5156 | 49.081
05/19/21 | 93 | 20210048.136 | Replicate H8 | 0.817 | 3677 | 11.473 | 857 | -33.199 | 13.6416 | 49.174
05/19/21 | 94 | 36-UWSIF-UT Glut 1-20210048.211 | 36-UWSIF-UT Glut 1- | 0.744 | 2271 | -2.888 | 645 | -29.709 | 9.2845 | 40.2746
05/19/21 | 95 | 36-UWSIF-UT Glut 1-20210048.212 | 36-UWSIF-UT Glut 1- | 0.755 | 2305 | -2.906 | 654 | -29.902 | 9.2734 | 40.2066
05/19/21 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 39-UWSIF-UW Glut 2- | 0.73 | 2418 | 29.767 | 680 | 22.186 | 10.1409 | 43.6557
05/19/21 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 39-UWSIF-UW Glut 2- | 0.768 | 2523 | 29.664 | 711 | 22.303 | 10.0161 | 43.3049
05/19/21 | 98 | 01-UWSIF-Liver-20210048.27 | 01-UWSIF-Liver- | 0.825 | 3496 | 8.476 | 877 | -19.405 | 12.8169 | 49.8267
05/19/21 | 99 | 01-UWSIF-Liver-20210048.28 | 01-UWSIF-Liver- | 0.736 | 3105 | 8.488 | 783 | -19.351 | 12.8068 | 49.7227
Comment
JOB: 2021-0048 Tray 2 Nelson
ANALYST: sme
SAMPLE COUNT: 0108
HE PSI: 1900
O2 PSI: 500
MASS 28:
MASS 29:
MASS 30:

## Sorted 3
Date | Row | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | %N | %C
05/19/21 | 1 | 36-UWSIF-UT Glut 1-20210048.31 | 36-UWSIF-UT Glut 1- | 0.795 | 2410 | -2.778 | 672 | -29.798 | 9.5882 | 40.6269
05/19/21 | 2 | 36-UWSIF-UT Glut 1-20210048.32 | 36-UWSIF-UT Glut 1- | 0.72 | 2125 | -2.828 | 604 | -30.029 | 9.5023 | 40.6284
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 | 1249 | -2.831 | 357 | -29.73 | 9.5228 | 41.0425
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 | 2223 | -2.847 | 628 | -29.877 | 9.5195 | 40.627
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 | 4785 | -2.893 | 1297 | -30.115 | 9.5211 | 40.8337
05/19/21 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 39-UWSIF-UW Glut 2- | 0.8 | 2580 | 29.724 | 728 | 22.144 | 10.3266 | 44.131
05/19/21 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 39-UWSIF-UW Glut 2- | 0.829 | 2680 | 29.762 | 755 | 22.232 | 10.341 | 44.1252
05/19/21 | 8 | 01-UWSIF-Liver-20210048.31 | 01-UWSIF-Liver- | 0.811 | 3340 | 8.509 | 844 | -19.515 | 13.101 | 50.3372
05/19/21 | 9 | 01-UWSIF-Liver-20210048.32 | 01-UWSIF-Liver- | 0.719 | 2940 | 8.452 | 747 | -19.55 | 13.0377 | 50.1599
05/19/21 | 10 | 20210048.137 | 921 | 0.846 | 3401 | 11.977 | 716 | -31.01 | 12.7789 | 40.8389
05/19/21 | 11 | 20210048.138 | 923 | 0.842 | 3675 | 12.085 | 884 | -33.528 | 13.7796 | 50.7746
05/19/21 | 12 | 20210048.138 | Replicate A11 | 0.829 | 3366 | 12.112 | 887 | -34.148 | 12.8465 | 51.782
05/19/21 | 13 | 20210048.139 | 925 | 0.801 | 3166 | 11.553 | 852 | -33.665 | 12.4896 | 51.4176
05/19/21 | 14 | 20210048.140 | 927 | 0.839 | 3573 | 12.023 | 879 | -33.66 | 13.4429 | 50.6637
05/19/21 | 15 | 20210048.141 | 929 | 0.843 | 3486 | 11.08 | 902 | -33.344 | 13.053 | 51.6324
05/19/21 | 16 | 20210048.142 | 931 | 0.812 | 3675 | 10.748 | 799 | -30.192 | 14.3203 | 47.4286
05/19/21 | 17 | 20210048.143 | 933 | 0.83 | 3740 | 9.55 | 849 | -32.288 | 14.2201 | 49.3048
05/19/21 | 18 | 20210048.144 | 935 | 0.825 | 3828 | 11.789 | 827 | -32.141 | 14.6815 | 48.495
05/19/21 | 19 | 20210048.145 | 937 | 0.84 | 2573 | 12.043 | 994 | -35.9 | 9.6718 | 57.4919
05/19/21 | 20 | 20210048.146 | 939 | 0.832 | 3539 | 11.959 | 876 | -33.604 | 13.4212 | 50.7269
05/19/21 | 21 | 20210048.147 | 941 | 0.814 | 3149 | 12.63 | 909 | -34.17 | 12.2121 | 53.8261
05/19/21 | 22 | 20210048.148 | 943 | 0.805 | 3678 | 12.729 | 821 | -31.437 | 14.432 | 49.1264
05/19/21 | 23 | 20210048.149 | 945 | 0.84 | 3436 | 12.043 | 915 | -34.253 | 12.8917 | 52.589
05/19/21 | 24 | 20210048.150 | 947 | 0.829 | 3487 | 12.022 | 881 | -33.772 | 13.2954 | 51.1734
05/19/21 | 25 | 20210048.151 | 949 | 0.826 | 3706 | 12.087 | 851 | -32.762 | 14.1459 | 49.5768
05/19/21 | 26 | 20210048.152 | 951 | 0.843 | 3951 | 12.309 | 846 | -31.093 | 14.7276 | 48.1699
05/19/21 | 27 | 20210048.153 | 953 | 0.827 | 3473 | 13.363 | 869 | -32.644 | 13.2715 | 50.7079
05/19/21 | 28 | 20210048.154 | 955 | 0.803 | 3584 | 9.979 | 824 | -31.884 | 14.0425 | 49.3257
05/19/21 | 29 | 20210048.155 | 957 | 0.846 | 4029 | 11.627 | 835 | -31.969 | 14.9942 | 47.4616
05/19/21 | 30 | 20210048.156 | 959 | 0.846 | 4035 | 12.117 | 849 | -29.876 | 15.0442 | 48.3398
05/19/21 | 31 | 20210048.157 | 961 | 0.824 | 3329 | 13.074 | 881 | -32.89 | 12.7703 | 51.5068
05/19/21 | 32 | 20210048.158 | 963 | 0.827 | 3793 | 9.809 | 840 | -31.099 | 14.4959 | 48.8611
05/19/21 | 33 | 20210048.158 | Replicate C8 | 0.821 | 3738 | 9.841 | 818 | -30.989 | 14.3791 | 47.9662
05/19/21 | 34 | 36-UWSIF-UT Glut 1-20210048.36 | 36-UWSIF-UT Glut 1- | 0.801 | 2382 | -2.943 | 674 | -29.867 | 9.4866 | 40.3482
05/19/21 | 35 | 36-UWSIF-UT Glut 1-20210048.37 | 36-UWSIF-UT Glut 1- | 0.786 | 2342 | -2.875 | 663 | -29.855 | 9.4862 | 40.5048
05/19/21 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 39-UWSIF-UW Glut 2- | 0.779 | 2511 | 29.645 | 706 | 22.272 | 10.2773 | 43.6597
05/19/21 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 39-UWSIF-UW Glut 2- | 0.726 | 2350 | 29.79 | 660 | 22.207 | 10.3461 | 43.753
05/19/21 | 38 | 01-UWSIF-Liver-20210048.33 | 01-UWSIF-Liver- | 0.75 | 3087 | 8.461 | 778 | -19.317 | 13.0074 | 49.7454
05/19/21 | 39 | 01-UWSIF-Liver-20210048.34 | 01-UWSIF-Liver- | 0.79 | 3251 | 8.439 | 817 | -19.31 | 12.9985 | 49.7138
05/19/21 | 40 | 20210048.159 | 965 | 0.825 | 3728 | 12.89 | 848 | -30.661 | 14.2494 | 49.4185
05/19/21 | 41 | 20210048.160 | 967 | 0.807 | 3330 | 11.46 | 890 | -34.016 | 13.0246 | 53.0575
05/19/21 | 42 | 20210048.161 | 969 | 0.809 | 3572 | 11.974 | 830 | -33.071 | 13.9622 | 49.3035
05/19/21 | 43 | 20210048.162 | 971 | 0.808 | 3329 | 11.648 | 867 | -33.489 | 12.9792 | 51.5328
05/19/21 | 44 | 20210048.163 | 973 | 0.826 | 3271 | 11.517 | 909 | -34.023 | 12.4862 | 52.9922
05/19/21 | 45 | 20210048.164 | 975 | 0.833 | 3733 | 11.822 | 852 | -32.845 | 14.1257 | 49.1767
05/19/21 | 46 | 20210048.165 | 749 | 0.842 | 4079 | 11.936 | 838 | -31.845 | 15.2703 | 47.7808
05/19/21 | 47 | 20210048.166 | 751 | 0.808 | 3503 | 11.085 | 832 | -31.533 | 13.6797 | 49.4367
05/19/21 | 48 | 20210048.167 | 753 | 0.827 | 3552 | 12.24 | 844 | -32.914 | 13.5426 | 49.0354
05/19/21 | 49 | 20210048.168 | 755 | 0.808 | 3639 | 12.506 | 821 | -32.251 | 14.1902 | 48.687
05/19/21 | 50 | 20210048.169 | 757 | 0.838 | 3540 | 11.715 | 877 | -32.679 | 13.308 | 50.3187
05/19/21 | 51 | 20210048.170 | 759 | 0.84 | 3943 | 12.063 | 846 | -29.339 | 14.7876 | 48.3501
05/19/21 | 52 | 20210048.171 | 761 | 0.816 | 2699 | 12.088 | 932 | -33.478 | 10.4086 | 54.979
05/19/21 | 53 | 20210048.172 | 763 | 0.803 | 3527 | 12.382 | 820 | -29.649 | 13.8225 | 48.9367
05/19/21 | 54 | 20210048.173 | 765 | 0.817 | 3676 | 11.275 | 818 | -31.919 | 14.1585 | 47.9658
05/19/21 | 55 | 20210048.174 | 767 | 0.827 | 3747 | 11.866 | 825 | -32.187 | 14.2313 | 47.7739
05/19/21 | 56 | 20210048.175 | 769 | 0.824 | 3328 | 11.862 | 851 | -33.38 | 12.6803 | 49.5222
05/19/21 | 57 | 20210048.176 | 771 | 0.837 | 3801 | 12.088 | 831 | -29.962 | 14.2434 | 47.4758
05/19/21 | 58 | 20210048.177 | 773 | 0.83 | 3793 | 11.284 | 830 | -31.856 | 14.3383 | 47.9216
05/19/21 | 59 | 20210048.178 | 775 | 0.82 | 3233 | 11.987 | 906 | -33.734 | 12.3557 | 53.0346
05/19/21 | 60 | 20210048.178 | Replicate E11 | 0.836 | 3739 | 12.051 | 883 | -32.671 | 14.0262 | 50.7223
05/19/21 | 61 | 20210048.179 | 777 | 0.845 | 3927 | 11.57 | 830 | -30.568 | 14.5829 | 47.0771
05/19/21 | 62 | 20210048.180 | 779 | 0.85 | 3312 | 11.1 | 903 | -32.516 | 12.2078 | 51.0175
05/19/21 | 63 | 20210048.181 | 781 | 0.81 | 3642 | 13.238 | 822 | -32.603 | 14.1266 | 48.5731
05/19/21 | 64 | 36-UWSIF-UT Glut 1-20210048.38 | 36-UWSIF-UT Glut 1- | 0.813 | 2426 | -2.897 | 689 | -29.809 | 9.3888 | 40.331
05/19/21 | 65 | 36-UWSIF-UT Glut 1-20210048.39 | 36-UWSIF-UT Glut 1- | 0.777 | 2310 | -2.858 | 654 | -29.891 | 9.3951 | 40.1745
05/19/21 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 39-UWSIF-UW Glut 2- | 0.766 | 2463 | 29.66 | 699 | 22.281 | 10.1668 | 43.6233
05/19/21 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 39-UWSIF-UW Glut 2- | 0.757 | 2434 | 29.59 | 691 | 22.319 | 10.1463 | 43.6093
05/19/21 | 68 | 01-UWSIF-Liver-20210048.35 | 01-UWSIF-Liver- | 0.771 | 3149 | 8.504 | 798 | -19.47 | 12.8204 | 49.4922
05/19/21 | 69 | 01-UWSIF-Liver-20210048.36 | 01-UWSIF-Liver- | 0.765 | 3126 | 8.477 | 790 | -19.398 | 12.8336 | 49.4161
05/19/21 | 70 | 20210048.182 | 783 | 0.808 | 3525 | 12.036 | 817 | -29.733 | 13.6847 | 48.3642
05/19/21 | 71 | 20210048.183 | 785 | 0.801 | 3042 | 12.168 | 875 | -32.422 | 11.8994 | 52.4117
05/19/21 | 72 | 20210048.184 | 787 | 0.815 | 3420 | 11.486 | 845 | -32.886 | 13.1625 | 49.6648
05/19/21 | 73 | 20210048.185 | 789 | 0.805 | 3737 | 10.734 | 798 | -30.796 | 14.5529 | 47.3997
05/19/21 | 74 | 20210048.186 | 791 | 0.848 | 3610 | 10.47 | 840 | -32.021 | 13.3536 | 47.4225
05/19/21 | 75 | 20210048.187 | 793 | 0.83 | 3535 | 12.036 | 851 | -32.26 | 13.3527 | 49.1869
05/19/21 | 76 | 20210048.188 | 795 | 0.845 | 3770 | 11.855 | 867 | -32.324 | 13.9805 | 49.2479
05/19/21 | 77 | 20210048.189 | 797 | 0.826 | 3642 | 12.351 | 845 | -32.25 | 13.8267 | 48.9535
05/19/21 | 78 | 20210048.190 | 799 | 0.845 | 3824 | 10.523 | 812 | -30.995 | 14.1741 | 45.9723
05/19/21 | 79 | 20210048.191 | 801 | 0.848 | 3465 | 11.878 | 908 | -33.069 | 12.787 | 51.4494
05/19/21 | 80 | 20210048.192 | 803 | 0.833 | 3778 | 11.932 | 824 | -31.12 | 14.2302 | 47.403
05/19/21 | 81 | 20210048.193 | 805 | 0.843 | 3974 | 12.232 | 845 | -31.869 | 14.7759 | 48.0091
05/19/21 | 82 | 20210048.194 | 807 | 0.818 | 3464 | 12.681 | 841 | -32.939 | 13.2718 | 49.2225
05/20/21 | 83 | 20210048.195 | 809 | 0.803 | 3549 | 12.029 | 819 | -31.108 | 13.9075 | 48.8085
05/20/21 | 84 | 20210048.196 | 811 | 0.843 | 3949 | 11.651 | 859 | -31.101 | 14.6545 | 48.807
05/20/21 | 85 | 20210048.197 | 813 | 0.831 | 3664 | 9.272 | 829 | -30.328 | 13.8199 | 47.8041
05/20/21 | 86 | 20210048.198 | 815 | 0.814 | 2726 | 11.121 | 940 | -34.145 | 10.4787 | 55.5948
05/20/21 | 87 | 20210048.199 | 817 | 0.802 | 3589 | 12.201 | 801 | -32.236 | 14.0605 | 47.9044
05/20/21 | 88 | 20210048.200 | 819 | 0.825 | 3726 | 11.63 | 830 | -32.44 | 14.1569 | 48.1518
05/20/21 | 89 | 20210048.201 | 821 | 0.831 | 3450 | 11.06 | 887 | -31.805 | 13.0082 | 51.3104
05/20/21 | 90 | 20210048.202 | 823 | 0.81 | 3555 | 13.059 | 819 | -32.344 | 13.7622 | 48.4566
05/20/21 | 91 | 20210048.203 | 825 | 0.837 | 3640 | 13.358 | 848 | -32.148 | 13.6014 | 48.4742
05/20/21 | 92 | 20210048.204 | 827 | 0.836 | 3737 | 12.014 | 844 | -32.394 | 13.9885 | 48.3495
05/20/21 | 93 | 20210048.204 | Replicate H8 | 0.832 | 3676 | 12.002 | 842 | -32.393 | 13.8373 | 48.5015
05/20/21 | 94 | 36-UWSIF-UT Glut 1-20210048.311 | 36-UWSIF-UT Glut 1- | 0.755 | 2214 | -2.888 | 632 | -29.902 | 9.2758 | 40.0269
05/20/21 | 95 | 36-UWSIF-UT Glut 1-20210048.312 | 36-UWSIF-UT Glut 1- | 0.756 | 2218 | -2.934 | 635 | -29.848 | 9.2588 | 40.0261
05/20/21 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 39-UWSIF-UW Glut 2- | 0.76 | 2426 | 29.631 | 691 | 22.346 | 10.0845 | 43.5184
05/20/21 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 39-UWSIF-UW Glut 2- | 0.706 | 2249 | 29.565 | 641 | 22.26 | 10.0794 | 43.4296
05/20/21 | 98 | 01-UWSIF-Liver-20210048.37 | 01-UWSIF-Liver- | 0.73 | 2946 | 8.508 | 751 | -19.579 | 12.6994 | 49.2395
05/20/21 | 99 | 01-UWSIF-Liver-20210048.38 | 01-UWSIF-Liver- | 0.723 | 2929 | 8.43 | 745 | -19.498 | 12.7474 | 49.3033
Comment
JOB: 2021-0048 Tray 3 Nelson
ANALYST: sme
SAMPLE COUNT: 0207
HE PSI: 1600
O2 PSI: 400
MASS 28: 39
MASS 29: 29
MASS 30: 763

## Sorted 4
Date | Row | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | %N | %C
05/20/21 | 1 | 36-UWSIF-UT Glut 1-20210048.41 | 36-UWSIF-UT Glut 1- | 0.794 | 2344 | -2.749 | 652 | -29.868 | 9.5306 | 40.5547
05/20/21 | 2 | 36-UWSIF-UT Glut 1-20210048.42 | 36-UWSIF-UT Glut 1- | 0.781 | 2239 | -2.874 | 638 | -29.831 | 9.4861 | 40.7892
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 | 1339 | -2.815 | 382 | -29.917 | 9.5695 | 40.9744
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 | 2291 | -2.753 | 650 | -29.767 | 9.4851 | 40.6739
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 | 4769 | -2.876 | 1284 | -30.028 | 9.5284 | 40.828
05/20/21 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 39-UWSIF-UW Glut 2- | 0.717 | 2213 | 29.706 | 629 | 22.375 | 10.2508 | 44.0368
05/20/21 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 39-UWSIF-UW Glut 2- | 0.807 | 2510 | 29.772 | 714 | 22.314 | 10.2806 | 44.2723
05/20/21 | 8 | 01-UWSIF-Liver-20210048.41 | 01-UWSIF-Liver- | 0.751 | 2961 | 8.514 | 750 | -19.534 | 12.9981 | 50.1285
05/20/21 | 9 | 01-UWSIF-Liver-20210048.42 | 01-UWSIF-Liver- | 0.809 | 3194 | 8.584 | 808 | -19.538 | 12.9926 | 50.145
05/20/21 | 10 | 20210048.205 | 829 | 0.844 | 3445 | 11.884 | 836 | -31.834 | 13.3758 | 49.778
05/20/21 | 11 | 20210048.206 | 831 | 0.807 | 2736 | 11.556 | 880 | -34.433 | 11.1422 | 54.8929
05/20/21 | 12 | 20210048.206 | Replicate A11 | 0.815 | 2690 | 11.582 | 867 | -34.482 | 10.8072 | 53.451
05/20/21 | 13 | 20210048.207 | 833 | 0.839 | 3641 | 11.888 | 837 | -31.448 | 14.2282 | 50.0282
05/20/21 | 14 | 20210048.208 | 835 | 0.828 | 3312 | 12.12 | 833 | -32.786 | 13.1082 | 50.5279
05/20/21 | 15 | 20210048.209 | 837 | 0.831 | 3319 | 12.522 | 849 | -33.494 | 13.0509 | 51.1799
05/20/21 | 16 | 20210048.210 | 839 | 0.805 | 3182 | 11.818 | 874 | -32.492 | 12.9241 | 54.3287
05/20/21 | 17 | 20210048.211 | 841 | 0.825 | 3319 | 11.895 | 845 | -31.71 | 13.1665 | 51.4145
05/20/21 | 18 | 20210048.212 | 843 | 0.822 | 3437 | 11.733 | 814 | -32.307 | 13.6307 | 49.4706
05/20/21 | 19 | 20210048.213 | 845 | 0.802 | 3635 | 10.58 | 755 | -31.063 | 14.7286 | 46.7683
05/20/21 | 20 | 20210048.214 | 002A | 0.827 | 2421 | 5.741 | 713 | -35.04 | 9.6423 | 43.0279
05/20/21 | 21 | 20210048.215 | 002B | 0.844 | 2514 | 5.722 | 733 | -34.858 | 9.7709 | 43.3295
05/20/21 | 22 | 20210048.216 | 002C | 0.812 | 2470 | 5.615 | 714 | -35.111 | 9.9807 | 43.8315
05/20/21 | 23 | 20210048.217 | 004A | 0.83 | 2824 | 8.532 | 822 | -35.118 | 11.1254 | 49.5493
05/20/21 | 24 | 20210048.218 | 004B | 0.821 | 2407 | 8.536 | 860 | -36.132 | 9.59 | 52.393
05/20/21 | 25 | 20210048.219 | 004C | 0.818 | 2046 | 7.962 | 875 | -37.046 | 8.2043 | 53.5824
05/20/21 | 26 | 20210048.220 | 006A | 0.843 | 2515 | 5.232 | 770 | -34.657 | 9.7748 | 45.5939
05/20/21 | 27 | 20210048.221 | 006B | 0.82 | 2457 | 5.346 | 745 | -34.753 | 9.8236 | 45.3654
05/20/21 | 28 | 20210048.222 | 006C | 0.841 | 2457 | 5.173 | 753 | -34.91 | 9.595 | 44.6923
05/20/21 | 29 | 20210048.223 | 008A | 0.834 | 2864 | 8.267 | 826 | -34.75 | 11.1724 | 49.3266
05/20/21 | 30 | 20210048.224 | 008B | 0.85 | 2741 | 7.812 | 830 | -34.941 | 10.5121 | 48.6797
05/20/21 | 31 | 20210048.225 | 008C | 0.81 | 2672 | 7.824 | 781 | -34.542 | 10.7626 | 47.9594
05/20/21 | 32 | 20210048.226 | 010A | 0.849 | 2429 | 5.923 | 780 | -35.541 | 9.3856 | 46.0383
05/20/21 | 33 | 20210048.226 | Replicate C8 | 0.843 | 2416 | 5.851 | 782 | -35.578 | 9.4198 | 46.4396
05/20/21 | 34 | 36-UWSIF-UT Glut 1-20210048.46 | 36-UWSIF-UT Glut 1- | 0.73 | 2106 | -2.824 | 598 | -29.87 | 9.5073 | 40.5561
05/20/21 | 35 | 36-UWSIF-UT Glut 1-20210048.47 | 36-UWSIF-UT Glut 1- | 0.798 | 2298 | -2.736 | 651 | -29.903 | 9.4423 | 40.5181
05/20/21 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 39-UWSIF-UW Glut 2- | 0.835 | 2618 | 29.682 | 737 | 22.128 | 10.2421 | 43.9303
05/20/21 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 39-UWSIF-UW Glut 2- | 0.722 | 2243 | 29.67 | 636 | 22.192 | 10.2443 | 43.8964
05/20/21 | 38 | 01-UWSIF-Liver-20210048.43 | 01-UWSIF-Liver- | 0.783 | 3130 | 8.554 | 790 | -19.547 | 12.9923 | 50.1568
05/20/21 | 39 | 01-UWSIF-Liver-20210048.44 | 01-UWSIF-Liver- | 0.727 | 2885 | 8.59 | 731 | -19.51 | 12.9464 | 49.9661
05/20/21 | 40 | 20210048.227 | 010B | 0.84 | 2527 | 5.928 | 789 | -35.398 | 9.8526 | 46.9475
05/20/21 | 41 | 20210048.228 | 010C | 0.814 | 2384 | 5.892 | 748 | -35.264 | 9.595 | 45.8262
05/20/21 | 42 | 20210048.229 | 012A | 0.827 | 2830 | 8.857 | 800 | -34.735 | 11.1343 | 48.1186
05/20/21 | 43 | 20210048.230 | 012B | 0.842 | 2717 | 8.145 | 782 | -34.36 | 10.5502 | 46.4063
05/20/21 | 44 | 20210048.231 | 012C | 0.811 | 2829 | 8.507 | 787 | -34.482 | 11.3471 | 48.2655
05/20/21 | 45 | 20210048.232 | 014A | 0.836 | 2905 | 8.546 | 759 | -34.031 | 11.3395 | 45.282
05/20/21 | 46 | 20210048.233 | 014B | 0.843 | 2628 | 8.387 | 842 | -35.024 | 10.1419 | 49.7424
05/20/21 | 47 | 20210048.234 | 014C | 0.815 | 2754 | 8.203 | 749 | -34.167 | 11.0009 | 45.6241
05/20/21 | 48 | 20210048.235 | 016A | 0.817 | 2626 | 8.424 | 760 | -34.537 | 10.5088 | 46.3416
05/20/21 | 49 | 20210048.236 | 016B | 0.835 | 2846 | 8.912 | 761 | -34.07 | 11.1154 | 45.3065
05/20/21 | 50 | 20210048.237 | 016C | 0.839 | 2824 | 8.845 | 813 | -34.686 | 10.9536 | 48.2336
05/20/21 | 51 | 20210048.238 | 018A | 0.819 | 2830 | 8.899 | 795 | -34.758 | 11.2371 | 48.1816
05/20/21 | 52 | 20210048.239 | 018B | 0.841 | 2613 | 9.105 | 844 | -35.373 | 10.1029 | 49.9366
05/20/21 | 53 | 20210048.240 | 018C | 0.848 | 2728 | 9.701 | 851 | -35.227 | 10.4402 | 49.882
05/20/21 | 54 | 36-UWSIF-UT Glut 1-20210048.411 | 36-UWSIF-UT Glut 1- | 0.851 | 2452 | -2.819 | 696 | -29.886 | 9.4194 | 40.587
05/20/21 | 55 | 36-UWSIF-UT Glut 1-20210048.412 | 36-UWSIF-UT Glut 1- | 0.83 | 2390 | -2.798 | 678 | -29.966 | 9.4571 | 40.6118
05/20/21 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 39-UWSIF-UW Glut 2- | 0.779 | 2415 | 29.663 | 685 | 22.381 | 10.2097 | 43.9331
05/20/21 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 39-UWSIF-UW Glut 2- | 0.855 | 2675 | 29.64 | 758 | 22.39 | 10.2211 | 44.0974
05/20/21 | 58 | 01-UWSIF-Liver-20210048.47 | 01-UWSIF-Liver- | 0.724 | 2878 | 8.55 | 730 | -19.537 | 12.9481 | 50.0556
05/21/21 | 59 | 01-UWSIF-Liver-20210048.48 | 01-UWSIF-Liver- | 0.723 | 2858 | 8.573 | 726 | -19.47 | 12.8789 | 49.828
Comment
JOB: 2021-0048 Tray 4 Nelson
ANALYST: gh
SAMPLE COUNT: 0306
HE PSI: 1250
O2 PSI: 350
MASS 28: 36
MASS 29: 27
MASS 30: 600

## % Calc 1
Date | Row | Identifier 1 | Identifier 2 | Amount | Area All N | Area All C | Known mg N | Known mg C | Calc mg N | Calc mg C | %N | %C
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 | 33.041 | 9.888 | 0.04065 | 0.17426 | 0.041064 | 0.17662 | 9.6168 | 41.3622
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 | 80.555 | 24.785 | 0.098342 | 0.42157 | 0.097748 | 0.4185 | 9.4625 | 40.513
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 | 238.593 | 74.631 | 0.28627 | 1.2272 | 0.28629 | 1.2278 | 9.5207 | 40.833
05/17/21 | 1 | 36-UWSIF-UT Glut 1-20210048.11 | 36-UWSIF-UT Glut 1- | 0.735 | 58.147 | 17.814 |  |  | 0.071015 | 0.30531 | 9.662 | 41.5389
05/17/21 | 2 | 36-UWSIF-UT Glut 1-20210048.12 | 36-UWSIF-UT Glut 1- | 0.813 | 61.755 | 18.984 |  |  | 0.07532 | 0.32431 | 9.2644 | 39.8903
05/17/21 | 3 | 36-UWSIF-UT Glut 1-20210048.13 | 36-UWSIF-UT Glut 1- | 0.427 | 33.041 | 9.888 |  |  | 0.041064 | 0.17662 | 9.6168 | 41.3622
05/17/21 | 4 | 36-UWSIF-UT Glut 1-20210048.14 | 36-UWSIF-UT Glut 1- | 1.033 | 80.555 | 24.785 |  |  | 0.097748 | 0.4185 | 9.4625 | 40.513
05/17/21 | 5 | 36-UWSIF-UT Glut 1-20210048.15 | 36-UWSIF-UT Glut 1- | 3.007 | 238.593 | 74.631 |  |  | 0.28629 | 1.2278 | 9.5207 | 40.833
05/17/21 | 6 | 39-UWSIF-UW Glut 2-20210048.11 | 39-UWSIF-UW Glut 2- | 0.728 | 61.705 | 18.809 |  |  | 0.07526 | 0.32147 | 10.3379 | 44.1575
05/17/21 | 7 | 39-UWSIF-UW Glut 2-20210048.12 | 39-UWSIF-UW Glut 2- | 0.747 | 63.349 | 19.389 |  |  | 0.077221 | 0.33088 | 10.3375 | 44.2951
05/17/21 | 8 | 01-UWSIF-Liver-20210048.11 | 01-UWSIF-Liver- | 0.713 | 77.134 | 21.18 |  |  | 0.093667 | 0.35996 | 13.137 | 50.4859
05/17/21 | 9 | 01-UWSIF-Liver-20210048.12 | 01-UWSIF-Liver- | 0.715 | 76.95 | 21.144 |  |  | 0.093447 | 0.35938 | 13.0696 | 50.263
05/17/21 | 10 | 20210048.001 | 501 | 0.823 | 95.758 | 23.311 |  |  | 0.11589 | 0.39457 | 14.0808 | 47.9424
05/17/21 | 11 | 20210048.002 | 503 | 0.805 | 93.641 | 23.519 |  |  | 0.11336 | 0.39794 | 14.082 | 49.4339
05/17/21 | 12 | 20210048.002 | Replicate A11 | 0.834 | 98.209 | 24.536 |  |  | 0.11881 | 0.41446 | 14.2457 | 49.695
05/17/21 | 13 | 20210048.003 | 505 | 0.841 | 97.305 | 24.501 |  |  | 0.11773 | 0.41389 | 13.9989 | 49.2138
05/17/21 | 14 | 20210048.004 | 507 | 0.846 | 93.176 | 25.107 |  |  | 0.1128 | 0.42373 | 13.3339 | 50.086
05/17/21 | 15 | 20210048.005 | 509 | 0.83 | 97.887 | 23.976 |  |  | 0.11843 | 0.40536 | 14.2681 | 48.839
05/17/21 | 16 | 20210048.006 | 511 | 0.847 | 101.977 | 24.132 |  |  | 0.1233 | 0.4079 | 14.5578 | 48.1578
05/17/21 | 17 | 20210048.007 | 513 | 0.832 | 81.599 | 25.909 |  |  | 0.098994 | 0.43675 | 11.8983 | 52.4939
05/17/21 | 18 | 20210048.008 | 515 | 0.839 | 99.493 | 23.571 |  |  | 0.12034 | 0.39879 | 14.3434 | 47.5313
05/17/21 | 19 | 20210048.009 | 517 | 0.84 | 95.083 | 24.744 |  |  | 0.11508 | 0.41783 | 13.7 | 49.7421
05/17/21 | 20 | 20210048.010 | 519 | 0.826 | 93.537 | 24.012 |  |  | 0.11324 | 0.40595 | 13.7089 | 49.1462
05/17/21 | 21 | 20210048.011 | 521 | 0.832 | 92.207 | 24.523 |  |  | 0.11165 | 0.41424 | 13.4193 | 49.7891
05/17/21 | 22 | 20210048.012 | 523 | 0.841 | 96.396 | 24.498 |  |  | 0.11665 | 0.41384 | 13.87 | 49.208
05/17/21 | 23 | 20210048.013 | 525 | 0.817 | 99.859 | 22.763 |  |  | 0.12078 | 0.38567 | 14.7831 | 47.2054
05/17/21 | 24 | 20210048.014 | 527 | 0.842 | 89.851 | 24.773 |  |  | 0.10884 | 0.4183 | 12.9262 | 49.6798
05/17/21 | 25 | 20210048.015 | 529 | 0.806 | 95.866 | 23.222 |  |  | 0.11601 | 0.39312 | 14.3938 | 48.7743
05/17/21 | 26 | 20210048.016 | 531 | 0.818 | 97.31 | 23.697 |  |  | 0.11774 | 0.40083 | 14.3933 | 49.0016
05/17/21 | 27 | 20210048.017 | 533 | 0.85 | 94.315 | 25.822 |  |  | 0.11416 | 0.43534 | 13.431 | 51.2161
05/17/21 | 28 | 20210048.018 | 535 | 0.811 | 98.813 | 23.354 |  |  | 0.11953 | 0.39526 | 14.7386 | 48.7378
05/17/21 | 29 | 20210048.019 | 537 | 0.828 | 99.618 | 23.679 |  |  | 0.12049 | 0.40054 | 14.552 | 48.3745
05/17/21 | 30 | 20210048.020 | 541 | 0.844 | 95.155 | 24.744 |  |  | 0.11517 | 0.41783 | 13.6453 | 49.5063
05/17/21 | 31 | 20210048.021 | 543 | 0.826 | 99.031 | 23.578 |  |  | 0.11979 | 0.3989 | 14.5024 | 48.2931
05/17/21 | 32 | 20210048.022 | 545 | 0.831 | 91.659 | 24.65 |  |  | 0.111 | 0.41631 | 13.3568 | 50.0971
05/17/21 | 33 | 20210048.022 | Replicate C8 | 0.829 | 91.406 | 24.947 |  |  | 0.11069 | 0.42113 | 13.3526 | 50.7997
05/17/21 | 34 | 36-UWSIF-UT Glut 1-20210048.16 | 36-UWSIF-UT Glut 1- | 0.842 | 65.632 | 20.24 |  |  | 0.079945 | 0.3447 | 9.4947 | 40.9385
05/17/21 | 35 | 36-UWSIF-UT Glut 1-20210048.17 | 36-UWSIF-UT Glut 1- | 0.766 | 59.303 | 18.21 |  |  | 0.072394 | 0.31174 | 9.451 | 40.6972
05/17/21 | 36 | 39-UWSIF-UW Glut 2-20210048.13 | 39-UWSIF-UW Glut 2- | 0.788 | 66.306 | 20.446 |  |  | 0.080749 | 0.34805 | 10.2473 | 44.1684
05/17/21 | 37 | 39-UWSIF-UW Glut 2-20210048.14 | 39-UWSIF-UW Glut 2- | 0.782 | 66.016 | 20.319 |  |  | 0.080403 | 0.34598 | 10.2817 | 44.2436
05/17/21 | 38 | 01-UWSIF-Liver-20210048.13 | 01-UWSIF-Liver- | 0.776 | 82.938 | 23.103 |  |  | 0.10059 | 0.39119 | 12.9628 | 50.4109
05/17/21 | 39 | 01-UWSIF-Liver-20210048.14 | 01-UWSIF-Liver- | 0.752 | 80.794 | 22.503 |  |  | 0.098033 | 0.38145 | 13.0363 | 50.7242
05/17/21 | 40 | 20210048.023 | 547 | 0.832 | 101.04 | 24.143 |  |  | 0.12219 | 0.40807 | 14.6859 | 49.0475
05/17/21 | 41 | 20210048.024 | 549 | 0.821 | 100.213 | 22.322 |  |  | 0.1212 | 0.37851 | 14.7625 | 46.1032
05/17/21 | 42 | 20210048.025 | 551 | 0.83 | 86.493 | 25.765 |  |  | 0.10483 | 0.43441 | 12.6304 | 52.3387
05/17/21 | 43 | 20210048.026 | 553 | 0.826 | 100.084 | 23.719 |  |  | 0.12105 | 0.40119 | 14.6545 | 48.5703
05/17/21 | 44 | 20210048.027 | 555 | 0.827 | 94.682 | 24.773 |  |  | 0.1146 | 0.4183 | 13.8575 | 50.5809
05/17/21 | 45 | 20210048.028 | 557 | 0.82 | 94.627 | 24.551 |  |  | 0.11454 | 0.4147 | 13.9678 | 50.5731
05/17/21 | 46 | 20210048.029 | 559 | 0.84 | 101.213 | 24.41 |  |  | 0.12239 | 0.41241 | 14.5706 | 49.0964
05/17/21 | 47 | 20210048.030 | 561 | 0.843 | 104.067 | 23.718 |  |  | 0.1258 | 0.40117 | 14.9226 | 47.5889
05/17/21 | 48 | 20210048.031 | 563 | 0.844 | 103.666 | 24.093 |  |  | 0.12532 | 0.40726 | 14.8483 | 48.2539
05/17/21 | 49 | 20210048.032 | 565 | 0.847 | 104.652 | 24.47 |  |  | 0.1265 | 0.41338 | 14.9346 | 48.8057
05/17/21 | 50 | 20210048.033 | 567 | 0.82 | 92.819 | 24.743 |  |  | 0.11238 | 0.41782 | 13.7048 | 50.9533
05/17/21 | 51 | 20210048.034 | 569 | 0.837 | 102.371 | 23.14 |  |  | 0.12377 | 0.39179 | 14.7879 | 46.8087
05/17/21 | 52 | 20210048.035 | 571 | 0.848 | 103.384 | 23.664 |  |  | 0.12498 | 0.4003 | 14.7386 | 47.2049
05/17/21 | 53 | 20210048.036 | 573 | 0.802 | 95.32 | 23.238 |  |  | 0.11536 | 0.39338 | 14.3844 | 49.0499
05/17/21 | 54 | 20210048.037 | 575 | 0.832 | 95.758 | 24.117 |  |  | 0.11589 | 0.40765 | 13.9285 | 48.9967
05/17/21 | 55 | 20210048.038 | 577 | 0.826 | 100.427 | 23.326 |  |  | 0.12146 | 0.39481 | 14.704 | 47.7977
05/17/21 | 56 | 20210048.039 | 579 | 0.836 | 73.574 | 27.637 |  |  | 0.08942 | 0.46481 | 10.6961 | 55.5989
05/17/21 | 57 | 20210048.040 | 581 | 0.836 | 96.305 | 24.2 |  |  | 0.11654 | 0.409 | 13.9399 | 48.9235
05/17/21 | 58 | 20210048.041 | 583 | 0.836 | 81.12 | 26.978 |  |  | 0.098422 | 0.45411 | 11.773 | 54.319
05/17/21 | 59 | 20210048.042 | 585 | 0.839 | 100.274 | 23.759 |  |  | 0.12127 | 0.40184 | 14.4545 | 47.8951
05/17/21 | 60 | 20210048.042 | Replicate E11 | 0.832 | 99.199 | 23.507 |  |  | 0.11999 | 0.39775 | 14.4219 | 47.8063
05/17/21 | 61 | 20210048.043 | 587 | 0.805 | 58.333 | 27.147 |  |  | 0.071237 | 0.45685 | 8.8494 | 56.7517
05/18/21 | 62 | 20210048.044 | 589 | 0.818 | 91.362 | 24.483 |  |  | 0.11064 | 0.4136 | 13.5258 | 50.5618
05/18/21 | 63 | 20210048.045 | 591 | 0.812 | 86.65 | 23.835 |  |  | 0.10502 | 0.40307 | 12.9334 | 49.6396
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.18 | 36-UWSIF-UT Glut 1- | 0.783 | 60.166 | 18.628 |  |  | 0.073424 | 0.31853 | 9.3773 | 40.6804
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.19 | 36-UWSIF-UT Glut 1- | 0.794 | 60.995 | 18.876 |  |  | 0.074413 | 0.32255 | 9.3719 | 40.624
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.15 | 39-UWSIF-UW Glut 2- | 0.735 | 61.22 | 18.943 |  |  | 0.074681 | 0.32364 | 10.1607 | 44.033
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.16 | 39-UWSIF-UW Glut 2- | 0.722 | 60.468 | 18.705 |  |  | 0.073784 | 0.31978 | 10.2194 | 44.2906
05/18/21 | 68 | 01-UWSIF-Liver-20210048.15 | 01-UWSIF-Liver- | 0.728 | 77.452 | 21.669 |  |  | 0.094046 | 0.3679 | 12.9184 | 50.5363
05/18/21 | 69 | 01-UWSIF-Liver-20210048.16 | 01-UWSIF-Liver- | 0.738 | 78.616 | 22.028 |  |  | 0.095435 | 0.37373 | 12.9316 | 50.6414
05/18/21 | 70 | 20210048.046 | 593 | 0.822 | 100.105 | 23.697 |  |  | 0.12107 | 0.40083 | 14.7289 | 48.7632
05/18/21 | 71 | 20210048.047 | 595 | 0.826 | 95.675 | 24.287 |  |  | 0.11579 | 0.41041 | 14.0177 | 49.6868
05/18/21 | 72 | 20210048.048 | 597 | 0.849 | 103.14 | 24.121 |  |  | 0.12469 | 0.40772 | 14.6869 | 48.0233
05/18/21 | 73 | 20210048.049 | 599 | 0.807 | 95.849 | 23.128 |  |  | 0.11599 | 0.39159 | 14.3735 | 48.5247
05/18/21 | 74 | 20210048.050 | 601 | 0.833 | 96.427 | 24.185 |  |  | 0.11668 | 0.40876 | 14.0076 | 49.0704
05/18/21 | 75 | 20210048.051 | 603 | 0.822 | 94.4 | 23.883 |  |  | 0.11427 | 0.40385 | 13.9009 | 49.1306
05/18/21 | 76 | 20210048.052 | 605 | 0.85 | 101.637 | 24.269 |  |  | 0.1229 | 0.41012 | 14.4587 | 48.2495
05/18/21 | 77 | 20210048.053 | 607 | 0.818 | 97.931 | 22.994 |  |  | 0.11848 | 0.38942 | 14.4838 | 47.6062
05/18/21 | 78 | 20210048.054 | 609 | 0.833 | 87.401 | 26.355 |  |  | 0.10592 | 0.44399 | 12.7149 | 53.3003
05/18/21 | 79 | 20210048.055 | 611 | 0.826 | 77.888 | 26.17 |  |  | 0.094566 | 0.44099 | 11.4487 | 53.3883
05/18/21 | 80 | 20210048.056 | 613 | 0.801 | 97.039 | 22.682 |  |  | 0.11741 | 0.38435 | 14.6584 | 47.9841
05/18/21 | 81 | 20210048.057 | 615 | 0.801 | 82.52 | 24.168 |  |  | 0.10009 | 0.40848 | 12.4959 | 50.9964
05/18/21 | 82 | 20210048.058 | 617 | 0.848 | 93.296 | 25.038 |  |  | 0.11295 | 0.42261 | 13.3194 | 49.8357
05/18/21 | 83 | 20210048.059 | 619 | 0.839 | 97.526 | 23.421 |  |  | 0.11799 | 0.39635 | 14.0637 | 47.241
05/18/21 | 84 | 20210048.060 | 623 | 0.817 | 98.305 | 23.432 |  |  | 0.11892 | 0.39653 | 14.5562 | 48.5349
05/18/21 | 85 | 20210048.061 | 625 | 0.813 | 96.6 | 23.191 |  |  | 0.11689 | 0.39262 | 14.3776 | 48.2924
05/18/21 | 86 | 20210048.062 | 627 | 0.816 | 91.351 | 23.34 |  |  | 0.11063 | 0.39504 | 13.5573 | 48.4113
05/18/21 | 87 | 20210048.063 | 629 | 0.843 | 101.212 | 23.692 |  |  | 0.12239 | 0.40075 | 14.5186 | 47.5388
05/18/21 | 88 | 20210048.064 | 631 | 0.825 | 87.694 | 24.495 |  |  | 0.10626 | 0.41379 | 12.8806 | 50.1564
05/18/21 | 89 | 20210048.065 | 635 | 0.821 | 97.804 | 23.852 |  |  | 0.11833 | 0.40335 | 14.4124 | 49.1291
05/18/21 | 90 | 20210048.066 | 637 | 0.836 | 98.554 | 24.146 |  |  | 0.11922 | 0.40812 | 14.2609 | 48.8186
05/18/21 | 91 | 20210048.067 | 639 | 0.83 | 95.028 | 23.797 |  |  | 0.11501 | 0.40246 | 13.8572 | 48.4888
05/18/21 | 92 | 20210048.068 | 641 | 0.848 | 102.497 | 23.639 |  |  | 0.12392 | 0.39989 | 14.6138 | 47.157
05/18/21 | 93 | 20210048.068 | Replicate H8 | 0.828 | 101.332 | 23.312 |  |  | 0.12254 | 0.39458 | 14.7989 | 47.6548
05/18/21 | 94 | 36-UWSIF-UT Glut 1-20210048.111 | 36-UWSIF-UT Glut 1- | 0.754 | 57.376 | 17.744 |  |  | 0.070096 | 0.30417 | 9.2965 | 40.3414
05/18/21 | 95 | 36-UWSIF-UT Glut 1-20210048.112 | 36-UWSIF-UT Glut 1- | 0.787 | 59.898 | 18.574 |  |  | 0.073104 | 0.31765 | 9.289 | 40.3623
05/18/21 | 96 | 39-UWSIF-UW Glut 2-20210048.17 | 39-UWSIF-UW Glut 2- | 0.703 | 58.195 | 17.968 |  |  | 0.071073 | 0.30781 | 10.1099 | 43.7854
05/18/21 | 97 | 39-UWSIF-UW Glut 2-20210048.18 | 39-UWSIF-UW Glut 2- | 0.83 | 68.489 | 21.316 |  |  | 0.083353 | 0.36217 | 10.0426 | 43.6353
05/18/21 | 98 | 01-UWSIF-Liver-20210048.17 | 01-UWSIF-Liver- | 0.753 | 79.08 | 22.148 |  |  | 0.095988 | 0.37568 | 12.7475 | 49.8914
05/18/21 | 99 | 01-UWSIF-Liver-20210048.18 | 01-UWSIF-Liver- | 0.735 | 77.237 | 21.645 |  |  | 0.09379 | 0.36751 | 12.7605 | 50.002
Comment
JOB: 2021-0048 Tray 1 Nelson
ANALYST: GH
SAMPLE COUNT: 0009
HE PSI: 2200
O2 PSI: 500
MASS 28: 42
MASS 29: 32
MASS 30: 748

## % Calc 2
Date | Row | Identifier 1 | Identifier 2 | Amount | Area All N | Area All C | Known mg N | Known mg C | Calc mg N | Calc mg C | %N | %C
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 | 35.554 | 10.822 | 0.044934 | 0.19262 | 0.044602 | 0.19127 | 9.4495 | 40.5235
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 | 55.169 | 17.1 | 0.068925 | 0.29546 | 0.069415 | 0.29726 | 9.5877 | 41.0583
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 | 115.77 | 36.572 | 0.14613 | 0.62643 | 0.14608 | 0.62601 | 9.5163 | 40.7823
05/18/21 | 1 | 36-UWSIF-UT Glut 1-20210048.21 | 36-UWSIF-UT Glut 1- | 0.769 | 58.412 | 18.189 |  |  | 0.073517 | 0.31565 | 9.5601 | 41.0465
05/18/21 | 2 | 36-UWSIF-UT Glut 1-20210048.22 | 36-UWSIF-UT Glut 1- | 0.824 | 61.881 | 19.239 |  |  | 0.077905 | 0.33338 | 9.4545 | 40.4581
05/18/21 | 3 | 36-UWSIF-UT Glut 1-20210048.23 | 36-UWSIF-UT Glut 1- | 0.472 | 35.554 | 10.822 |  |  | 0.044602 | 0.19127 | 9.4495 | 40.5235
05/18/21 | 4 | 36-UWSIF-UT Glut 1-20210048.24 | 36-UWSIF-UT Glut 1- | 0.724 | 55.169 | 17.1 |  |  | 0.069415 | 0.29726 | 9.5877 | 41.0583
05/18/21 | 5 | 36-UWSIF-UT Glut 1-20210048.25 | 36-UWSIF-UT Glut 1- | 1.535 | 115.77 | 36.572 |  |  | 0.14608 | 0.62601 | 9.5163 | 40.7823
05/18/21 | 6 | 39-UWSIF-UW Glut 2-20210048.21 | 39-UWSIF-UW Glut 2- | 0.742 | 60.675 | 18.924 |  |  | 0.07638 | 0.32806 | 10.2938 | 44.2125
05/18/21 | 7 | 39-UWSIF-UW Glut 2-20210048.22 | 39-UWSIF-UW Glut 2- | 0.736 | 57.082 | 17.769 |  |  | 0.071835 | 0.30856 | 9.7602 | 41.9235
05/18/21 | 8 | 01-UWSIF-Liver-20210048.21 | 01-UWSIF-Liver- | 0.723 | 74.933 | 21.062 |  |  | 0.094416 | 0.36415 | 13.059 | 50.3669
05/18/21 | 9 | 01-UWSIF-Liver-20210048.22 | 01-UWSIF-Liver- | 0.717 | 74.418 | 20.924 |  |  | 0.093765 | 0.36182 | 13.0774 | 50.4634
05/18/21 | 10 | 20210048.069 | 643 | 0.806 | 94.551 | 23.019 |  |  | 0.11923 | 0.39719 | 14.7932 | 49.2795
05/18/21 | 11 | 20210048.070 | 645 | 0.814 | 97.932 | 23.096 |  |  | 0.12351 | 0.39849 | 15.1732 | 48.9549
05/18/21 | 12 | 20210048.070 | Replicate A11 | 0.84 | 100.845 | 23.834 |  |  | 0.12719 | 0.41095 | 15.1423 | 48.9229
05/18/21 | 13 | 20210048.071 | 651 | 0.824 | 98.653 | 22.852 |  |  | 0.12442 | 0.39437 | 15.0998 | 47.8608
05/18/21 | 14 | 20210048.072 | 653 | 0.845 | 98.287 | 23.982 |  |  | 0.12396 | 0.41345 | 14.6697 | 48.9291
05/18/21 | 15 | 20210048.073 | 655 | 0.84 | 98.009 | 24.178 |  |  | 0.12361 | 0.41676 | 14.7152 | 49.6143
05/18/21 | 16 | 20210048.074 | 657 | 0.837 | 97.155 | 23.828 |  |  | 0.12253 | 0.41085 | 14.6388 | 49.0862
05/18/21 | 17 | 20210048.075 | 659 | 0.838 | 92.484 | 24.696 |  |  | 0.11662 | 0.42551 | 13.9163 | 50.7763
05/18/21 | 18 | 20210048.076 | 661 | 0.846 | 96.962 | 23.514 |  |  | 0.12228 | 0.40555 | 14.4542 | 47.9373
05/18/21 | 19 | 20210048.077 | 663 | 0.818 | 97.948 | 22.875 |  |  | 0.12353 | 0.39476 | 15.1015 | 48.2594
05/18/21 | 20 | 20210048.078 | 665 | 0.824 | 74.083 | 24.012 |  |  | 0.093341 | 0.41396 | 11.3278 | 50.2376
05/18/21 | 21 | 20210048.079 | 667 | 0.802 | 94.077 | 22.192 |  |  | 0.11863 | 0.38323 | 14.7922 | 47.7844
05/18/21 | 22 | 20210048.080 | 669 | 0.836 | 100.005 | 23.236 |  |  | 0.12613 | 0.40086 | 15.0876 | 47.9493
05/18/21 | 23 | 20210048.081 | 671 | 0.828 | 99.335 | 22.989 |  |  | 0.12528 | 0.39669 | 15.131 | 47.909
05/18/21 | 24 | 20210048.082 | 673 | 0.816 | 96.149 | 22.484 |  |  | 0.12125 | 0.38816 | 14.8596 | 47.5687
05/18/21 | 25 | 20210048.083 | 703 | 0.832 | 84.658 | 24.568 |  |  | 0.10672 | 0.42334 | 12.8267 | 50.8828
05/18/21 | 26 | 20210048.084 | 705 | 0.801 | 84.919 | 23.318 |  |  | 0.10705 | 0.40224 | 13.3644 | 50.2173
05/18/21 | 27 | 20210048.085 | 709 | 0.835 | 98.832 | 23.324 |  |  | 0.12465 | 0.40234 | 14.928 | 48.1847
05/18/21 | 28 | 20210048.086 | 713 | 0.829 | 86.526 | 24.14 |  |  | 0.10908 | 0.41612 | 13.1582 | 50.1952
05/18/21 | 29 | 20210048.087 | 715 | 0.81 | 75.43 | 25.503 |  |  | 0.095045 | 0.43913 | 11.7339 | 54.2136
05/18/21 | 30 | 20210048.088 | 719 | 0.829 | 94.661 | 23.243 |  |  | 0.11937 | 0.40097 | 14.3995 | 48.3685
05/18/21 | 31 | 20210048.089 | 721 | 0.814 | 92.591 | 22.959 |  |  | 0.11675 | 0.39618 | 14.3432 | 48.6707
05/18/21 | 32 | 20210048.090 | 723 | 0.825 | 97.687 | 22.78 |  |  | 0.1232 | 0.39316 | 14.9333 | 47.6555
05/18/21 | 33 | 20210048.090 | Replicate C8 | 0.81 | 97.101 | 22.497 |  |  | 0.12246 | 0.38838 | 15.1184 | 47.9481
05/18/21 | 34 | 36-UWSIF-UT Glut 1-20210048.26 | 36-UWSIF-UT Glut 1- | 0.731 | 54.821 | 16.754 |  |  | 0.068975 | 0.29142 | 9.4356 | 39.866
05/18/21 | 35 | 36-UWSIF-UT Glut 1-20210048.27 | 36-UWSIF-UT Glut 1- | 0.829 | 62.107 | 19.131 |  |  | 0.078191 | 0.33155 | 9.432 | 39.9942
05/18/21 | 36 | 39-UWSIF-UW Glut 2-20210048.23 | 39-UWSIF-UW Glut 2- | 0.769 | 62.828 | 19.379 |  |  | 0.079103 | 0.33574 | 10.2865 | 43.6591
05/18/21 | 37 | 39-UWSIF-UW Glut 2-20210048.24 | 39-UWSIF-UW Glut 2- | 0.841 | 68.061 | 21.105 |  |  | 0.085723 | 0.36488 | 10.193 | 43.3863
05/18/21 | 38 | 01-UWSIF-Liver-20210048.23 | 01-UWSIF-Liver- | 0.736 | 75.675 | 21.205 |  |  | 0.095355 | 0.36657 | 12.9558 | 49.8053
05/18/21 | 39 | 01-UWSIF-Liver-20210048.24 | 01-UWSIF-Liver- | 0.731 | 75.322 | 21.061 |  |  | 0.094908 | 0.36414 | 12.9834 | 49.8134
05/18/21 | 40 | 20210048.091 | 725 | 0.827 | 91.21 | 23.97 |  |  | 0.11501 | 0.41325 | 13.9065 | 49.9696
05/18/21 | 41 | 20210048.092 | 727 | 0.807 | 93.758 | 22.634 |  |  | 0.11823 | 0.39069 | 14.6505 | 48.413
05/18/21 | 42 | 20210048.093 | 729 | 0.85 | 96.422 | 24.238 |  |  | 0.1216 | 0.41777 | 14.3059 | 49.1498
05/18/21 | 43 | 20210048.094 | 731 | 0.823 | 85.401 | 24.677 |  |  | 0.10766 | 0.42518 | 13.0812 | 51.6628
05/18/21 | 44 | 20210048.095 | 733 | 0.837 | 100.22 | 23.118 |  |  | 0.1264 | 0.39886 | 15.1021 | 47.654
05/18/21 | 45 | 20210048.096 | 735 | 0.85 | 95.703 | 24.231 |  |  | 0.12069 | 0.41765 | 14.1989 | 49.1359
05/18/21 | 46 | 20210048.097 | 739 | 0.826 | 95.119 | 22.565 |  |  | 0.11995 | 0.38953 | 14.522 | 47.1583
05/18/21 | 47 | 20210048.098 | 741 | 0.825 | 93.373 | 23.09 |  |  | 0.11774 | 0.39839 | 14.2719 | 48.2899
05/18/21 | 48 | 20210048.099 | 747 | 0.849 | 96.799 | 24.229 |  |  | 0.12208 | 0.41762 | 14.3789 | 49.1898
05/18/21 | 49 | 20210048.100 | 847 | 0.815 | 57.407 | 26.362 |  |  | 0.072246 | 0.45363 | 8.8645 | 55.6604
05/18/21 | 50 | 20210048.101 | 849 | 0.833 | 82.414 | 25.104 |  |  | 0.10388 | 0.43239 | 12.4706 | 51.908
05/18/21 | 51 | 20210048.102 | 851 | 0.82 | 93.634 | 22.111 |  |  | 0.11807 | 0.38186 | 14.3991 | 46.5687
05/18/21 | 52 | 20210048.103 | 853 | 0.833 | 88.986 | 24.043 |  |  | 0.11219 | 0.41448 | 13.4686 | 49.7576
05/18/21 | 53 | 20210048.104 | 855 | 0.825 | 88.561 | 23.652 |  |  | 0.11166 | 0.40788 | 13.534 | 49.44
05/18/21 | 54 | 20210048.105 | 857 | 0.848 | 87.561 | 25.126 |  |  | 0.11039 | 0.43277 | 13.0178 | 51.0336
05/18/21 | 55 | 20210048.106 | 859 | 0.848 | 90.663 | 24.336 |  |  | 0.11431 | 0.41943 | 13.4805 | 49.4608
05/18/21 | 56 | 20210048.107 | 861 | 0.806 | 88.855 | 22.973 |  |  | 0.11203 | 0.39642 | 13.8992 | 49.1831
05/18/21 | 57 | 20210048.108 | 863 | 0.812 | 93.265 | 22.376 |  |  | 0.11761 | 0.38634 | 14.4835 | 47.5784
05/18/21 | 58 | 20210048.109 | 865 | 0.827 | 90.09 | 23.787 |  |  | 0.11359 | 0.41016 | 13.7352 | 49.596
05/18/21 | 59 | 20210048.110 | 867 | 0.816 | 53.539 | 28.367 |  |  | 0.067353 | 0.48748 | 8.254 | 59.7406
05/18/21 | 60 | 20210048.110 | Replicate E11 | 0.845 | 65.207 | 28.306 |  |  | 0.082113 | 0.48645 | 9.7175 | 57.5684
05/18/21 | 61 | 20210048.111 | 869 | 0.834 |  |  |  |  | -0.000374 | 0.008563 | -0.044844 | 1.0267
05/18/21 | 62 | 20210048.112 | 871 | 0.825 | 94.137 | 22.881 |  |  | 0.11871 | 0.39486 | 14.389 | 47.8622
05/18/21 | 63 | 20210048.113 | 873 | 0.839 | 97.381 | 23.092 |  |  | 0.12281 | 0.39843 | 14.638 | 47.4881
05/18/21 | 64 | 36-UWSIF-UT Glut 1-20210048.28 | 36-UWSIF-UT Glut 1- | 0.827 | 61.137 | 19.265 |  |  | 0.076964 | 0.33381 | 9.3064 | 40.3644
05/18/21 | 65 | 36-UWSIF-UT Glut 1-20210048.29 | 36-UWSIF-UT Glut 1- | 0.791 | 58.525 | 18.334 |  |  | 0.07366 | 0.3181 | 9.3123 | 40.2144
05/18/21 | 66 | 39-UWSIF-UW Glut 2-20210048.25 | 39-UWSIF-UW Glut 2- | 0.716 | 57.827 | 18.163 |  |  | 0.072777 | 0.31521 | 10.1644 | 44.0236
05/18/21 | 67 | 39-UWSIF-UW Glut 2-20210048.26 | 39-UWSIF-UW Glut 2- | 0.798 | 64.379 | 20.353 |  |  | 0.081065 | 0.35218 | 10.1586 | 44.1332
05/18/21 | 68 | 01-UWSIF-Liver-20210048.25 | 01-UWSIF-Liver- | 0.741 | 75.521 | 21.509 |  |  | 0.09516 | 0.3717 | 12.8421 | 50.1619
05/18/21 | 69 | 01-UWSIF-Liver-20210048.26 | 01-UWSIF-Liver- | 0.767 | 78.608 | 22.318 |  |  | 0.099065 | 0.38536 | 12.9159 | 50.2422
05/18/21 | 70 | 20210048.114 | 875 | 0.826 | 95.143 | 23.044 |  |  | 0.11998 | 0.39761 | 14.5257 | 48.1374
05/18/21 | 71 | 20210048.115 | 877 | 0.832 | 76.749 | 25.901 |  |  | 0.096713 | 0.44585 | 11.6242 | 53.5877
05/18/21 | 72 | 20210048.116 | 879 | 0.824 | 96.447 | 22.671 |  |  | 0.12163 | 0.39132 | 14.7611 | 47.49
05/18/21 | 73 | 20210048.117 | 881 | 0.845 | 98.717 | 23.321 |  |  | 0.1245 | 0.40229 | 14.7341 | 47.6085
05/18/21 | 74 | 20210048.118 | 883 | 0.841 | 95.416 | 23.041 |  |  | 0.12033 | 0.39756 | 14.3076 | 47.2728
05/18/21 | 75 | 20210048.119 | 885 | 0.807 | 78.285 | 24.626 |  |  | 0.098657 | 0.42432 | 12.2251 | 52.5804
05/18/21 | 76 | 20210048.120 | 887 | 0.849 | 99.87 | 23.317 |  |  | 0.12596 | 0.40222 | 14.8365 | 47.3762
05/18/21 | 77 | 20210048.121 | 889 | 0.811 | 94.151 | 22.201 |  |  | 0.11873 | 0.38338 | 14.6396 | 47.2728
05/18/21 | 78 | 20210048.122 | 891 | 0.801 | 87.8 | 22.23 |  |  | 0.11069 | 0.38387 | 13.8194 | 47.9241
05/18/21 | 79 | 20210048.123 | 893 | 0.813 | 89.194 | 22.034 |  |  | 0.11246 | 0.38056 | 13.8323 | 46.8097
05/18/21 | 80 | 20210048.124 | 895 | 0.812 | 88.147 | 23.726 |  |  | 0.11113 | 0.40913 | 13.6862 | 50.3854
05/18/21 | 81 | 20210048.125 | 897 | 0.841 | 93.643 | 23.034 |  |  | 0.11808 | 0.39745 | 14.041 | 47.2587
05/19/21 | 82 | 20210048.126 | 899 | 0.81 | 84.803 | 23.132 |  |  | 0.1069 | 0.3991 | 13.1978 | 49.2717
05/19/21 | 83 | 20210048.127 | 901 | 0.811 | 74.392 | 24.729 |  |  | 0.093732 | 0.42606 | 11.5576 | 52.5355
05/19/21 | 84 | 20210048.128 | 903 | 0.832 | 87.02 | 24.384 |  |  | 0.10971 | 0.42024 | 13.1859 | 50.5094
05/19/21 | 85 | 20210048.129 | 905 | 0.8 | 75.394 | 24.436 |  |  | 0.094999 | 0.42112 | 11.8749 | 52.6395
05/19/21 | 86 | 20210048.130 | 907 | 0.846 | 83.525 | 25.851 |  |  | 0.10529 | 0.44501 | 12.4451 | 52.6011
05/19/21 | 87 | 20210048.131 | 909 | 0.807 | 91.789 | 22.548 |  |  | 0.11574 | 0.38924 | 14.3419 | 48.2331
05/19/21 | 88 | 20210048.132 | 911 | 0.844 | 95.944 | 23.529 |  |  | 0.121 | 0.4058 | 14.3359 | 48.0809
05/19/21 | 89 | 20210048.133 | 913 | 0.815 | 91.818 | 23.093 |  |  | 0.11578 | 0.39844 | 14.2056 | 48.8886
05/19/21 | 90 | 20210048.134 | 915 | 0.838 | 90.825 | 24.158 |  |  | 0.11452 | 0.41642 | 13.6658 | 49.6924
05/19/21 | 91 | 20210048.135 | 917 | 0.822 | 95.613 | 22.592 |  |  | 0.12058 | 0.38998 | 14.6687 | 47.4433
05/19/21 | 92 | 20210048.136 | 919 | 0.817 | 87.586 | 23.244 |  |  | 0.11042 | 0.40099 | 13.5156 | 49.081
05/19/21 | 93 | 20210048.136 | Replicate H8 | 0.817 | 88.4 | 23.289 |  |  | 0.11145 | 0.40175 | 13.6416 | 49.174
05/19/21 | 94 | 36-UWSIF-UT Glut 1-20210048.211 | 36-UWSIF-UT Glut 1- | 0.744 | 54.902 | 17.241 |  |  | 0.069077 | 0.29964 | 9.2845 | 40.2746
05/19/21 | 95 | 36-UWSIF-UT Glut 1-20210048.212 | 36-UWSIF-UT Glut 1- | 0.755 | 55.643 | 17.473 |  |  | 0.070014 | 0.30356 | 9.2734 | 40.2066
05/19/21 | 96 | 39-UWSIF-UW Glut 2-20210048.27 | 39-UWSIF-UW Glut 2- | 0.73 | 58.816 | 18.369 |  |  | 0.074028 | 0.31869 | 10.1409 | 43.6557
05/19/21 | 97 | 39-UWSIF-UW Glut 2-20210048.28 | 39-UWSIF-UW Glut 2- | 0.768 | 61.105 | 19.192 |  |  | 0.076924 | 0.33258 | 10.0161 | 43.3049
05/19/21 | 98 | 01-UWSIF-Liver-20210048.27 | 01-UWSIF-Liver- | 0.825 | 83.884 | 23.841 |  |  | 0.10574 | 0.41107 | 12.8169 | 49.8267
05/19/21 | 99 | 01-UWSIF-Liver-20210048.28 | 01-UWSIF-Liver- | 0.736 | 74.808 | 21.169 |  |  | 0.094258 | 0.36596 | 12.8068 | 49.7227
Comment
JOB: 2021-0048 Tray 2 Nelson
ANALYST: sme
SAMPLE COUNT: 0108
HE PSI: 1900
O2 PSI: 500
MASS 28:
MASS 29:
MASS 30:

## % Calc 3
Date | Row | Identifier 1 | Identifier 2 | Amount | Area All N | Area All C | Known mg N | Known mg C | Calc mg N | Calc mg C | %N | %C
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 | 31.047 | 9.514 | 0.040555 | 0.17385 | 0.040567 | 0.17484 | 9.5228 | 41.0425
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 | 54.634 | 17.058 | 0.071305 | 0.30567 | 0.071301 | 0.3043 | 9.5195 | 40.627
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 | 115.219 | 36.875 | 0.15023 | 0.64398 | 0.15024 | 0.64436 | 9.5211 | 40.8337
05/19/21 | 1 | 36-UWSIF-UT Glut 1-20210048.31 | 36-UWSIF-UT Glut 1- | 0.795 | 58.414 | 18.147 |  |  | 0.076226 | 0.32298 | 9.5882 | 40.6269
05/19/21 | 2 | 36-UWSIF-UT Glut 1-20210048.32 | 36-UWSIF-UT Glut 1- | 0.72 | 52.42 | 16.372 |  |  | 0.068416 | 0.29252 | 9.5023 | 40.6284
05/19/21 | 3 | 36-UWSIF-UT Glut 1-20210048.33 | 36-UWSIF-UT Glut 1- | 0.426 | 31.047 | 9.514 |  |  | 0.040567 | 0.17484 | 9.5228 | 41.0425
05/19/21 | 4 | 36-UWSIF-UT Glut 1-20210048.34 | 36-UWSIF-UT Glut 1- | 0.749 | 54.634 | 17.058 |  |  | 0.071301 | 0.3043 | 9.5195 | 40.627
05/19/21 | 5 | 36-UWSIF-UT Glut 1-20210048.35 | 36-UWSIF-UT Glut 1- | 1.578 | 115.219 | 36.875 |  |  | 0.15024 | 0.64436 | 9.5211 | 40.8337
05/19/21 | 6 | 39-UWSIF-UW Glut 2-20210048.31 | 39-UWSIF-UW Glut 2- | 0.8 | 63.315 | 19.899 |  |  | 0.082612 | 0.35305 | 10.3266 | 44.131
05/19/21 | 7 | 39-UWSIF-UW Glut 2-20210048.32 | 39-UWSIF-UW Glut 2- | 0.829 | 65.705 | 20.642 |  |  | 0.085727 | 0.3658 | 10.341 | 44.1252
05/19/21 | 8 | 01-UWSIF-Liver-20210048.31 | 01-UWSIF-Liver- | 0.811 | 81.455 | 23.115 |  |  | 0.10625 | 0.40823 | 13.101 | 50.3372
05/19/21 | 9 | 01-UWSIF-Liver-20210048.32 | 01-UWSIF-Liver- | 0.719 | 71.856 | 20.342 |  |  | 0.093741 | 0.36065 | 13.0377 | 50.1599
05/19/21 | 10 | 20210048.137 | 921 | 0.846 | 82.883 | 19.459 |  |  | 0.10811 | 0.3455 | 12.7789 | 40.8389
05/19/21 | 11 | 20210048.138 | 923 | 0.842 | 88.957 | 24.239 |  |  | 0.11602 | 0.42752 | 13.7796 | 50.7746
05/19/21 | 12 | 20210048.138 | Replicate A11 | 0.829 | 81.646 | 24.341 |  |  | 0.1065 | 0.42927 | 12.8465 | 51.782
05/19/21 | 13 | 20210048.139 | 925 | 0.801 | 76.691 | 23.326 |  |  | 0.10004 | 0.41186 | 12.4896 | 51.4176
05/19/21 | 14 | 20210048.140 | 927 | 0.839 | 86.472 | 24.096 |  |  | 0.11279 | 0.42507 | 13.4429 | 50.6637
05/19/21 | 15 | 20210048.141 | 929 | 0.843 | 84.362 | 24.69 |  |  | 0.11004 | 0.43526 | 13.053 | 51.6324
05/19/21 | 16 | 20210048.142 | 931 | 0.812 | 89.154 | 21.768 |  |  | 0.11628 | 0.38512 | 14.3203 | 47.4286
05/19/21 | 17 | 20210048.143 | 933 | 0.83 | 90.494 | 23.173 |  |  | 0.11803 | 0.40923 | 14.2201 | 49.3048
05/19/21 | 18 | 20210048.144 | 935 | 0.825 | 92.87 | 22.64 |  |  | 0.12112 | 0.40008 | 14.6815 | 48.495
05/19/21 | 19 | 20210048.145 | 937 | 0.84 | 62.264 | 27.468 |  |  | 0.081243 | 0.48293 | 9.6718 | 57.4919
05/19/21 | 20 | 20210048.146 | 939 | 0.832 | 85.611 | 23.92 |  |  | 0.11166 | 0.42205 | 13.4212 | 50.7269
05/19/21 | 21 | 20210048.147 | 941 | 0.814 | 76.204 | 24.858 |  |  | 0.099407 | 0.43814 | 12.2121 | 53.8261
05/19/21 | 22 | 20210048.148 | 943 | 0.805 | 89.075 | 22.371 |  |  | 0.11618 | 0.39547 | 14.432 | 49.1264
05/19/21 | 23 | 20210048.149 | 945 | 0.84 | 83.022 | 25.068 |  |  | 0.10829 | 0.44175 | 12.8917 | 52.589
05/19/21 | 24 | 20210048.150 | 947 | 0.829 | 84.502 | 24.047 |  |  | 0.11022 | 0.42423 | 13.2954 | 51.1734
05/19/21 | 25 | 20210048.151 | 949 | 0.826 | 89.587 | 23.189 |  |  | 0.11684 | 0.4095 | 14.1459 | 49.5768
05/19/21 | 26 | 20210048.152 | 951 | 0.843 | 95.196 | 22.989 |  |  | 0.12415 | 0.40607 | 14.7276 | 48.1699
05/19/21 | 27 | 20210048.153 | 953 | 0.827 | 84.146 | 23.763 |  |  | 0.10976 | 0.41935 | 13.2715 | 50.7079
05/19/21 | 28 | 20210048.154 | 955 | 0.803 | 86.453 | 22.407 |  |  | 0.11276 | 0.39609 | 14.0425 | 49.3257
05/19/21 | 29 | 20210048.155 | 957 | 0.846 | 97.266 | 22.724 |  |  | 0.12685 | 0.40152 | 14.9942 | 47.4616
05/19/21 | 30 | 20210048.156 | 959 | 0.846 | 97.591 | 23.157 |  |  | 0.12727 | 0.40896 | 15.0442 | 48.3398
05/19/21 | 31 | 20210048.157 | 961 | 0.824 | 80.671 | 24.058 |  |  | 0.10523 | 0.42442 | 12.7703 | 51.5068
05/19/21 | 32 | 20210048.158 | 963 | 0.827 | 91.917 | 22.873 |  |  | 0.11988 | 0.40408 | 14.4959 | 48.8611
05/19/21 | 33 | 20210048.158 | Replicate C8 | 0.821 | 90.514 | 22.274 |  |  | 0.11805 | 0.3938 | 14.3791 | 47.9662
05/19/21 | 34 | 36-UWSIF-UT Glut 1-20210048.36 | 36-UWSIF-UT Glut 1- | 0.801 | 58.231 | 18.159 |  |  | 0.075988 | 0.32319 | 9.4866 | 40.3482
05/19/21 | 35 | 36-UWSIF-UT Glut 1-20210048.37 | 36-UWSIF-UT Glut 1- | 0.786 | 57.136 | 17.878 |  |  | 0.074561 | 0.31837 | 9.4862 | 40.5048
05/19/21 | 36 | 39-UWSIF-UW Glut 2-20210048.33 | 39-UWSIF-UW Glut 2- | 0.779 | 61.356 | 19.145 |  |  | 0.08006 | 0.34011 | 10.2773 | 43.6597
05/19/21 | 37 | 39-UWSIF-UW Glut 2-20210048.34 | 39-UWSIF-UW Glut 2- | 0.726 | 57.559 | 17.836 |  |  | 0.075112 | 0.31765 | 10.3461 | 43.753
05/19/21 | 38 | 01-UWSIF-Liver-20210048.33 | 01-UWSIF-Liver- | 0.75 | 74.783 | 21.067 |  |  | 0.097555 | 0.37309 | 13.0074 | 49.7454
05/19/21 | 39 | 01-UWSIF-Liver-20210048.34 | 01-UWSIF-Liver- | 0.79 | 78.722 | 22.212 |  |  | 0.10269 | 0.39274 | 12.9985 | 49.7138
05/19/21 | 40 | 20210048.159 | 965 | 0.825 | 90.134 | 23.084 |  |  | 0.11756 | 0.4077 | 14.2494 | 49.4185
05/19/21 | 41 | 20210048.160 | 967 | 0.807 | 80.58 | 24.277 |  |  | 0.10511 | 0.42817 | 13.0246 | 53.0575
05/19/21 | 42 | 20210048.161 | 969 | 0.809 | 86.601 | 22.569 |  |  | 0.11295 | 0.39887 | 13.9622 | 49.3035
05/19/21 | 43 | 20210048.162 | 971 | 0.808 | 80.398 | 23.59 |  |  | 0.10487 | 0.41639 | 12.9792 | 51.5328
05/19/21 | 44 | 20210048.163 | 973 | 0.826 | 79.066 | 24.833 |  |  | 0.10314 | 0.43772 | 12.4862 | 52.9922
05/19/21 | 45 | 20210048.164 | 975 | 0.833 | 90.218 | 23.197 |  |  | 0.11767 | 0.40964 | 14.1257 | 49.1767
05/19/21 | 46 | 20210048.165 | 749 | 0.842 | 98.59 | 22.77 |  |  | 0.12858 | 0.40231 | 15.2703 | 47.7808
05/19/21 | 47 | 20210048.166 | 751 | 0.808 | 84.742 | 22.603 |  |  | 0.11053 | 0.39945 | 13.6797 | 49.4367
05/19/21 | 48 | 20210048.167 | 753 | 0.827 | 85.867 | 22.957 |  |  | 0.112 | 0.40552 | 13.5426 | 49.0354
05/19/21 | 49 | 20210048.168 | 755 | 0.808 | 87.908 | 22.25 |  |  | 0.11466 | 0.39339 | 14.1902 | 48.687
05/19/21 | 50 | 20210048.169 | 757 | 0.838 | 85.501 | 23.898 |  |  | 0.11152 | 0.42167 | 13.308 | 50.3187
05/19/21 | 51 | 20210048.170 | 759 | 0.84 | 95.244 | 22.993 |  |  | 0.12422 | 0.40614 | 14.7876 | 48.3501
05/19/21 | 52 | 20210048.171 | 761 | 0.816 | 65.097 | 25.469 |  |  | 0.084934 | 0.44863 | 10.4086 | 54.979
05/19/21 | 53 | 20210048.172 | 763 | 0.803 | 85.097 | 22.225 |  |  | 0.11099 | 0.39296 | 13.8225 | 48.9367
05/19/21 | 54 | 20210048.173 | 765 | 0.817 | 88.689 | 22.162 |  |  | 0.11567 | 0.39188 | 14.1585 | 47.9658
05/19/21 | 55 | 20210048.174 | 767 | 0.827 | 90.238 | 22.349 |  |  | 0.11769 | 0.39509 | 14.2313 | 47.7739
05/19/21 | 56 | 20210048.175 | 769 | 0.824 | 80.102 | 23.105 |  |  | 0.10449 | 0.40806 | 12.6803 | 49.5222
05/19/21 | 57 | 20210048.176 | 771 | 0.837 | 91.408 | 22.482 |  |  | 0.11922 | 0.39737 | 14.2434 | 47.4758
05/19/21 | 58 | 20210048.177 | 773 | 0.83 | 91.247 | 22.504 |  |  | 0.11901 | 0.39775 | 14.3383 | 47.9216
05/19/21 | 59 | 20210048.178 | 775 | 0.82 | 77.67 | 24.668 |  |  | 0.10132 | 0.43488 | 12.3557 | 53.0346
05/19/21 | 60 | 20210048.178 | Replicate E11 | 0.836 | 89.905 | 24.036 |  |  | 0.11726 | 0.42404 | 14.0262 | 50.7223
05/19/21 | 61 | 20210048.179 | 777 | 0.845 | 94.484 | 22.507 |  |  | 0.12323 | 0.3978 | 14.5829 | 47.0771
05/19/21 | 62 | 20210048.180 | 779 | 0.85 | 79.55 | 24.596 |  |  | 0.10377 | 0.43365 | 12.2078 | 51.0175
05/19/21 | 63 | 20210048.181 | 781 | 0.81 | 87.73 | 22.253 |  |  | 0.11443 | 0.39344 | 14.1266 | 48.5731
05/19/21 | 64 | 36-UWSIF-UT Glut 1-20210048.38 | 36-UWSIF-UT Glut 1- | 0.813 | 58.494 | 18.433 |  |  | 0.076331 | 0.32789 | 9.3888 | 40.331
05/19/21 | 65 | 36-UWSIF-UT Glut 1-20210048.39 | 36-UWSIF-UT Glut 1- | 0.777 | 55.938 | 17.516 |  |  | 0.073 | 0.31216 | 9.3951 | 40.1745
05/19/21 | 66 | 39-UWSIF-UW Glut 2-20210048.35 | 39-UWSIF-UW Glut 2- | 0.766 | 59.681 | 18.798 |  |  | 0.077877 | 0.33415 | 10.1668 | 43.6233
05/19/21 | 67 | 39-UWSIF-UW Glut 2-20210048.36 | 39-UWSIF-UW Glut 2- | 0.757 | 58.86 | 18.563 |  |  | 0.076808 | 0.33012 | 10.1463 | 43.6093
05/19/21 | 68 | 01-UWSIF-Liver-20210048.35 | 01-UWSIF-Liver- | 0.771 | 75.773 | 21.562 |  |  | 0.098845 | 0.38158 | 12.8204 | 49.4922
05/19/21 | 69 | 01-UWSIF-Liver-20210048.36 | 01-UWSIF-Liver- | 0.765 | 75.26 | 21.355 |  |  | 0.098177 | 0.37803 | 12.8336 | 49.4161
05/19/21 | 70 | 20210048.182 | 783 | 0.808 | 84.773 | 22.098 |  |  | 0.11057 | 0.39078 | 13.6847 | 48.3642
05/19/21 | 71 | 20210048.183 | 785 | 0.801 | 73.063 | 23.79 |  |  | 0.095314 | 0.41982 | 11.8994 | 52.4117
05/19/21 | 72 | 20210048.184 | 787 | 0.815 | 82.242 | 22.913 |  |  | 0.10727 | 0.40477 | 13.1625 | 49.6648
05/19/21 | 73 | 20210048.185 | 789 | 0.805 | 89.822 | 21.561 |  |  | 0.11715 | 0.38157 | 14.5529 | 47.3997
05/19/21 | 74 | 20210048.186 | 791 | 0.848 | 86.819 | 22.76 |  |  | 0.11324 | 0.40214 | 13.3536 | 47.4225
05/19/21 | 75 | 20210048.187 | 793 | 0.83 | 84.969 | 23.116 |  |  | 0.11083 | 0.40825 | 13.3527 | 49.1869
05/19/21 | 76 | 20210048.188 | 795 | 0.845 | 90.577 | 23.576 |  |  | 0.11813 | 0.41615 | 13.9805 | 49.2479
05/19/21 | 77 | 20210048.189 | 797 | 0.826 | 87.564 | 22.889 |  |  | 0.11421 | 0.40436 | 13.8267 | 48.9535
05/19/21 | 78 | 20210048.190 | 799 | 0.845 | 91.833 | 21.963 |  |  | 0.11977 | 0.38847 | 14.1741 | 45.9723
05/19/21 | 79 | 20210048.191 | 801 | 0.848 | 83.132 | 24.75 |  |  | 0.10843 | 0.43629 | 12.787 | 51.4494
05/19/21 | 80 | 20210048.192 | 803 | 0.833 | 90.886 | 22.336 |  |  | 0.11854 | 0.39487 | 14.2302 | 47.403
05/19/21 | 81 | 20210048.193 | 805 | 0.843 | 95.509 | 22.91 |  |  | 0.12456 | 0.40472 | 14.7759 | 48.0091
05/19/21 | 82 | 20210048.194 | 807 | 0.818 | 83.231 | 22.789 |  |  | 0.10856 | 0.40264 | 13.2718 | 49.2225
05/20/21 | 83 | 20210048.195 | 809 | 0.803 | 85.621 | 22.165 |  |  | 0.11168 | 0.39193 | 13.9075 | 48.8085
05/20/21 | 84 | 20210048.196 | 811 | 0.843 | 94.723 | 23.302 |  |  | 0.12354 | 0.41144 | 14.6545 | 48.807
05/20/21 | 85 | 20210048.197 | 813 | 0.831 | 88.051 | 22.475 |  |  | 0.11484 | 0.39725 | 13.8199 | 47.8041
05/20/21 | 86 | 20210048.198 | 815 | 0.814 | 65.375 | 25.697 |  |  | 0.085297 | 0.45254 | 10.4787 | 55.5948
05/20/21 | 87 | 20210048.199 | 817 | 0.802 | 86.456 | 21.714 |  |  | 0.11277 | 0.38419 | 14.0605 | 47.9044
05/20/21 | 88 | 20210048.200 | 819 | 0.825 | 89.548 | 22.475 |  |  | 0.11679 | 0.39725 | 14.1569 | 48.1518
05/20/21 | 89 | 20210048.201 | 821 | 0.831 | 82.874 | 24.173 |  |  | 0.1081 | 0.42639 | 13.0082 | 51.3104
05/20/21 | 90 | 20210048.202 | 823 | 0.81 | 85.465 | 22.198 |  |  | 0.11147 | 0.3925 | 13.7622 | 48.4566
05/20/21 | 91 | 20210048.203 | 825 | 0.837 | 87.284 | 22.969 |  |  | 0.11384 | 0.40573 | 13.6014 | 48.4742
05/20/21 | 92 | 20210048.204 | 827 | 0.836 | 89.663 | 22.88 |  |  | 0.11694 | 0.4042 | 13.9885 | 48.3495
05/20/21 | 93 | 20210048.204 | Replicate H8 | 0.832 | 88.268 | 22.841 |  |  | 0.11513 | 0.40353 | 13.8373 | 48.5015
05/20/21 | 94 | 36-UWSIF-UT Glut 1-20210048.311 | 36-UWSIF-UT Glut 1- | 0.755 | 53.66 | 16.936 |  |  | 0.070032 | 0.3022 | 9.2758 | 40.0269
05/20/21 | 95 | 36-UWSIF-UT Glut 1-20210048.312 | 36-UWSIF-UT Glut 1- | 0.756 | 53.633 | 16.959 |  |  | 0.069997 | 0.3026 | 9.2588 | 40.0261
05/20/21 | 96 | 39-UWSIF-UW Glut 2-20210048.37 | 39-UWSIF-UW Glut 2- | 0.76 | 58.733 | 18.599 |  |  | 0.076642 | 0.33074 | 10.0845 | 43.5184
05/20/21 | 97 | 39-UWSIF-UW Glut 2-20210048.38 | 39-UWSIF-UW Glut 2- | 0.706 | 54.526 | 17.193 |  |  | 0.07116 | 0.30661 | 10.0794 | 43.4296
05/20/21 | 98 | 01-UWSIF-Liver-20210048.37 | 01-UWSIF-Liver- | 0.73 | 71.061 | 20.272 |  |  | 0.092705 | 0.35945 | 12.6994 | 49.2395
05/20/21 | 99 | 01-UWSIF-Liver-20210048.38 | 01-UWSIF-Liver- | 0.723 | 70.645 | 20.098 |  |  | 0.092163 | 0.35646 | 12.7474 | 49.3033

## % Calc 4
Date | Row | Identifier 1 | Identifier 2 | Amount | Area All N | Area All C | Known mg N | Known mg C | Calc mg N | Calc mg C | %N | %C
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 | 33.177 | 10.228 | 0.044839 | 0.19222 | 0.045072 | 0.19299 | 9.5695 | 40.9744
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 | 56.272 | 17.694 | 0.07616 | 0.32648 | 0.075881 | 0.32539 | 9.4851 | 40.6739
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 | 115.459 | 36.757 | 0.1547 | 0.66316 | 0.15484 | 0.66345 | 9.5284 | 40.828
05/20/21 | 1 | 36-UWSIF-UT Glut 1-20210048.41 | 36-UWSIF-UT Glut 1- | 0.794 | 56.116 | 17.503 |  |  | 0.075673 | 0.322 | 9.5306 | 40.5547
05/20/21 | 2 | 36-UWSIF-UT Glut 1-20210048.42 | 36-UWSIF-UT Glut 1- | 0.781 | 54.927 | 17.309 |  |  | 0.074087 | 0.31856 | 9.4861 | 40.7892
05/20/21 | 3 | 36-UWSIF-UT Glut 1-20210048.43 | 36-UWSIF-UT Glut 1- | 0.471 | 33.177 | 10.228 |  |  | 0.045072 | 0.19299 | 9.5695 | 40.9744
05/20/21 | 4 | 36-UWSIF-UT Glut 1-20210048.44 | 36-UWSIF-UT Glut 1- | 0.8 | 56.272 | 17.694 |  |  | 0.075881 | 0.32539 | 9.4851 | 40.6739
05/20/21 | 5 | 36-UWSIF-UT Glut 1-20210048.45 | 36-UWSIF-UT Glut 1- | 1.625 | 115.459 | 36.757 |  |  | 0.15484 | 0.66345 | 9.5284 | 40.828
05/20/21 | 6 | 39-UWSIF-UW Glut 2-20210048.41 | 39-UWSIF-UW Glut 2- | 0.717 | 54.486 | 17.15 |  |  | 0.073498 | 0.31574 | 10.2508 | 44.0368
05/20/21 | 7 | 39-UWSIF-UW Glut 2-20210048.42 | 39-UWSIF-UW Glut 2- | 0.807 | 61.582 | 19.492 |  |  | 0.082964 | 0.35728 | 10.2806 | 44.2723
05/20/21 | 8 | 01-UWSIF-Liver-20210048.41 | 01-UWSIF-Liver- | 0.751 | 72.565 | 20.574 |  |  | 0.097616 | 0.37647 | 12.9981 | 50.1285
05/20/21 | 9 | 01-UWSIF-Liver-20210048.42 | 01-UWSIF-Liver- | 0.809 | 78.183 | 22.221 |  |  | 0.10511 | 0.40567 | 12.9926 | 50.145
05/20/21 | 10 | 20210048.205 | 829 | 0.844 | 84.016 | 23.036 |  |  | 0.11289 | 0.42013 | 13.3758 | 49.778
05/20/21 | 11 | 20210048.206 | 831 | 0.807 | 66.794 | 24.325 |  |  | 0.089917 | 0.44299 | 11.1422 | 54.8929
05/20/21 | 12 | 20210048.206 | Replicate A11 | 0.815 | 65.416 | 23.91 |  |  | 0.088079 | 0.43563 | 10.8072 | 53.451
05/20/21 | 13 | 20210048.207 | 833 | 0.839 | 88.876 | 23.014 |  |  | 0.11937 | 0.41974 | 14.2282 | 50.0282
05/20/21 | 14 | 20210048.208 | 835 | 0.828 | 80.751 | 22.937 |  |  | 0.10854 | 0.41837 | 13.1082 | 50.5279
05/20/21 | 15 | 20210048.209 | 837 | 0.831 | 80.689 | 23.328 |  |  | 0.10845 | 0.4253 | 13.0509 | 51.1799
05/20/21 | 16 | 20210048.210 | 839 | 0.805 | 77.38 | 24.007 |  |  | 0.10404 | 0.43735 | 12.9241 | 54.3287
05/20/21 | 17 | 20210048.211 | 841 | 0.825 | 80.817 | 23.264 |  |  | 0.10862 | 0.42417 | 13.1665 | 51.4145
05/20/21 | 18 | 20210048.212 | 843 | 0.822 | 83.381 | 22.276 |  |  | 0.11204 | 0.40665 | 13.6307 | 49.4706
05/20/21 | 19 | 20210048.213 | 845 | 0.802 | 87.938 | 20.496 |  |  | 0.11812 | 0.37508 | 14.7286 | 46.7683
05/20/21 | 20 | 20210048.214 | 002A | 0.827 | 59.166 | 19.411 |  |  | 0.079741 | 0.35584 | 9.6423 | 43.0279
05/20/21 | 21 | 20210048.215 | 002B | 0.844 | 61.209 | 19.967 |  |  | 0.082467 | 0.3657 | 9.7709 | 43.3295
05/20/21 | 22 | 20210048.216 | 002C | 0.812 | 60.142 | 19.415 |  |  | 0.081043 | 0.35591 | 9.9807 | 43.8315
05/20/21 | 23 | 20210048.217 | 004A | 0.83 | 68.611 | 22.536 |  |  | 0.092341 | 0.41126 | 11.1254 | 49.5493
05/20/21 | 24 | 20210048.218 | 004B | 0.821 | 58.411 | 23.601 |  |  | 0.078734 | 0.43015 | 9.59 | 52.393
05/20/21 | 25 | 20210048.219 | 004C | 0.818 | 49.698 | 24.061 |  |  | 0.067111 | 0.4383 | 8.2043 | 53.5824
05/20/21 | 26 | 20210048.220 | 006A | 0.843 | 61.16 | 21.019 |  |  | 0.082401 | 0.38436 | 9.7748 | 45.5939
05/20/21 | 27 | 20210048.221 | 006B | 0.82 | 59.775 | 20.322 |  |  | 0.080554 | 0.372 | 9.8236 | 45.3654
05/20/21 | 28 | 20210048.222 | 006C | 0.841 | 59.88 | 20.54 |  |  | 0.080694 | 0.37586 | 9.595 | 44.6923
05/20/21 | 29 | 20210048.223 | 008A | 0.834 | 69.238 | 22.543 |  |  | 0.093177 | 0.41138 | 11.1724 | 49.3266
05/20/21 | 30 | 20210048.224 | 008B | 0.85 | 66.371 | 22.678 |  |  | 0.089353 | 0.41378 | 10.5121 | 48.6797
05/20/21 | 31 | 20210048.225 | 008C | 0.81 | 64.74 | 21.251 |  |  | 0.087177 | 0.38847 | 10.7626 | 47.9594
05/20/21 | 32 | 20210048.226 | 010A | 0.849 | 59.123 | 21.386 |  |  | 0.079684 | 0.39087 | 9.3856 | 46.0383
05/20/21 | 33 | 20210048.226 | Replicate C8 | 0.843 | 58.917 | 21.421 |  |  | 0.079409 | 0.39149 | 9.4198 | 46.4396
05/20/21 | 34 | 36-UWSIF-UT Glut 1-20210048.46 | 36-UWSIF-UT Glut 1- | 0.73 | 51.416 | 16.04 |  |  | 0.069403 | 0.29606 | 9.5073 | 40.5561
05/20/21 | 35 | 36-UWSIF-UT Glut 1-20210048.47 | 36-UWSIF-UT Glut 1- | 0.798 | 55.874 | 17.578 |  |  | 0.07535 | 0.32333 | 9.4423 | 40.5181
05/20/21 | 36 | 39-UWSIF-UW Glut 2-20210048.43 | 39-UWSIF-UW Glut 2- | 0.835 | 63.499 | 20.03 |  |  | 0.085522 | 0.36682 | 10.2421 | 43.9303
05/20/21 | 37 | 39-UWSIF-UW Glut 2-20210048.44 | 39-UWSIF-UW Glut 2- | 0.722 | 54.835 | 17.217 |  |  | 0.073964 | 0.31693 | 10.2443 | 43.8964
05/20/21 | 38 | 01-UWSIF-Liver-20210048.43 | 01-UWSIF-Liver- | 0.783 | 75.649 | 21.491 |  |  | 0.10173 | 0.39273 | 12.9923 | 50.1568
05/20/21 | 39 | 01-UWSIF-Liver-20210048.44 | 01-UWSIF-Liver- | 0.727 | 69.945 | 19.829 |  |  | 0.094121 | 0.36325 | 12.9464 | 49.9661
05/20/21 | 40 | 20210048.227 | 010B | 0.84 | 61.43 | 21.583 |  |  | 0.082762 | 0.39436 | 9.8526 | 46.9475
05/20/21 | 41 | 20210048.228 | 010C | 0.814 | 57.938 | 20.38 |  |  | 0.078103 | 0.37302 | 9.595 | 45.8262
05/20/21 | 42 | 20210048.229 | 012A | 0.827 | 68.416 | 21.785 |  |  | 0.092081 | 0.39794 | 11.1343 | 48.1186
05/20/21 | 43 | 20210048.230 | 012B | 0.842 | 65.981 | 21.379 |  |  | 0.088833 | 0.39074 | 10.5502 | 46.4063
05/20/21 | 44 | 20210048.231 | 012C | 0.811 | 68.374 | 21.418 |  |  | 0.092025 | 0.39143 | 11.3471 | 48.2655
05/20/21 | 45 | 20210048.232 | 014A | 0.836 | 70.453 | 20.692 |  |  | 0.094798 | 0.37856 | 11.3395 | 45.282
05/20/21 | 46 | 20210048.233 | 014B | 0.843 | 63.48 | 22.991 |  |  | 0.085496 | 0.41933 | 10.1419 | 49.7424
05/20/21 | 47 | 20210048.234 | 014C | 0.815 | 66.599 | 20.313 |  |  | 0.089657 | 0.37184 | 11.0009 | 45.6241
05/20/21 | 48 | 20210048.235 | 016A | 0.817 | 63.75 | 20.695 |  |  | 0.085856 | 0.37861 | 10.5088 | 46.3416
05/20/21 | 49 | 20210048.236 | 016B | 0.835 | 68.965 | 20.678 |  |  | 0.092813 | 0.37831 | 11.1154 | 45.3065
05/20/21 | 50 | 20210048.237 | 016C | 0.839 | 68.281 | 22.165 |  |  | 0.091901 | 0.40468 | 10.9536 | 48.2336
05/20/21 | 51 | 20210048.238 | 018A | 0.819 | 68.379 | 21.597 |  |  | 0.092032 | 0.39461 | 11.2371 | 48.1816
05/20/21 | 52 | 20210048.239 | 018B | 0.841 | 63.082 | 23.027 |  |  | 0.084965 | 0.41997 | 10.1029 | 49.9366
05/20/21 | 53 | 20210048.240 | 018C | 0.848 | 65.756 | 23.198 |  |  | 0.088533 | 0.423 | 10.4402 | 49.882
05/20/21 | 54 | 36-UWSIF-UT Glut 1-20210048.411 | 36-UWSIF-UT Glut 1- | 0.851 | 59.479 | 18.822 |  |  | 0.080159 | 0.3454 | 9.4194 | 40.587
05/20/21 | 55 | 36-UWSIF-UT Glut 1-20210048.412 | 36-UWSIF-UT Glut 1- | 0.83 | 58.231 | 18.353 |  |  | 0.078494 | 0.33708 | 9.4571 | 40.6118
05/20/21 | 56 | 39-UWSIF-UW Glut 2-20210048.47 | 39-UWSIF-UW Glut 2- | 0.779 | 59.01 | 18.644 |  |  | 0.079533 | 0.34224 | 10.2097 | 43.9331
05/20/21 | 57 | 39-UWSIF-UW Glut 2-20210048.48 | 39-UWSIF-UW Glut 2- | 0.855 | 64.9 | 20.606 |  |  | 0.087391 | 0.37703 | 10.2211 | 44.0974
05/20/21 | 58 | 01-UWSIF-Liver-20210048.47 | 01-UWSIF-Liver- | 0.724 | 69.663 | 19.781 |  |  | 0.093744 | 0.3624 | 12.9481 | 50.0556
05/21/21 | 59 | 01-UWSIF-Liver-20210048.48 | 01-UWSIF-Liver- | 0.723 | 69.191 | 19.66 |  |  | 0.093115 | 0.36026 | 12.8789 | 49.828