## Final Report
Final Report
Job 2022-0028
|  | Isotope(s) requested: | δ13C, δ15N
|  | Instrument Used: | Carlo Erba 1110 Elemental Analyzer coupled to a Thermo Delta V IRMS
|  | Analytical Code: | 010 (combustion/reduction organics)
|  | Units: | δ13C and δ15N values are reported w.r.t. VPDB and AIR-N2 respectively, expressed in per mil.
|  | Principal Investigator: | Nannette Nelson
|  | Job submission contact: | Nannette Nelson
|  | Sample Material(s): | Chironomid Tissue
|  | Project: | Flathead Lake MeHg/SI Study
|  | Date Submited: | 2022-03-09
|  | Number of unknown samples analyzed: | 9
|  | Quality Assurance Reference Material 1: | 36-UWSIF-Glutamic 1
|  | Quality Assurance Reference Material 2: | 39-UWSIF-Glutamic 2
|  | Quality Control Reference Material 3: | 315-UWSIF-Chitin
|  | Number of reference samples analyzed: | 13
Quality Assurance Data |  |  |  |  | Quality Control Data
Reference Material 1 | Known |  | Known |  | Reference Material 3 | Known | Normalized | Known | Normalized
36-UWSIF-Glutamic | δ13C VPDB |  | δ15N AIR-N2 |  | 315-UWSIF-Chitin | δ13C VPDB | δ13C VPDB | δ15N AIR-N2 | δ15N AIR-N2
average (N=5) | -28.28 |  | -4.62 |  | average (N=4) | -18.96 | -18.9728 | 3.39 | 3.2873
standard uncertainty | 0.067243 |  | 0.09869 |  | standard uncertainty |  | 0.078035 |  | 0.37419
Reference Material 2 | Known |  | Known |  | long-term | δ13C VPDB | -19.26 | δ15N AIR-N2 | 2.99
39-UWSIF-Glutamic | δ13C VPDB |  | δ15N AIR-N2 |  | acceptable range | 2 σ = 0.3 | -18.66 | 2 σ = 0.4 | 3.79
average (N=4) | 24.36 |  | 27.89 |  | Weight Percent | Known | Measured | Known | Measured
standard uncertainty | 0.13804 |  | 0.056247 |  |  | Wt% C | Wt% C | Wt% N | Wt% N
[5 empty] | average (N=4) | 42.26 | 43.0733 | 7.25 | 7.5019
[5 empty] | standard uncertainty |  | 0.19585 |  | 0.064856
[5 empty] | percent error (%) |  | 1.9245 |  | 3.4741
Record Keeping |  |  |  |  | Quality Assurance Approval
Date Reported: | 2022-06-09 |  |  |  | Reviewer: | Craig Cook
Date Invoiced: | 2022-06-09 |  |  |  | Title: | Facility Director
Initial: | cjm |  |  |  | Date Reviewed: | 2022-06-09
Comments: |  |  |  |  | Comments: | Data meet all QAQC criteria in the SOP.
Analytical Comments:
*Sample weight percents are calculated using the sample weights reported by the user and are dependent upon the accuracy of these values.
Quality Control Color Legend
Yellow | Peak amplitudes too low for reliable results. Use with extreme caution or rerun the sample.
Green | Weight percent values wrong. Use with caution.
Pink | Potential outlier. Use with caution.
Blue | Sample lost during the analysis. Reload and rerun the sample.
Olive | Possible sample ID problem. Check loading documents.
SIF ID | Sample ID | δ15N | δ13C | Wt% N* | Wt% C* | C:N ratio | Comments
20220028.001 | Yellow Bay Site #1 | 6.5488 | -30.8891 | 9.2273 | 48.7806 | 5.2865
20220028.002 | Yellow Bay Site #2 | 6.8231 | -31.8891 | 9.2626 | 47.4546 | 5.1233
20220028.003 | Yellow Bay Site #3 | [5 empty] | Sample lost during the analysis.
20220028.004 | Midway Site #4 | [5 empty] | Sample lost during the analysis.
20220028.005 | Midway Site #5 | 7.0575 | -31.7222 | 9.4991 | 48.3931 | 5.0945
20220028.006 | Midway Site #6 | 5.306 | -30.998 | 9.7798 | 46.852 | 4.7907
20220028.007 | Blue Bay Site #7 | 7.9681 | -31.6363 | 7.023 | 41.22 | 5.8693
20220028.008 | Blue Bay Site #8 | 6.6904 | -31.8112 | 7.9386 | 35.16 | 4.429
20220028.009 | Blue Bay Site #9 | 8.2224 | -32.2447 | 9.2767 | 37.7506 | 4.0694
Stable Isotope Facility

## QAQC, calculations
Line | QA Reference Material 1 | δ15N | δ13C | Line | QA Reference Material 2 | δ15N | δ13C | Line | QC Reference Material 3 | δ15N | δ13C |  | wt% Nitrogen | wt'% Carbon |  | wt% Nitrogen | wt'% Carbon
|  | normalized | normalized |  |  | normalized | normalized |  |  | normalized | normalized |  | measured values | measured values |  | absolute error | absolute error
3 | 36-UWSIF-UT Glut 1-20220028.13 | -4.5652 | -28.1698 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 27.9031 | 24.3676 | 8 | 315-UWSIF-Chitin20220028.11 | 3.3282 | -19 |  | 7.4845 | 43.2138 |  | 0.23454 | 0.95377
4 | 36-UWSIF-UT Glut 1-20220028.14 | -4.5962 | -28.2607 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 27.8523 | 24.5484 | 9 | 315-UWSIF-Chitin20220028.12 | 3.6833 | -18.8591 |  | 7.5927 | 43.0458 |  | 0.34271 | 0.78578
5 | 36-UWSIF-UT Glut 1-20220028.15 | -4.7907 | -28.3187 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 27.963 | 24.3157 | 23 | 315-UWSIF-Chitin20220028.15 | 2.7806 | -19.0369 |  | 7.4912 | 43.2268 |  | 0.24124 | 0.96678
19 | 36-UWSIF-UT Glut 1-20220028.16 | -4.5563 | -28.3336 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 27.8393 | 24.2198 | 24 | 315-UWSIF-Chitin20220028.16 | 3.3571 | -18.995 |  | 7.439 | 42.8068 |  | 0.18899 | 0.54679
20 | 36-UWSIF-UT Glut 1-20220028.17 | -4.5722 | -28.3147 | 4 | δ15N known | 27.89 |  | 4 | δ15N known | 3.39 |  | wt% N known | 7.25
5 | δ15N known | -4.62 |  |  | δ13C known |  | 24.36 |  | δ13C known |  | -18.96 | wt% C known |  | 42.26
| δ13C known |  | -28.28 |  | average | 27.8894 | 24.3629 |  | average | 3.2873 | -18.9728 | average | 7.5019 | 43.0733 | average | 0.25187 | 0.81328
| average | -4.6161 | -28.2795 |  | std uncertainty | 0.056247 | 0.13804 |  | std uncertainty | 0.37419 | 0.078035 | std uncertainty | 0.064856 | 0.19585 | percent error | 3.4741 | 1.9245
| std uncertainty | 0.09869 | 0.067243 | [5 empty] | long-term | 3.79 | -18.66
[9 empty] | acceptable range | 2.99 | -19.26

## Analysis Information
Principal of operation
●The Finnigan DeltaPlus XP is run in continuous flow mode and connected to a Costech 4010 elemental analyzer via a Finnigan ConFlo III interface.
●The Finnigan Delta V is run in continuous flow mode and connected to either a Carlo Erba 1110 or a Thermo Flash isolink elemental analyzer via a Finnigan ConFlo IV interface.
●A sample contained within a tin capsule is dropped into a combustion reactor held at 1020°C. For nitrogen and carbon analysis, the combustion reactor contains chromium oxide for oxidation and silvered cobaltous/cobaltic oxide for removal of sulfur. When the tin capsule is exposed to a gas flow temporarily enriched with ultra-high purity oxygen (25-30 mL min-1), flash combustion occurs which raises the temperature of the sample to >1700°C. The encapsulated sample, depending on its composition, combusts generating one or more of these gases: N2, NxOx, CO2, and H2O. The reduction reactor contains reduced copper wires for the reduction of nitrogen oxides to N2 and the removal of excess O2. An adsorption trap containing magnesium perchlorate removes the H2O. The remaining N2 and CO2 gases travel through a Porapak Q chromatographic column (80-100 mL min-1 at 50˚C) as the carrier gas and then moves to the ConFlo III open-split interface.
●The gas then moves through the ConFlo III to the Finnigan DeltaPlus XP mass spectrometer. Reference gases are pulsed by the ConFlo III into the mass spectrometer for proper mass balance.
QA/QC
Quality assessment of carbon and nitrogen isotope composition of solids is based on the standard uncertainty of the known value of the secondary laboratory reference material calculated on multiple analyses. For carbon isotope composition, if the standard uncertainty is greater than 0.15‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.3‰). The carbon isotopic composition is reported in per mil relative to VPDB scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST 8542 sucrose ANU, respectively are -26.24‰, +37.76‰, and -10.05‰. For nitrogen isotope ratio composition, if the standard uncertainty is greater than 0.2‰, the unknowns are re-analyzed (until the 2-sigma expanded standard uncertainty of the result is better than 0.4‰). The nitrogen isotopic composition is reported in per mil relative to nitrogen in air on a scale such that USGS40 glutamic acid, USGS41 glutamic acid, and NIST8549 potassium nitrate, respectively are -4.52‰, +47.57‰, and 4.36‰.
Normalization
Whenever isotopic analyses are performed, reference materials must be included with the unknowns. In general two ‘bracketing’ reference materials should be used for correcting purposes. These reference materials ideally would be similar in chemical complexity to the unknowns but have very different isotopic compositions. Using two such reference materials, you can generate two linear equations, one for each referenece material. Since these linear equations will have the same slope and intercept, combining the two equations will generate a single linear equation that can be used to correct the unknowns.
The linear equation is derived as follows:
Slope (m): m=[δRM 1-known – δRM 2-known] / [δRM 1-measured – δRM 2-measured]
Intercept (b): b= δRM 1-known – [m* δRM 1-measured]
Normalization equation:
δSA-COR = m* δSA-measured + b
Where:
δRM 1-known = the accepted delta value of reference material 1
δRM 2-known = the accepted delta value of reference material 2
δRM 1-measured = the measured delta value of reference material 1
δRM 2-measured = the measured delta value of reference material 2
δSA-measured = measured delta value of sample
δSA-COR = corrected delta value of sample
13C and 15N analyses of solids

## Contact
For questions about the analysis, please contact:
|  | Chandelle Macdonald | Master Technician
| email: | cmacdon1@uwyo.edu
| Phone: | (307) 766-6373
Contact information for UW Stable Isotope Facility:
| Personnel: | Dr. David G. Williams | Faculty Director
dgw@uwyo.edu
|  | Craig Cook | Facility Director
ccook21@uwyo.edu
|  | Chandelle Macdonald | Master Technician
cmacdon1@uwyo.edu
| Address: | UWYO Stable Isotope Facility
University of Wyoming
Berry Biodiversity Center Rm 214
Laramie, WY 82071
| email: | uwyosif@uwyo.edu
| Phone: | (307) 766-6373
| Fax: | (307) 766-6403

## Tray 1
Date | Line | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Nitrogen % | Carbon % | C:N | Corr. d15N | Corr. d13C
04/19/22 | 10 | 20220028.001 | Yellow Bay Site #1 | 1.807 | 4814 | 9.237 | 11861 | -21.62 | 9.2273 | 48.7806 | 5.2865 | 6.5488 | -30.8891
04/19/22 | 11 | 20220028.002 | Yellow Bay Site #2 | 1.133 | 2995 | 9.512 | 7668 | -22.621 | 9.2626 | 47.4546 | 5.1233 | 6.8231 | -31.8891
04/19/22 | 12 | 20220028.003 | Yellow Bay Site #3 | 0.856 | [9 empty] | Samples dropped together
04/19/22 | 13 | 20220028.004 | Midway Site #4 | 0.968 | 4679 | 10.848 | 10798 | -21.498 | 16.7921 | 81.3248 | 4.843 | 8.1556 | -30.7672
04/19/22 | 14 | 20220028.005 | Midway Site #5 | 0.852 | 2275 | 9.747 | 5914 | -22.454 | 9.4991 | 48.3931 | 5.0945 | 7.0575 | -31.7222
04/19/22 | 15 | 20220028.006 | Midway Site #6 | 1.467 | 4112 | 7.991 | 9545 | -21.729 | 9.7798 | 46.852 | 4.7907 | 5.306 | -30.998
04/19/22 | 16 | 20220028.007 | Blue Bay Site #7 | 1 | 1952 | 10.66 | 5881 | -22.368 | 7.023 | 41.22 | 5.8693 | 7.9681 | -31.6363
04/19/22 | 17 | 20220028.008 | Blue Bay Site #8 | 1.222 | 2732 | 9.379 | 6156 | -22.543 | 7.9386 | 35.16 | 4.429 | 6.6904 | -31.8112
04/19/22 | 18 | 20220028.009 | Blue Bay Site #9 | 1.394 | 3689 | 10.915 | 7505 | -22.977 | 9.2767 | 37.7506 | 4.0694 | 8.2224 | -32.2447
Reference Check
04/19/22 | 8 | 315-UWSIF-Chitin20220028.11 | 315-UWSIF-Chitin | 0.822 | 1722 | 6.008 | 5103 | -9.719 | 7.4845 | 43.2138 | 5.7737 | 3.3282 | -19
04/19/22 | 9 | 315-UWSIF-Chitin20220028.12 | 315-UWSIF-Chitin | 0.829 | 1765 | 6.364 | 5144 | -9.578 | 7.5927 | 43.0458 | 5.6694 | 3.6833 | -18.8591
04/19/22 | 23 | 315-UWSIF-Chitin20220028.15 | 315-UWSIF-Chitin | 0.814 | 1698 | 5.459 | 5064 | -9.756 | 7.4912 | 43.2268 | 5.7703 | 2.7806 | -19.0369
04/19/22 | 24 | 315-UWSIF-Chitin20220028.16 | 315-UWSIF-Chitin | 0.816 | 1697 | 6.037 | 5053 | -9.714 | 7.439 | 42.8068 | 5.7544 | 3.3571 | -18.995
[5 empty] | average | 5.967 |  | -9.6917 | 7.5019 | 43.0733 |  | 3.2873 | -18.9728
[5 empty] | std. dev. | 0.37517 |  | 0.078113 | 0.064856 | 0.19585 |  | 0.37419 | 0.078035
Reference Material
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 | 1079 | -1.906 | 2435 | -18.898 | 9.5656 | 41.0825 | 4.2948 | -4.5652 | -28.1698
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 | 2721 | -1.937 | 5956 | -18.989 | 9.4942 | 40.6598 | 4.2826 | -4.5962 | -28.2607
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 | 8590 | -2.132 | 15966 | -19.047 | 9.5207 | 40.8165 | 4.2871 | -4.7907 | -28.3187
04/19/22 | 19 | 36-UWSIF-UT Glut 1-20220028.16 | 36-UWSIF-UT Glut 1- | 0.825 | 2219 | -1.897 | 4885 | -19.062 | 9.5572 | 40.5003 | 4.2377 | -4.5563 | -28.3336
04/19/22 | 20 | 36-UWSIF-UT Glut 1-20220028.17 | 36-UWSIF-UT Glut 1- | 0.761 | 2025 | -1.913 | 4491 | -19.043 | 9.4871 | 40.3431 | 4.2524 | -4.5722 | -28.3147
[5 empty] | average | -1.957 |  | -19.0078 | 9.525 | 40.6805 |  | -4.6161 | -28.2795
[5 empty] | std. dev. | 0.098947 |  | 0.06731 | 0.035676 | 0.28583 |  | 0.09869 | 0.067243
04/19/22 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 39-UWSIF-UW Glut 2- | 0.759 | 2201 | 30.647 | 4870 | 33.692 | 10.2984 | 44.0575 | 4.2781 | 27.9031 | 24.3676
04/19/22 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 39-UWSIF-UW Glut 2- | 0.809 | 2356 | 30.596 | 5179 | 33.873 | 10.262 | 43.8061 | 4.2688 | 27.8523 | 24.5484
04/19/22 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 39-UWSIF-UW Glut 2- | 0.752 | 2198 | 30.707 | 4816 | 33.64 | 10.4098 | 43.9355 | 4.2206 | 27.963 | 24.3157
04/19/22 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 39-UWSIF-UW Glut 2- | 0.741 | 2151 | 30.583 | 4756 | 33.544 | 10.3258 | 43.9222 | 4.2536 | 27.8393 | 24.2198
[5 empty] | average | 30.6332 |  | 33.6872 | 10.324 | 43.9303 |  | 27.8894 | 24.3629
[5 empty] | std. dev. | 0.056394 |  | 0.13818 | 0.062886 | 0.1028 |  | 0.056247 | 0.13804
15N Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -1.957 | -4.616
|  | UWSIF 39- UW Glutamic 2 | 30.6332 | 27.888
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF315 (Chitin) | 3.2873 | 3.39 | 0.10271
Percentage
|  | UWSIF315 (Chitin) | 7.5019 | 7.25 | 0.25187
13C Normalization
|  |  | Meas. | Actual
|  | UWSIF 36- UT Glutamic | -19.0078 | -28.279
|  | UWSIF 39- UW Glutamic 2 | 33.6872 | 24.362
Lab QC Check
|  | Ref. Check | Corrected | Actual | Acuracy
|  | UWSIF315 (Chitin) | -18.9728 | -18.96 | 0.012758
Percentage
|  | UWSIF315 (Chitin) | 43.0733 | 42.26 | 0.81328
|  | Date Analyzed: | 2022-04-20
|  | Analyst: | CR
|  | Instrument: | Carlo Erba

## Original 1
Date | Row | Identifier 1 | Identifier 2 | Amount | Component | Peak Nr | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | Amt% | Area All | Preparation | Comment | Area 44 | Area 28 | Is Ref _ | rR 45CO2/44CO2 | rR 29N2/28N2 | R 13C/12C | R 15N/14N | AT% 13C/12C | AT% 15N/14N | BGD 28 | BGD 29 | BGD 30 | BGD 44 | BGD 45 | BGD 46 | Time Code | Sample Dilution | Ampl 29
04/19/22 | 1 | 36-UWSIF-UT Glut 1-20220028.11 | 36-UWSIF-UT Glut 1- | 0.811 |  | 1 | 5770 | 0.044 |  |  | 22.6769 | 106.264 |  | 2022-0028 Nelson Tray 1 |  | 105.5 | 0 |  | 0.72467 |  | 0.0036784 |  | 0.36649 | BGD 28: 28.2; BGD 29: 20.9; BGD 30: 914.2 | 2022/04/19 12:29:23 | 0 | 4180
04/19/22 | 1 | 36-UWSIF-UT Glut 1-20220028.11 | 36-UWSIF-UT Glut 1- | 0.811 |  | 2 | 5768 | 0 |  |  | 22.6828 | 106.292 |  | 2022-0028 Nelson Tray 1 |  | 105.527 | 1 |  | 0.72464 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.0; BGD 29: 20.7; BGD 30: 966.9 | 2022/04/19 12:29:23 | 0 | 4179
04/19/22 | 1 | 36-UWSIF-UT Glut 1-20220028.11 | 36-UWSIF-UT Glut 1- | 0.811 | Nitrogen | 3 | 2242 | -1.81 |  |  | 9.1507 | 42.631 |  | 2022-0028 Nelson Tray 1 |  | 42.325 | 0 |  | 0.72333 |  | 0.0036715 |  | 0.36581 | BGD 28: 28.3; BGD 29: 20.9; BGD 30: 901.7 | 2022/04/19 12:29:23 | 0 | 1622
04/19/22 | 1 | 36-UWSIF-UT Glut 1-20220028.11 | 36-UWSIF-UT Glut 1- | 0.811 | Carbon | 4 |  |  | 4884 | -18.943 | 83.2989 | 111.487 |  | 2022-0028 Nelson Tray 1 | 109.761 |  | 0 | 1.1514 |  | 0.010968 |  | 1.0849 |  | BGD 44: -0.1; BGD 45: -0.1; BGD 46: -0.8 | 2022/04/19 12:29:23 | 78
04/19/22 | 1 | 36-UWSIF-UT Glut 1-20220028.11 | 36-UWSIF-UT Glut 1- | 0.811 |  | 5 |  |  | 5371 | -39.8 | 76.4128 | 100.214 |  | 2022-0028 Nelson Tray 1 | 98.703 |  | 1 | 1.127 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.7; BGD 45: 0.8; BGD 46: 0.2 | 2022/04/19 12:29:23 | 78
04/19/22 | 1 | 36-UWSIF-UT Glut 1-20220028.11 | 36-UWSIF-UT Glut 1- | 0.811 |  | 6 |  |  | 5368 | -40.057 | 76.718 | 100.702 |  | 2022-0028 Nelson Tray 1 | 99.184 |  | 0 | 1.1267 |  | 0.010732 |  | 1.0618 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 1.6 | 2022/04/19 12:29:23 | 78
04/19/22 | 2 | 36-UWSIF-UT Glut 1-20220028.12 | 36-UWSIF-UT Glut 1- | 0.841 |  | 1 | 5771 | 0.088 |  |  | 21.8359 | 106.107 |  | CR 4/19/2022 |  | 105.343 | 0 |  | 0.7246 |  | 0.0036785 |  | 0.3665 | BGD 28: 29.0; BGD 29: 21.8; BGD 30: 821.0 | 2022/04/19 12:39:08 | 0 | 4181
04/19/22 | 2 | 36-UWSIF-UT Glut 1-20220028.12 | 36-UWSIF-UT Glut 1- | 0.841 |  | 2 | 5781 | 0 |  |  | 21.8561 | 106.206 |  | CR 4/19/2022 |  | 105.442 | 1 |  | 0.72454 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.5; BGD 29: 21.4; BGD 30: 901.6 | 2022/04/19 12:39:08 | 0 | 4188
04/19/22 | 2 | 36-UWSIF-UT Glut 1-20220028.12 | 36-UWSIF-UT Glut 1- | 0.841 | Nitrogen | 3 | 2279 | -1.831 |  |  | 9.1833 | 44.373 |  | CR 4/19/2022 |  | 44.054 | 0 |  | 0.72321 |  | 0.0036715 |  | 0.3658 | BGD 28: 28.9; BGD 29: 21.6; BGD 30: 858.4 | 2022/04/19 12:39:08 | 0 | 1649
04/19/22 | 2 | 36-UWSIF-UT Glut 1-20220028.12 | 36-UWSIF-UT Glut 1- | 0.841 | Carbon | 4 |  |  | 4995 | -19.032 | 82.7515 | 115.75 |  | CR 4/19/2022 | 113.959 |  | 0 | 1.1516 |  | 0.010967 |  | 1.0848 |  | BGD 44: 0.1; BGD 45: 0.2; BGD 46: -0.5 | 2022/04/19 12:39:08 | 78
04/19/22 | 2 | 36-UWSIF-UT Glut 1-20220028.12 | 36-UWSIF-UT Glut 1- | 0.841 |  | 5 |  |  | 5389 | -39.8 | 73.9644 | 100.674 |  | CR 4/19/2022 | 99.155 |  | 1 | 1.1274 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 0.4 | 2022/04/19 12:39:08 | 78
04/19/22 | 2 | 36-UWSIF-UT Glut 1-20220028.12 | 36-UWSIF-UT Glut 1- | 0.841 |  | 6 |  |  | 5368 | -40.006 | 74.0714 | 100.852 |  | CR 4/19/2022 | 99.33 |  | 0 | 1.1271 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.7; BGD 45: 2.0; BGD 46: 1.7 | 2022/04/19 12:39:08 | 78
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 |  | 1 | 5777 | 0.107 |  |  | 9.52 | 106.256 |  | He PSI: 1000 |  | 105.492 | 0 |  | 0.72462 |  | 0.0036786 |  | 0.36651 | BGD 28: 29.0; BGD 29: 21.9; BGD 30: 820.8 | 2022/04/19 12:48:54 | 0 | 4186
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 |  | 2 | 5783 | 0 |  |  | 9.52 | 106.335 |  | He PSI: 1000 |  | 105.57 | 1 |  | 0.72454 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.5; BGD 29: 21.5; BGD 30: 902.6 | 2022/04/19 12:48:54 | 0 | 4190
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 | Nitrogen | 3 | 1079 | -1.906 |  |  | 9.52 | 21.249 |  | He PSI: 1000 |  | 21.097 | 0 |  | 0.72316 |  | 0.0036712 |  | 0.36578 | BGD 28: 28.9; BGD 29: 21.7; BGD 30: 853.9 | 2022/04/19 12:48:54 | 0 | 781
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 | Carbon | 4 |  |  | 2435 | -18.898 | 40.81 | 55.168 |  | He PSI: 1000 | 54.313 |  | 0 | 1.1519 |  | 0.010969 |  | 1.085 |  | BGD 44: 0.1; BGD 45: 0.3; BGD 46: -0.4 | 2022/04/19 12:48:54 | 78
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 |  | 5 |  |  | 5387 | -39.8 | 40.81 | 100.621 |  | He PSI: 1000 | 99.102 |  | 1 | 1.1275 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.3; BGD 45: 0.5; BGD 46: -0.2 | 2022/04/19 12:48:54 | 78
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 |  | 6 |  |  | 5379 | -40.047 | 40.81 | 100.699 |  | He PSI: 1000 | 99.18 |  | 0 | 1.1272 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.3; BGD 45: 1.6; BGD 46: 1.2 | 2022/04/19 12:48:54 | 78
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 |  | 1 | 5769 | 0.098 |  |  | 9.52 | 105.991 |  | O2 PSI: 1800 |  | 105.229 | 0 |  | 0.72464 |  | 0.0036786 |  | 0.36651 | BGD 28: 28.6; BGD 29: 21.6; BGD 30: 818.3 | 2022/04/19 12:58:40 | 0 | 4179
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 |  | 2 | 5779 | 0 |  |  | 9.52 | 106.131 |  | O2 PSI: 1800 |  | 105.367 | 1 |  | 0.72457 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.1; BGD 29: 21.1; BGD 30: 900.3 | 2022/04/19 12:58:40 | 0 | 4187
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 | Nitrogen | 3 | 2721 | -1.937 |  |  | 9.52 | 52.54 |  | O2 PSI: 1800 |  | 52.162 | 0 |  | 0.72316 |  | 0.0036711 |  | 0.36577 | BGD 28: 28.6; BGD 29: 21.4; BGD 30: 859.9 | 2022/04/19 12:58:40 | 0 | 1969
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 | Carbon | 4 |  |  | 5956 | -18.989 | 40.81 | 138.31 |  | O2 PSI: 1800 | 136.171 |  | 0 | 1.1516 |  | 0.010968 |  | 1.0849 |  | BGD 44: 0.0; BGD 45: 0.1; BGD 46: -0.5 | 2022/04/19 12:58:40 | 78
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 |  | 5 |  |  | 5371 | -39.8 | 40.81 | 100.34 |  | O2 PSI: 1800 | 98.826 |  | 1 | 1.1274 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 0.6 | 2022/04/19 12:58:40 | 78
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 |  | 6 |  |  | 5370 | -40.01 | 40.81 | 100.517 |  | O2 PSI: 1800 | 99.001 |  | 0 | 1.1271 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.8; BGD 45: 2.1; BGD 46: 1.9 | 2022/04/19 12:58:40 | 78
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 |  | 1 | 5747 | 0.108 |  |  | 9.52 | 105.75 |  | M28: 28 |  | 104.989 | 0 |  | 0.72475 |  | 0.0036786 |  | 0.36651 | BGD 28: 28.8; BGD 29: 21.9; BGD 30: 830.0 | 2022/04/19 13:08:25 | 0 | 4165
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 |  | 2 | 5749 | 0 |  |  | 9.52 | 105.9 |  | M28: 28 |  | 105.138 | 1 |  | 0.72467 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.3; BGD 29: 21.4; BGD 30: 911.7 | 2022/04/19 13:08:25 | 0 | 4166
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 | Nitrogen | 3 | 8590 | -2.132 |  |  | 9.52 | 158.921 |  | M28: 28 |  | 157.78 | 0 |  | 0.72312 |  | 0.0036704 |  | 0.36569 | BGD 28: 28.8; BGD 29: 21.7; BGD 30: 874.6 | 2022/04/19 13:08:25 | 0 | 6214
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 | Carbon | 4 |  |  | 15966 | -19.047 | 40.81 | 421.985 |  | M28: 28 | 415.471 |  | 0 | 1.151 |  | 0.010967 |  | 1.0848 |  | BGD 44: 0.1; BGD 45: 0.2; BGD 46: -0.3 | 2022/04/19 13:08:25 | 78
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 |  | 5 |  |  | 5353 | -39.8 | 40.81 | 100.051 |  | M28: 28 | 98.542 |  | 1 | 1.1272 |  | 0.010735 |  | 1.0621 |  | BGD 44: 3.2; BGD 45: 3.7; BGD 46: 3.9 | 2022/04/19 13:08:25 | 78
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 |  | 6 |  |  | 5350 | -39.978 | 40.81 | 100.282 |  | M28: 28 | 98.769 |  | 0 | 1.127 |  | 0.010733 |  | 1.0619 |  | BGD 44: 3.5; BGD 45: 4.0; BGD 46: 4.3 | 2022/04/19 13:08:25 | 78
04/19/22 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 39-UWSIF-UW Glut 2- | 0.759 |  | 1 | 5735 | 0.109 |  |  | 25.1295 | 105.586 |  | M29: 21 |  | 104.826 | 0 |  | 0.72484 |  | 0.0036786 |  | 0.36651 | BGD 28: 30.0; BGD 29: 23.2; BGD 30: 866.3 | 2022/04/19 13:18:11 | 0 | 4156
04/19/22 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 39-UWSIF-UW Glut 2- | 0.759 |  | 2 | 5742 | 0 |  |  | 25.1165 | 105.53 |  | M29: 21 |  | 104.771 | 1 |  | 0.72476 |  | 0.0036782 |  | 0.36647 | BGD 28: 29.4; BGD 29: 22.6; BGD 30: 947.5 | 2022/04/19 13:18:11 | 0 | 4161
04/19/22 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 39-UWSIF-UW Glut 2- | 0.759 | Nitrogen | 3 | 2201 | 30.647 |  |  | 10.329 | 42.987 |  | M29: 21 |  | 42.668 | 0 |  | 0.74697 |  | 0.0037909 |  | 0.37766 | BGD 28: 29.7; BGD 29: 22.7; BGD 30: 900.1 | 2022/04/19 13:18:11 | 0 | 1645
04/19/22 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 39-UWSIF-UW Glut 2- | 0.759 | Carbon | 4 |  |  | 4870 | 33.692 | 92.1566 | 112.755 |  | M29: 21 | 110.947 |  | 0 | 1.2091 |  | 0.011557 |  | 1.1425 |  | BGD 44: 0.5; BGD 45: 0.7; BGD 46: 0.2 | 2022/04/19 13:18:11 | 78
04/19/22 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 39-UWSIF-UW Glut 2- | 0.759 |  | 5 |  |  | 5350 | -39.8 | 83.6965 | 99.969 |  | M29: 21 | 98.461 |  | 1 | 1.1272 |  | 0.010735 |  | 1.0621 |  | BGD 44: 1.0; BGD 45: 1.3; BGD 46: 0.8 | 2022/04/19 13:18:11 | 78
04/19/22 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 39-UWSIF-UW Glut 2- | 0.759 |  | 6 |  |  | 5348 | -40.035 | 83.8439 | 100.186 |  | M29: 21 | 98.674 |  | 0 | 1.127 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.8; BGD 45: 2.2; BGD 46: 2.1 | 2022/04/19 13:18:11 | 78
04/19/22 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 39-UWSIF-UW Glut 2- | 0.809 |  | 1 | 5734 | 0.089 |  |  | 23.5515 | 105.473 |  | M30: 408 |  | 104.714 | 0 |  | 0.72482 |  | 0.0036785 |  | 0.3665 | BGD 28: 29.1; BGD 29: 22.3; BGD 30: 863.4 | 2022/04/19 13:27:56 | 0 | 4156
04/19/22 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 39-UWSIF-UW Glut 2- | 0.809 |  | 2 | 5741 | 0 |  |  | 23.5876 | 105.637 |  | M30: 408 |  | 104.877 | 1 |  | 0.72476 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.6; BGD 29: 21.8; BGD 30: 946.6 | 2022/04/19 13:27:56 | 0 | 4160
04/19/22 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 39-UWSIF-UW Glut 2- | 0.809 | Nitrogen | 3 | 2356 | 30.596 |  |  | 10.2958 | 45.69 |  | M30: 408 |  | 45.351 | 0 |  | 0.74693 |  | 0.0037907 |  | 0.37764 | BGD 28: 29.0; BGD 29: 22.0; BGD 30: 902.3 | 2022/04/19 13:27:56 | 0 | 1760
04/19/22 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 39-UWSIF-UW Glut 2- | 0.809 | Carbon | 4 |  |  | 5179 | 33.873 | 90.5582 | 119.676 |  | M30: 408 | 117.758 |  | 0 | 1.2092 |  | 0.011559 |  | 1.1427 |  | BGD 44: 0.3; BGD 45: 0.4; BGD 46: -0.2 | 2022/04/19 13:27:56 | 78
04/19/22 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 39-UWSIF-UW Glut 2- | 0.809 |  | 5 |  |  | 5359 | -39.8 | 78.7116 | 100.263 |  | M30: 408 | 98.75 |  | 1 | 1.1273 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.9; BGD 45: 1.2; BGD 46: 0.7 | 2022/04/19 13:27:56 | 78
04/19/22 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 39-UWSIF-UW Glut 2- | 0.809 |  | 6 |  |  | 5363 | -40.029 | 78.6936 | 100.235 |  | M30: 408 | 98.723 |  | 0 | 1.127 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.7; BGD 45: 2.1; BGD 46: 1.9 | 2022/04/19 13:27:56 | 78
04/19/22 | 8 | 315-UWSIF-Chitin20220028.11 | 315-UWSIF-Chitin | 0.822 |  | 1 | 5755 | 0.095 |  |  | 23.2424 | 105.766 |  | Peak Center: 2.976 |  | 105.004 | 0 |  | 0.72481 |  | 0.0036785 |  | 0.36651 | BGD 28: 29.0; BGD 29: 22.2; BGD 30: 870.7 | 2022/04/19 13:37:42 | 0 | 4171
04/19/22 | 8 | 315-UWSIF-Chitin20220028.11 | 315-UWSIF-Chitin | 0.822 |  | 2 | 5742 | 0 |  |  | 23.2507 | 105.804 |  | Peak Center: 2.976 |  | 105.043 | 1 |  | 0.72474 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.5; BGD 29: 21.7; BGD 30: 953.6 | 2022/04/19 13:37:42 | 0 | 4161
04/19/22 | 8 | 315-UWSIF-Chitin20220028.11 | 315-UWSIF-Chitin | 0.822 | Nitrogen | 3 | 1722 | 6.008 |  |  | 7.4922 | 33.721 |  | Peak Center: 2.976 |  | 33.477 | 0 |  | 0.72909 |  | 0.0037003 |  | 0.36867 | BGD 28: 28.9; BGD 29: 21.9; BGD 30: 906.2 | 2022/04/19 13:37:42 | 0 | 1256
04/19/22 | 8 | 315-UWSIF-Chitin20220028.11 | 315-UWSIF-Chitin | 0.822 | Carbon | 4 |  |  | 5103 | -9.719 | 89.2894 | 119.962 |  | Peak Center: 2.976 | 118.094 |  | 0 | 1.1616 |  | 0.011071 |  | 1.095 |  | BGD 44: 0.2; BGD 45: 0.4; BGD 46: -0.2 | 2022/04/19 13:37:42 | 78
04/19/22 | 8 | 315-UWSIF-Chitin20220028.11 | 315-UWSIF-Chitin | 0.822 |  | 5 |  |  | 5362 | -39.8 | 77.5218 | 100.351 |  | Peak Center: 2.976 | 98.837 |  | 1 | 1.1272 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 0.6 | 2022/04/19 13:37:42 | 78
04/19/22 | 8 | 315-UWSIF-Chitin20220028.11 | 315-UWSIF-Chitin | 0.822 |  | 6 |  |  | 5365 | -40.023 | 77.6299 | 100.523 |  | Peak Center: 2.976 | 99.007 |  | 0 | 1.127 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.7; BGD 45: 2.1; BGD 46: 1.9 | 2022/04/19 13:37:42 | 78
04/19/22 | 9 | 315-UWSIF-Chitin20220028.12 | 315-UWSIF-Chitin | 0.829 |  | 1 | 5758 | 0.102 |  |  | 23.0887 | 105.964 |  | Instrument: Carlo Erba 1110 EA coupled to Thermo Delta V IRMS |  | 105.202 | 0 |  | 0.72483 |  | 0.0036786 |  | 0.36651 | BGD 28: 28.9; BGD 29: 22.1; BGD 30: 874.8 | 2022/04/19 13:47:27 | 0 | 4173
04/19/22 | 9 | 315-UWSIF-Chitin20220028.12 | 315-UWSIF-Chitin | 0.829 |  | 2 | 5765 | 0 |  |  | 23.107 | 106.05 |  | Instrument: Carlo Erba 1110 EA coupled to Thermo Delta V IRMS |  | 105.287 | 1 |  | 0.72475 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.4; BGD 29: 21.6; BGD 30: 958.7 | 2022/04/19 13:47:27 | 0 | 4178
04/19/22 | 9 | 315-UWSIF-Chitin20220028.12 | 315-UWSIF-Chitin | 0.829 | Nitrogen | 3 | 1765 | 6.364 |  |  | 7.6022 | 34.512 |  | Instrument: Carlo Erba 1110 EA coupled to Thermo Delta V IRMS |  | 34.262 | 0 |  | 0.72936 |  | 0.0037016 |  | 0.3688 | BGD 28: 28.8; BGD 29: 21.9; BGD 30: 910.6 | 2022/04/19 13:47:27 | 0 | 1288
04/19/22 | 9 | 315-UWSIF-Chitin20220028.12 | 315-UWSIF-Chitin | 0.829 | Carbon | 4 |  |  | 5144 | -9.578 | 88.8555 | 120.527 |  | Instrument: Carlo Erba 1110 EA coupled to Thermo Delta V IRMS | 118.65 |  | 0 | 1.1618 |  | 0.011073 |  | 1.0952 |  | BGD 44: 0.2; BGD 45: 0.4; BGD 46: -0.2 | 2022/04/19 13:47:27 | 78
04/19/22 | 9 | 315-UWSIF-Chitin20220028.12 | 315-UWSIF-Chitin | 0.829 |  | 5 |  |  | 5391 | -39.8 | 76.9909 | 100.549 |  | Instrument: Carlo Erba 1110 EA coupled to Thermo Delta V IRMS | 99.032 |  | 1 | 1.1272 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 0.6 | 2022/04/19 13:47:27 | 78
04/19/22 | 9 | 315-UWSIF-Chitin20220028.12 | 315-UWSIF-Chitin | 0.829 |  | 6 |  |  | 5369 | -40.008 | 77.0682 | 100.674 |  | Instrument: Carlo Erba 1110 EA coupled to Thermo Delta V IRMS | 99.155 |  | 0 | 1.127 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.7; BGD 45: 2.0; BGD 46: 1.9 | 2022/04/19 13:47:27 | 78
04/19/22 | 10 | 20220028.001 | Yellow Bay Site #1 | 1.807 |  | 1 | 5765 | 0.1 |  |  | 10.5962 | 106.002 |  | Comments: |  | 105.24 | 0 |  | 0.72482 |  | 0.0036786 |  | 0.36651 | BGD 28: 28.8; BGD 29: 22.0; BGD 30: 879.5 | 2022/04/19 13:57:13 | 0 | 4178
04/19/22 | 10 | 20220028.001 | Yellow Bay Site #1 | 1.807 |  | 2 | 5760 | 0 |  |  | 10.609 | 106.132 |  | Comments: |  | 105.369 | 1 |  | 0.72474 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.3; BGD 29: 21.6; BGD 30: 962.7 | 2022/04/19 13:57:13 | 0 | 4175
04/19/22 | 10 | 20220028.001 | Yellow Bay Site #1 | 1.807 | Nitrogen | 3 | 4814 | 9.237 |  |  | 9.2463 | 92.304 |  | Comments: |  | 91.634 | 0 |  | 0.73144 |  | 0.0037122 |  | 0.36984 | BGD 28: 28.8; BGD 29: 21.8; BGD 30: 919.8 | 2022/04/19 13:57:13 | 0 | 3523
04/19/22 | 10 | 20220028.001 | Yellow Bay Site #1 | 1.807 | Carbon | 4 |  |  | 11861 | -21.62 | 66.3038 | 302.118 |  | Comments: | 297.46 |  | 0 | 1.1482 |  | 0.010939 |  | 1.082 |  | BGD 44: 0.2; BGD 45: 0.3; BGD 46: -0.2 | 2022/04/19 13:57:13 | 78
04/19/22 | 10 | 20220028.001 | Yellow Bay Site #1 | 1.807 |  | 5 |  |  | 5387 | -39.8 | 35.3426 | 100.624 |  | Comments: | 99.106 |  | 1 | 1.127 |  | 0.010735 |  | 1.0621 |  | BGD 44: 2.3; BGD 45: 2.7; BGD 46: 2.6 | 2022/04/19 13:57:13 | 78
04/19/22 | 10 | 20220028.001 | Yellow Bay Site #1 | 1.807 |  | 6 |  |  | 5373 | -39.983 | 35.4395 | 100.964 |  | Comments: | 99.441 |  | 0 | 1.1268 |  | 0.010733 |  | 1.0619 |  | BGD 44: 2.7; BGD 45: 3.2; BGD 46: 3.4 | 2022/04/19 13:57:13 | 78
04/19/22 | 11 | 20220028.002 | Yellow Bay Site #2 | 1.133 |  | 1 | 5764 | 0.099 |  |  | 16.89 | 105.941 |  |  |  | 105.178 | 0 |  | 0.72481 |  | 0.0036786 |  | 0.36651 | BGD 28: 29.5; BGD 29: 22.8; BGD 30: 900.6 | 2022/04/19 14:06:58 | 0 | 4177
04/19/22 | 11 | 20220028.002 | Yellow Bay Site #2 | 1.133 |  | 2 | 5765 | 0 |  |  | 16.8972 | 105.987 |  |  |  | 105.224 | 1 |  | 0.72474 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.9; BGD 29: 22.2; BGD 30: 983.4 | 2022/04/19 14:06:58 | 0 | 4178
04/19/22 | 11 | 20220028.002 | Yellow Bay Site #2 | 1.133 | Nitrogen | 3 | 2995 | 9.512 |  |  | 9.2986 | 57.898 |  |  |  | 57.478 | 0 |  | 0.73163 |  | 0.0037132 |  | 0.36995 | BGD 28: 29.2; BGD 29: 22.4; BGD 30: 935.6 | 2022/04/19 14:06:58 | 0 | 2193
04/19/22 | 11 | 20220028.002 | Yellow Bay Site #2 | 1.133 | Carbon | 4 |  |  | 7668 | -22.621 | 86.821 | 183.113 |  |  | 180.288 |  | 0 | 1.1473 |  | 0.010927 |  | 1.0809 |  | BGD 44: 0.4; BGD 45: 0.6; BGD 46: 0.1 | 2022/04/19 14:06:58 | 78
04/19/22 | 11 | 20220028.002 | Yellow Bay Site #2 | 1.133 |  | 5 |  |  | 5381 | -39.8 | 56.3875 | 100.669 |  |  | 99.15 |  | 1 | 1.127 |  | 0.010735 |  | 1.0621 |  | BGD 44: 1.5; BGD 45: 1.8; BGD 46: 1.5 | 2022/04/19 14:06:58 | 78
04/19/22 | 11 | 20220028.002 | Yellow Bay Site #2 | 1.133 |  | 6 |  |  | 5384 | -39.977 | 56.4597 | 100.828 |  |  | 99.307 |  | 0 | 1.1268 |  | 0.010733 |  | 1.0619 |  | BGD 44: 2.2; BGD 45: 2.5; BGD 46: 2.5 | 2022/04/19 14:06:58 | 78
04/19/22 | 12 | 20220028.003 | Yellow Bay Site #3 | 0.856 |  | 1 | 5771 | 0.135 |  |  | 22.3752 | 106.036 |  |  |  | 105.273 | 0 |  | 0.72481 |  | 0.0036787 |  | 0.36652 | BGD 28: 29.1; BGD 29: 22.4; BGD 30: 902.6 | 2022/04/19 14:16:44 | 0 | 4183
04/19/22 | 12 | 20220028.003 | Yellow Bay Site #3 | 0.856 |  | 2 | 5768 | 0 |  |  | 22.3768 | 106.043 |  |  |  | 105.28 | 1 |  | 0.72471 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.5; BGD 29: 21.9; BGD 30: 986.2 | 2022/04/19 14:16:44 | 0 | 4180
04/19/22 | 12 | 20220028.003 | Yellow Bay Site #3 | 0.856 |  | 3 |  |  | 5391 | -39.8 | 74.6003 | 100.612 |  |  | 99.094 |  | 1 | 1.1273 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.0; BGD 45: 0.1; BGD 46: -0.6 | 2022/04/19 14:16:44 | 78
04/19/22 | 12 | 20220028.003 | Yellow Bay Site #3 | 0.856 |  | 4 |  |  | 5387 | -40.127 | 74.8263 | 100.988 |  |  | 99.465 |  | 0 | 1.1269 |  | 0.010732 |  | 1.0618 |  | BGD 44: 1.0; BGD 45: 1.2; BGD 46: 0.9 | 2022/04/19 14:16:44 | 78
04/19/22 | 13 | 20220028.004 | Midway Site #4 | 0.968 |  | 1 | 5776 | 0.091 |  |  | 19.8047 | 106.135 |  |  |  | 105.372 | 0 |  | 0.72477 |  | 0.0036785 |  | 0.36651 | BGD 28: 28.1; BGD 29: 21.3; BGD 30: 885.2 | 2022/04/19 14:26:29 | 0 | 4186
04/19/22 | 13 | 20220028.004 | Midway Site #4 | 0.968 |  | 2 | 5772 | 0 |  |  | 19.8124 | 106.177 |  |  |  | 105.413 | 1 |  | 0.7247 |  | 0.0036782 |  | 0.36647 | BGD 28: 27.7; BGD 29: 21.0; BGD 30: 970.9 | 2022/04/19 14:26:29 | 0 | 4183
04/19/22 | 13 | 20220028.004 | Midway Site #4 | 0.968 | Nitrogen | 3 | 4679 | 10.848 |  |  | 16.83 | 89.971 |  |  |  | 89.316 | 0 |  | 0.73256 |  | 0.0037181 |  | 0.37043 | BGD 28: 28.2; BGD 29: 21.3; BGD 30: 931.6 | 2022/04/19 14:26:29 | 0 | 3430
04/19/22 | 13 | 20220028.004 | Midway Site #4 | 0.968 | Carbon | 4 |  |  | 10798 | -21.498 | 121.1396 | 269.497 |  |  | 265.34 |  | 0 | 1.1482 |  | 0.01094 |  | 1.0821 |  | BGD 44: 0.1; BGD 45: 0.2; BGD 46: -0.3 | 2022/04/19 14:26:29 | 78
04/19/22 | 13 | 20220028.004 | Midway Site #4 | 0.968 |  | 5 |  |  | 5389 | -39.8 | 66.0666 | 100.796 |  |  | 99.275 |  | 1 | 1.1269 |  | 0.010735 |  | 1.0621 |  | BGD 44: 2.0; BGD 45: 2.3; BGD 46: 2.2 | 2022/04/19 14:26:29 | 78
04/19/22 | 13 | 20220028.004 | Midway Site #4 | 0.968 |  | 6 |  |  | 5380 | -39.987 | 66.0537 | 100.772 |  |  | 99.252 |  | 0 | 1.1267 |  | 0.010733 |  | 1.0619 |  | BGD 44: 2.5; BGD 45: 2.9; BGD 46: 2.9 | 2022/04/19 14:26:29 | 78
04/19/22 | 14 | 20220028.005 | Midway Site #5 | 0.852 |  | 1 | 5756 | 0.094 |  |  | 22.4486 | 105.884 |  |  |  | 105.122 | 0 |  | 0.72475 |  | 0.0036785 |  | 0.36651 | BGD 28: 29.2; BGD 29: 22.5; BGD 30: 910.4 | 2022/04/19 14:36:15 | 0 | 4171
04/19/22 | 14 | 20220028.005 | Midway Site #5 | 0.852 |  | 2 | 5762 | 0 |  |  | 22.4467 | 105.874 |  |  |  | 105.113 | 1 |  | 0.72469 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.6; BGD 29: 22.0; BGD 30: 993.7 | 2022/04/19 14:36:15 | 0 | 4176
04/19/22 | 14 | 20220028.005 | Midway Site #5 | 0.852 | Nitrogen | 3 | 2275 | 9.747 |  |  | 9.5293 | 44.528 |  |  |  | 44.204 | 0 |  | 0.73175 |  | 0.0037141 |  | 0.37003 | BGD 28: 28.9; BGD 29: 22.1; BGD 30: 943.9 | 2022/04/19 14:36:15 | 0 | 1666
04/19/22 | 14 | 20220028.005 | Midway Site #5 | 0.852 | Carbon | 4 |  |  | 5914 | -22.454 | 96.5072 | 139.724 |  |  | 137.568 |  | 0 | 1.1474 |  | 0.010929 |  | 1.0811 |  | BGD 44: 0.4; BGD 45: 0.5; BGD 46: 0.1 | 2022/04/19 14:36:15 | 78
04/19/22 | 14 | 20220028.005 | Midway Site #5 | 0.852 |  | 5 |  |  | 5388 | -39.8 | 74.9681 | 100.641 |  |  | 99.123 |  | 1 | 1.1269 |  | 0.010735 |  | 1.0621 |  | BGD 44: 1.1; BGD 45: 1.4; BGD 46: 1.0 | 2022/04/19 14:36:15 | 78
04/19/22 | 14 | 20220028.005 | Midway Site #5 | 0.852 |  | 6 |  |  | 5393 | -39.999 | 75.0468 | 100.771 |  |  | 99.252 |  | 0 | 1.1267 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.9; BGD 45: 2.2; BGD 46: 2.1 | 2022/04/19 14:36:15 | 78
04/19/22 | 15 | 20220028.006 | Midway Site #6 | 1.467 |  | 1 | 5755 | 0.124 |  |  | 13.0157 | 105.703 |  |  |  | 104.942 | 0 |  | 0.72474 |  | 0.0036787 |  | 0.36652 | BGD 28: 28.7; BGD 29: 22.1; BGD 30: 908.3 | 2022/04/19 14:46:00 | 0 | 4171
04/19/22 | 15 | 20220028.006 | Midway Site #6 | 1.467 |  | 2 | 5768 | 0 |  |  | 13.0463 | 105.956 |  |  |  | 105.193 | 1 |  | 0.72465 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.2; BGD 29: 21.6; BGD 30: 992.8 | 2022/04/19 14:46:00 | 0 | 4180
04/19/22 | 15 | 20220028.006 | Midway Site #6 | 1.467 | Nitrogen | 3 | 4112 | 7.991 |  |  | 9.8101 | 79.348 |  |  |  | 78.773 | 0 |  | 0.73044 |  | 0.0037076 |  | 0.36939 | BGD 28: 28.7; BGD 29: 21.8; BGD 30: 948.2 | 2022/04/19 14:46:00 | 0 | 3005
04/19/22 | 15 | 20220028.006 | Midway Site #6 | 1.467 | Carbon | 4 |  |  | 9545 | -21.729 | 76.221 | 234.916 |  |  | 231.294 |  | 0 | 1.1479 |  | 0.010937 |  | 1.0819 |  | BGD 44: 0.3; BGD 45: 0.4; BGD 46: -0.1 | 2022/04/19 14:46:00 | 78
04/19/22 | 15 | 20220028.006 | Midway Site #6 | 1.467 |  | 5 |  |  | 5385 | -39.8 | 43.547 | 100.662 |  |  | 99.144 |  | 1 | 1.1268 |  | 0.010735 |  | 1.0621 |  | BGD 44: 1.8; BGD 45: 2.2; BGD 46: 2.0 | 2022/04/19 14:46:00 | 78
04/19/22 | 15 | 20220028.006 | Midway Site #6 | 1.467 |  | 6 |  |  | 5378 | -40.01 | 43.6173 | 100.862 |  |  | 99.341 |  | 0 | 1.1266 |  | 0.010733 |  | 1.0619 |  | BGD 44: 2.4; BGD 45: 2.8; BGD 46: 2.8 | 2022/04/19 14:46:00 | 78
04/19/22 | 16 | 20220028.007 | Blue Bay Site #7 | 1 |  | 1 | 5758 | 0.119 |  |  | 19.0957 | 105.712 |  |  |  | 104.951 | 0 |  | 0.72473 |  | 0.0036786 |  | 0.36652 | BGD 28: 29.1; BGD 29: 22.5; BGD 30: 921.5 | 2022/04/19 14:55:46 | 0 | 4172
04/19/22 | 16 | 20220028.007 | Blue Bay Site #7 | 1 |  | 2 | 5770 | 0 |  |  | 19.1481 | 106.007 |  |  |  | 105.244 | 1 |  | 0.72464 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.5; BGD 29: 22.0; BGD 30: 1005.9 | 2022/04/19 14:55:46 | 0 | 4181
04/19/22 | 16 | 20220028.007 | Blue Bay Site #7 | 1 | Nitrogen | 3 | 1952 | 10.66 |  |  | 7.0388 | 38.569 |  |  |  | 38.289 | 0 |  | 0.73237 |  | 0.0037174 |  | 0.37036 | BGD 28: 28.9; BGD 29: 22.1; BGD 30: 953.0 | 2022/04/19 14:55:46 | 0 | 1430
04/19/22 | 16 | 20220028.007 | Blue Bay Site #7 | 1 | Carbon | 4 |  |  | 5881 | -22.368 | 82.2078 | 139.686 |  |  | 137.53 |  | 0 | 1.1473 |  | 0.01093 |  | 1.0812 |  | BGD 44: 0.4; BGD 45: 0.5; BGD 46: 0.0 | 2022/04/19 14:55:46 | 78
04/19/22 | 16 | 20220028.007 | Blue Bay Site #7 | 1 |  | 5 |  |  | 5379 | -39.8 | 63.8614 | 100.619 |  |  | 99.102 |  | 1 | 1.1268 |  | 0.010735 |  | 1.0621 |  | BGD 44: 1.2; BGD 45: 1.4; BGD 46: 1.1 | 2022/04/19 14:55:46 | 78
04/19/22 | 16 | 20220028.007 | Blue Bay Site #7 | 1 |  | 6 |  |  | 5381 | -40.011 | 63.9341 | 100.76 |  |  | 99.241 |  | 0 | 1.1265 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.9; BGD 45: 2.2; BGD 46: 2.2 | 2022/04/19 14:55:46 | 78
04/19/22 | 17 | 20220028.008 | Blue Bay Site #8 | 1.222 |  | 1 | 5751 | 0.081 |  |  | 15.6325 | 105.753 |  |  |  | 104.992 | 0 |  | 0.72471 |  | 0.0036785 |  | 0.3665 | BGD 28: 28.7; BGD 29: 22.1; BGD 30: 919.0 | 2022/04/19 15:05:31 | 0 | 4167
04/19/22 | 17 | 20220028.008 | Blue Bay Site #8 | 1.222 |  | 2 | 5758 | 0 |  |  | 15.6394 | 105.8 |  |  |  | 105.039 | 1 |  | 0.72465 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.2; BGD 29: 21.6; BGD 30: 1004.2 | 2022/04/19 15:05:31 | 0 | 4172
04/19/22 | 17 | 20220028.008 | Blue Bay Site #8 | 1.222 | Nitrogen | 3 | 2732 | 9.379 |  |  | 7.9689 | 53.48 |  |  |  | 53.092 | 0 |  | 0.73145 |  | 0.0037127 |  | 0.3699 | BGD 28: 28.6; BGD 29: 21.8; BGD 30: 955.3 | 2022/04/19 15:05:31 | 0 | 1999
04/19/22 | 17 | 20220028.008 | Blue Bay Site #8 | 1.222 | Carbon | 4 |  |  | 6156 | -22.543 | 69.3317 | 145.728 |  |  | 143.48 |  | 0 | 1.1471 |  | 0.010928 |  | 1.081 |  | BGD 44: 0.3; BGD 45: 0.4; BGD 46: -0.1 | 2022/04/19 15:05:31 | 78
04/19/22 | 17 | 20220028.008 | Blue Bay Site #8 | 1.222 |  | 5 |  |  | 5384 | -39.8 | 52.2489 | 100.593 |  |  | 99.076 |  | 1 | 1.1267 |  | 0.010735 |  | 1.0621 |  | BGD 44: 1.2; BGD 45: 1.4; BGD 46: 1.1 | 2022/04/19 15:05:31 | 78
04/19/22 | 17 | 20220028.008 | Blue Bay Site #8 | 1.222 |  | 6 |  |  | 5377 | -39.977 | 52.2693 | 100.642 |  |  | 99.124 |  | 0 | 1.1265 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.9; BGD 45: 2.2; BGD 46: 2.1 | 2022/04/19 15:05:31 | 78
04/19/22 | 18 | 20220028.009 | Blue Bay Site #9 | 1.394 |  | 1 | 5757 | 0.118 |  |  | 13.7004 | 105.727 |  |  |  | 104.966 | 0 |  | 0.7247 |  | 0.0036786 |  | 0.36651 | BGD 28: 28.7; BGD 29: 22.1; BGD 30: 924.2 | 2022/04/19 15:15:17 | 0 | 4171
04/19/22 | 18 | 20220028.009 | Blue Bay Site #9 | 1.394 |  | 2 | 5750 | 0 |  |  | 13.7004 | 105.727 |  |  |  | 104.966 | 1 |  | 0.72462 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.2; BGD 29: 21.6; BGD 30: 1009.6 | 2022/04/19 15:15:17 | 0 | 4166
04/19/22 | 18 | 20220028.009 | Blue Bay Site #9 | 1.394 | Nitrogen | 3 | 3689 | 10.915 |  |  | 9.3097 | 71.468 |  |  |  | 70.949 | 0 |  | 0.73252 |  | 0.0037183 |  | 0.37046 | BGD 28: 28.6; BGD 29: 21.9; BGD 30: 963.1 | 2022/04/19 15:15:17 | 0 | 2703
04/19/22 | 18 | 20220028.009 | Blue Bay Site #9 | 1.394 | Carbon | 4 |  |  | 7505 | -22.977 | 69.6427 | 179.161 |  |  | 176.4 |  | 0 | 1.1465 |  | 0.010923 |  | 1.0805 |  | BGD 44: 0.3; BGD 45: 0.4; BGD 46: -0.0 | 2022/04/19 15:15:17 | 78
04/19/22 | 18 | 20220028.009 | Blue Bay Site #9 | 1.394 |  | 5 |  |  | 5390 | -39.8 | 45.8388 | 100.692 |  |  | 99.174 |  | 1 | 1.1267 |  | 0.010735 |  | 1.0621 |  | BGD 44: 1.4; BGD 45: 1.7; BGD 46: 1.4 | 2022/04/19 15:15:17 | 78
04/19/22 | 18 | 20220028.009 | Blue Bay Site #9 | 1.394 |  | 6 |  |  | 5386 | -39.963 | 45.8733 | 100.786 |  |  | 99.266 |  | 0 | 1.1265 |  | 0.010733 |  | 1.0619 |  | BGD 44: 2.1; BGD 45: 2.4; BGD 46: 2.5 | 2022/04/19 15:15:17 | 78
04/19/22 | 19 | 36-UWSIF-UT Glut 1-20220028.16 | 36-UWSIF-UT Glut 1- | 0.825 |  | 1 | 5749 | 0.089 |  |  | 23.1418 | 105.691 |  |  |  | 104.931 | 0 |  | 0.72467 |  | 0.0036785 |  | 0.3665 | BGD 28: 28.8; BGD 29: 22.2; BGD 30: 932.7 | 2022/04/19 15:25:02 | 0 | 4165
04/19/22 | 19 | 36-UWSIF-UT Glut 1-20220028.16 | 36-UWSIF-UT Glut 1- | 0.825 |  | 2 | 5759 | 0 |  |  | 23.1877 | 105.904 |  |  |  | 105.142 | 1 |  | 0.72461 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.3; BGD 29: 21.8; BGD 30: 1017.9 | 2022/04/19 15:25:02 | 0 | 4173
04/19/22 | 19 | 36-UWSIF-UT Glut 1-20220028.16 | 36-UWSIF-UT Glut 1- | 0.825 | Nitrogen | 3 | 2219 | -1.897 |  |  | 9.5863 | 43.367 |  |  |  | 43.056 | 0 |  | 0.72323 |  | 0.0036712 |  | 0.36578 | BGD 28: 28.8; BGD 29: 22.0; BGD 30: 969.5 | 2022/04/19 15:25:02 | 0 | 1606
04/19/22 | 19 | 36-UWSIF-UT Glut 1-20220028.16 | 36-UWSIF-UT Glut 1- | 0.825 | Carbon | 4 |  |  | 4885 | -19.062 | 84.729 | 112.662 |  |  | 110.919 |  | 0 | 1.1509 |  | 0.010967 |  | 1.0848 |  | BGD 44: 0.3; BGD 45: 0.5; BGD 46: 0.1 | 2022/04/19 15:25:02 | 78
04/19/22 | 19 | 36-UWSIF-UT Glut 1-20220028.16 | 36-UWSIF-UT Glut 1- | 0.825 |  | 5 |  |  | 5389 | -39.8 | 77.4643 | 100.71 |  |  | 99.191 |  | 1 | 1.1267 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.9; BGD 45: 1.2; BGD 46: 0.7 | 2022/04/19 15:25:02 | 78
04/19/22 | 19 | 36-UWSIF-UT Glut 1-20220028.16 | 36-UWSIF-UT Glut 1- | 0.825 |  | 6 |  |  | 5388 | -40.021 | 77.5332 | 100.82 |  |  | 99.3 |  | 0 | 1.1265 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.7; BGD 45: 2.0; BGD 46: 1.9 | 2022/04/19 15:25:02 | 78
04/19/22 | 20 | 36-UWSIF-UT Glut 1-20220028.17 | 36-UWSIF-UT Glut 1- | 0.761 |  | 1 | 5754 | 0.12 |  |  | 25.1129 | 105.798 |  |  |  | 105.036 | 0 |  | 0.72467 |  | 0.0036786 |  | 0.36652 | BGD 28: 28.5; BGD 29: 21.9; BGD 30: 931.2 | 2022/04/19 15:34:47 | 0 | 4169
04/19/22 | 20 | 36-UWSIF-UT Glut 1-20220028.17 | 36-UWSIF-UT Glut 1- | 0.761 |  | 2 | 5769 | 0 |  |  | 25.1215 | 105.834 |  |  |  | 105.073 | 1 |  | 0.72459 |  | 0.0036782 |  | 0.36647 | BGD 28: 28.0; BGD 29: 21.5; BGD 30: 1015.0 | 2022/04/19 15:34:47 | 0 | 4180
04/19/22 | 20 | 36-UWSIF-UT Glut 1-20220028.17 | 36-UWSIF-UT Glut 1- | 0.761 | Nitrogen | 3 | 2025 | -1.913 |  |  | 9.5104 | 39.664 |  |  |  | 39.379 | 0 |  | 0.7232 |  | 0.0036712 |  | 0.36577 | BGD 28: 28.4; BGD 29: 21.7; BGD 30: 968.9 | 2022/04/19 15:34:47 | 0 | 1465
04/19/22 | 20 | 36-UWSIF-UT Glut 1-20220028.17 | 36-UWSIF-UT Glut 1- | 0.761 | Carbon | 4 |  |  | 4491 | -19.043 | 85.7075 | 103.276 |  |  | 101.678 |  | 0 | 1.151 |  | 0.010967 |  | 1.0848 |  | BGD 44: 0.3; BGD 45: 0.4; BGD 46: -0.1 | 2022/04/19 15:34:47 | 78
04/19/22 | 20 | 36-UWSIF-UT Glut 1-20220028.17 | 36-UWSIF-UT Glut 1- | 0.761 |  | 5 |  |  | 5394 | -39.8 | 84.0643 | 100.835 |  |  | 99.315 |  | 1 | 1.1267 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.8; BGD 45: 1.0; BGD 46: 0.6 | 2022/04/19 15:34:47 | 78
04/19/22 | 20 | 36-UWSIF-UT Glut 1-20220028.17 | 36-UWSIF-UT Glut 1- | 0.761 |  | 6 |  |  | 5387 | -40.015 | 84.0876 | 100.87 |  |  | 99.349 |  | 0 | 1.1265 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.6; BGD 45: 1.9; BGD 46: 1.8 | 2022/04/19 15:34:47 | 78
04/19/22 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 39-UWSIF-UW Glut 2- | 0.752 |  | 1 | 5756 | 0.099 |  |  | 25.3864 | 105.683 |  |  |  | 104.923 | 0 |  | 0.72465 |  | 0.0036786 |  | 0.36651 | BGD 28: 28.5; BGD 29: 21.8; BGD 30: 932.0 | 2022/04/19 15:44:33 | 0 | 4170
04/19/22 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 39-UWSIF-UW Glut 2- | 0.752 |  | 2 | 5764 | 0 |  |  | 25.4207 | 105.828 |  |  |  | 105.067 | 1 |  | 0.72458 |  | 0.0036782 |  | 0.36647 | BGD 28: 27.9; BGD 29: 21.4; BGD 30: 1017.8 | 2022/04/19 15:44:33 | 0 | 4176
04/19/22 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 39-UWSIF-UW Glut 2- | 0.752 | Nitrogen | 3 | 2198 | 30.707 |  |  | 10.4409 | 43.052 |  |  |  | 42.733 | 0 |  | 0.74683 |  | 0.0037911 |  | 0.37768 | BGD 28: 28.4; BGD 29: 21.6; BGD 30: 970.4 | 2022/04/19 15:44:33 | 0 | 1643
04/19/22 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 39-UWSIF-UW Glut 2- | 0.752 | Carbon | 4 |  |  | 4816 | 33.64 | 92.1144 | 111.37 |  |  | 109.586 |  | 0 | 1.2084 |  | 0.011556 |  | 1.1424 |  | BGD 44: 0.3; BGD 45: 0.4; BGD 46: -0.1 | 2022/04/19 15:44:33 | 78
04/19/22 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 39-UWSIF-UW Glut 2- | 0.752 |  | 5 |  |  | 5394 | -39.8 | 85.0009 | 100.734 |  |  | 99.215 |  | 1 | 1.1266 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 0.6 | 2022/04/19 15:44:33 | 78
04/19/22 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 39-UWSIF-UW Glut 2- | 0.752 |  | 6 |  |  | 5393 | -40.007 | 85.1209 | 100.909 |  |  | 99.388 |  | 0 | 1.1264 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.6; BGD 45: 2.0; BGD 46: 1.8 | 2022/04/19 15:44:33 | 78
04/19/22 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 39-UWSIF-UW Glut 2- | 0.741 |  | 1 | 5761 | 0.104 |  |  | 25.7658 | 105.694 |  |  |  | 104.933 | 0 |  | 0.72464 |  | 0.0036786 |  | 0.36651 | BGD 28: 28.4; BGD 29: 21.9; BGD 30: 935.0 | 2022/04/19 15:54:18 | 0 | 4174
04/19/22 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 39-UWSIF-UW Glut 2- | 0.741 |  | 2 | 5762 | 0 |  |  | 25.8055 | 105.859 |  |  |  | 105.098 | 1 |  | 0.72457 |  | 0.0036782 |  | 0.36647 | BGD 28: 27.9; BGD 29: 21.4; BGD 30: 1021.1 | 2022/04/19 15:54:18 | 0 | 4174
04/19/22 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 39-UWSIF-UW Glut 2- | 0.741 | Nitrogen | 3 | 2151 | 30.583 |  |  | 10.3553 | 42.068 |  |  |  | 41.756 | 0 |  | 0.74672 |  | 0.0037907 |  | 0.37764 | BGD 28: 28.4; BGD 29: 21.6; BGD 30: 974.5 | 2022/04/19 15:54:18 | 0 | 1607
04/19/22 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 39-UWSIF-UW Glut 2- | 0.741 | Carbon | 4 |  |  | 4756 | 33.544 | 92.3476 | 109.663 |  |  | 107.906 |  | 0 | 1.2083 |  | 0.011555 |  | 1.1423 |  | BGD 44: 0.2; BGD 45: 0.4; BGD 46: -0.1 | 2022/04/19 15:54:18 | 78
04/19/22 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 39-UWSIF-UW Glut 2- | 0.741 |  | 5 |  |  | 5396 | -39.8 | 86.3696 | 100.888 |  |  | 99.366 |  | 1 | 1.1266 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 0.6 | 2022/04/19 15:54:18 | 78
04/19/22 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 39-UWSIF-UW Glut 2- | 0.741 |  | 6 |  |  | 5386 | -40.018 | 86.3634 | 100.879 |  |  | 99.358 |  | 0 | 1.1264 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.6; BGD 45: 2.0; BGD 46: 1.8 | 2022/04/19 15:54:18 | 78
04/19/22 | 23 | 315-UWSIF-Chitin20220028.15 | 315-UWSIF-Chitin | 0.814 |  | 1 | 5756 | 0.106 |  |  | 23.4533 | 105.685 |  |  |  | 104.925 | 0 |  | 0.72463 |  | 0.0036786 |  | 0.36651 | BGD 28: 28.4; BGD 29: 21.9; BGD 30: 937.3 | 2022/04/19 16:04:03 | 0 | 4170
04/19/22 | 23 | 315-UWSIF-Chitin20220028.15 | 315-UWSIF-Chitin | 0.814 |  | 2 | 5758 | 0 |  |  | 23.455 | 105.693 |  |  |  | 104.933 | 1 |  | 0.72456 |  | 0.0036782 |  | 0.36647 | BGD 28: 27.9; BGD 29: 21.4; BGD 30: 1020.9 | 2022/04/19 16:04:03 | 0 | 4172
04/19/22 | 23 | 315-UWSIF-Chitin20220028.15 | 315-UWSIF-Chitin | 0.814 | Nitrogen | 3 | 1698 | 5.459 |  |  | 7.4983 | 33.418 |  |  |  | 33.177 | 0 |  | 0.72851 |  | 0.0036983 |  | 0.36846 | BGD 28: 28.4; BGD 29: 21.7; BGD 30: 971.7 | 2022/04/19 16:04:03 | 0 | 1238
04/19/22 | 23 | 315-UWSIF-Chitin20220028.15 | 315-UWSIF-Chitin | 0.814 | Carbon | 4 |  |  | 5064 | -9.756 | 89.4951 | 118.802 |  |  | 116.953 |  | 0 | 1.161 |  | 0.011071 |  | 1.095 |  | BGD 44: 0.2; BGD 45: 0.4; BGD 46: -0.1 | 2022/04/19 16:04:03 | 78
04/19/22 | 23 | 315-UWSIF-Chitin20220028.15 | 315-UWSIF-Chitin | 0.814 |  | 5 |  |  | 5395 | -39.8 | 78.5543 | 100.778 |  |  | 99.258 |  | 1 | 1.1266 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.9; BGD 45: 1.2; BGD 46: 0.8 | 2022/04/19 16:04:03 | 78
04/19/22 | 23 | 315-UWSIF-Chitin20220028.15 | 315-UWSIF-Chitin | 0.814 |  | 6 |  |  | 5397 | -40.013 | 78.671 | 100.962 |  |  | 99.44 |  | 0 | 1.1264 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.7; BGD 45: 2.0; BGD 46: 1.9 | 2022/04/19 16:04:03 | 78
04/19/22 | 24 | 315-UWSIF-Chitin20220028.16 | 315-UWSIF-Chitin | 0.816 |  | 1 | 5763 | 0.091 |  |  | 23.4568 | 105.966 |  |  |  | 105.203 | 0 |  | 0.72461 |  | 0.0036785 |  | 0.36651 | BGD 28: 28.5; BGD 29: 21.9; BGD 30: 939.1 | 2022/04/19 16:13:49 | 0 | 4175
04/19/22 | 24 | 315-UWSIF-Chitin20220028.16 | 315-UWSIF-Chitin | 0.816 |  | 2 | 5765 | 0 |  |  | 23.4721 | 106.036 |  |  |  | 105.273 | 1 |  | 0.72454 |  | 0.0036782 |  | 0.36647 | BGD 28: 27.9; BGD 29: 21.4; BGD 30: 1026.0 | 2022/04/19 16:13:49 | 0 | 4176
04/19/22 | 24 | 315-UWSIF-Chitin20220028.16 | 315-UWSIF-Chitin | 0.816 | Nitrogen | 3 | 1697 | 6.037 |  |  | 7.4455 | 33.264 |  |  |  | 33.023 | 0 |  | 0.72891 |  | 0.0037004 |  | 0.36868 | BGD 28: 28.4; BGD 29: 21.6; BGD 30: 976.2 | 2022/04/19 16:13:49 | 0 | 1238
04/19/22 | 24 | 315-UWSIF-Chitin20220028.16 | 315-UWSIF-Chitin | 0.816 | Carbon | 4 |  |  | 5053 | -9.714 | 88.7605 | 117.915 |  |  | 116.079 |  | 0 | 1.161 |  | 0.011072 |  | 1.095 |  | BGD 44: 0.2; BGD 45: 0.3; BGD 46: -0.1 | 2022/04/19 16:13:49 | 78
04/19/22 | 24 | 315-UWSIF-Chitin20220028.16 | 315-UWSIF-Chitin | 0.816 |  | 5 |  |  | 5398 | -39.8 | 78.351 | 100.761 |  |  | 99.242 |  | 1 | 1.1265 |  | 0.010735 |  | 1.0621 |  | BGD 44: 0.9; BGD 45: 1.1; BGD 46: 0.8 | 2022/04/19 16:13:49 | 78
04/19/22 | 24 | 315-UWSIF-Chitin20220028.16 | 315-UWSIF-Chitin | 0.816 |  | 6 |  |  | 5389 | -40.028 | 78.4646 | 100.941 |  |  | 99.419 |  | 0 | 1.1263 |  | 0.010733 |  | 1.0619 |  | BGD 44: 1.7; BGD 45: 2.0; BGD 46: 1.9 | 2022/04/19 16:13:49 | 78

## Sorted 1
Date | Row | Identifier 1 | Identifier 2 | Amount | Ampl 28 | d 15N/14N | Ampl 44 | d 13C/12C | %N | %C
04/19/22 | 1 | 36-UWSIF-UT Glut 1-20220028.11 | 36-UWSIF-UT Glut 1- | 0.811 | 2242 | -1.81 | 4884 | -18.943 | 9.5592 | 40.7809
04/19/22 | 2 | 36-UWSIF-UT Glut 1-20220028.12 | 36-UWSIF-UT Glut 1- | 0.841 | 2279 | -1.831 | 4995 | -19.032 | 9.5902 | 40.7906
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 | 1079 | -1.906 | 2435 | -18.898 | 9.5656 | 41.0825
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 | 2721 | -1.937 | 5956 | -18.989 | 9.4942 | 40.6598
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 | 8590 | -2.132 | 15966 | -19.047 | 9.5207 | 40.8165
04/19/22 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 39-UWSIF-UW Glut 2- | 0.759 | 2201 | 30.647 | 4870 | 33.692 | 10.2984 | 44.0575
04/19/22 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 39-UWSIF-UW Glut 2- | 0.809 | 2356 | 30.596 | 5179 | 33.873 | 10.262 | 43.8061
04/19/22 | 8 | 315-UWSIF-Chitin20220028.11 | 315-UWSIF-Chitin | 0.822 | 1722 | 6.008 | 5103 | -9.719 | 7.4845 | 43.2138
04/19/22 | 9 | 315-UWSIF-Chitin20220028.12 | 315-UWSIF-Chitin | 0.829 | 1765 | 6.364 | 5144 | -9.578 | 7.5927 | 43.0458
04/19/22 | 10 | 20220028.001 | Yellow Bay Site #1 | 1.807 | 4814 | 9.237 | 11861 | -21.62 | 9.2273 | 48.7806
04/19/22 | 11 | 20220028.002 | Yellow Bay Site #2 | 1.133 | 2995 | 9.512 | 7668 | -22.621 | 9.2626 | 47.4546
04/19/22 | 12 | 20220028.003 | Yellow Bay Site #3 | 0.856
04/19/22 | 13 | 20220028.004 | Midway Site #4 | 0.968 | 4679 | 10.848 | 10798 | -21.498 | 16.7921 | 81.3248
04/19/22 | 14 | 20220028.005 | Midway Site #5 | 0.852 | 2275 | 9.747 | 5914 | -22.454 | 9.4991 | 48.3931
04/19/22 | 15 | 20220028.006 | Midway Site #6 | 1.467 | 4112 | 7.991 | 9545 | -21.729 | 9.7798 | 46.852
04/19/22 | 16 | 20220028.007 | Blue Bay Site #7 | 1 | 1952 | 10.66 | 5881 | -22.368 | 7.023 | 41.22
04/19/22 | 17 | 20220028.008 | Blue Bay Site #8 | 1.222 | 2732 | 9.379 | 6156 | -22.543 | 7.9386 | 35.16
04/19/22 | 18 | 20220028.009 | Blue Bay Site #9 | 1.394 | 3689 | 10.915 | 7505 | -22.977 | 9.2767 | 37.7506
04/19/22 | 19 | 36-UWSIF-UT Glut 1-20220028.16 | 36-UWSIF-UT Glut 1- | 0.825 | 2219 | -1.897 | 4885 | -19.062 | 9.5572 | 40.5003
04/19/22 | 20 | 36-UWSIF-UT Glut 1-20220028.17 | 36-UWSIF-UT Glut 1- | 0.761 | 2025 | -1.913 | 4491 | -19.043 | 9.4871 | 40.3431
04/19/22 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 39-UWSIF-UW Glut 2- | 0.752 | 2198 | 30.707 | 4816 | 33.64 | 10.4098 | 43.9355
04/19/22 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 39-UWSIF-UW Glut 2- | 0.741 | 2151 | 30.583 | 4756 | 33.544 | 10.3258 | 43.9222
04/19/22 | 23 | 315-UWSIF-Chitin20220028.15 | 315-UWSIF-Chitin | 0.814 | 1698 | 5.459 | 5064 | -9.756 | 7.4912 | 43.2268
04/19/22 | 24 | 315-UWSIF-Chitin20220028.16 | 315-UWSIF-Chitin | 0.816 | 1697 | 6.037 | 5053 | -9.714 | 7.439 | 42.8068
Comment
2022-0028 Nelson Tray 1
CR 4/19/2022
He PSI: 1000
O2 PSI: 1800
M28: 28
M29: 21
M30: 408
Peak Center: 2.976
Instrument: Carlo Erba 1110 EA coupled to Thermo Delta V IRMS
Comments:

## % Calc 1
Date | Row | Identifier 1 | Identifier 2 | Amount | Area All N | Area All C | Known mg N | Known mg C | Calc mg N | Calc mg C | %N | %C
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 | 21.249 | 55.168 | 0.038937 | 0.16691 | 0.039123 | 0.16803 | 9.5656 | 41.0825
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 | 52.54 | 138.31 | 0.095581 | 0.40973 | 0.095322 | 0.40822 | 9.4942 | 40.6598
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 | 158.921 | 421.985 | 0.28636 | 1.2276 | 0.28638 | 1.2278 | 9.5207 | 40.8165
04/19/22 | 1 | 36-UWSIF-UT Glut 1-20220028.11 | 36-UWSIF-UT Glut 1- | 0.811 | 42.631 | 111.487 |  |  | 0.077525 | 0.33073 | 9.5592 | 40.7809
04/19/22 | 2 | 36-UWSIF-UT Glut 1-20220028.12 | 36-UWSIF-UT Glut 1- | 0.841 | 44.373 | 115.75 |  |  | 0.080654 | 0.34305 | 9.5902 | 40.7906
04/19/22 | 3 | 36-UWSIF-UT Glut 1-20220028.13 | 36-UWSIF-UT Glut 1- | 0.409 | 21.249 | 55.168 |  |  | 0.039123 | 0.16803 | 9.5656 | 41.0825
04/19/22 | 4 | 36-UWSIF-UT Glut 1-20220028.14 | 36-UWSIF-UT Glut 1- | 1.004 | 52.54 | 138.31 |  |  | 0.095322 | 0.40822 | 9.4942 | 40.6598
04/19/22 | 5 | 36-UWSIF-UT Glut 1-20220028.15 | 36-UWSIF-UT Glut 1- | 3.008 | 158.921 | 421.985 |  |  | 0.28638 | 1.2278 | 9.5207 | 40.8165
04/19/22 | 6 | 39-UWSIF-UW Glut 2-20220028.11 | 39-UWSIF-UW Glut 2- | 0.759 | 42.987 | 112.755 |  |  | 0.078165 | 0.3344 | 10.2984 | 44.0575
04/19/22 | 7 | 39-UWSIF-UW Glut 2-20220028.12 | 39-UWSIF-UW Glut 2- | 0.809 | 45.69 | 119.676 |  |  | 0.083019 | 0.35439 | 10.262 | 43.8061
04/19/22 | 8 | 315-UWSIF-Chitin20220028.11 | 315-UWSIF-Chitin | 0.822 | 33.721 | 119.962 |  |  | 0.061523 | 0.35522 | 7.4845 | 43.2138
04/19/22 | 9 | 315-UWSIF-Chitin20220028.12 | 315-UWSIF-Chitin | 0.829 | 34.512 | 120.527 |  |  | 0.062944 | 0.35685 | 7.5927 | 43.0458
04/19/22 | 10 | 20220028.001 | Yellow Bay Site #1 | 1.807 | 92.304 | 302.118 |  |  | 0.16674 | 0.88147 | 9.2273 | 48.7806
04/19/22 | 11 | 20220028.002 | Yellow Bay Site #2 | 1.133 | 57.898 | 183.113 |  |  | 0.10494 | 0.53766 | 9.2626 | 47.4546
04/19/22 | 12 | 20220028.003 | Yellow Bay Site #3 | 0.856
04/19/22 | 13 | 20220028.004 | Midway Site #4 | 0.968 | 89.971 | 269.497 |  |  | 0.16255 | 0.78722 | 16.7921 | 81.3248
04/19/22 | 14 | 20220028.005 | Midway Site #5 | 0.852 | 44.528 | 139.724 |  |  | 0.080932 | 0.41231 | 9.4991 | 48.3931
04/19/22 | 15 | 20220028.006 | Midway Site #6 | 1.467 | 79.348 | 234.916 |  |  | 0.14347 | 0.68732 | 9.7798 | 46.852
04/19/22 | 16 | 20220028.007 | Blue Bay Site #7 | 1 | 38.569 | 139.686 |  |  | 0.07023 | 0.4122 | 7.023 | 41.22
04/19/22 | 17 | 20220028.008 | Blue Bay Site #8 | 1.222 | 53.48 | 145.728 |  |  | 0.09701 | 0.42966 | 7.9386 | 35.16
04/19/22 | 18 | 20220028.009 | Blue Bay Site #9 | 1.394 | 71.468 | 179.161 |  |  | 0.12932 | 0.52624 | 9.2767 | 37.7506
04/19/22 | 19 | 36-UWSIF-UT Glut 1-20220028.16 | 36-UWSIF-UT Glut 1- | 0.825 | 43.367 | 112.662 |  |  | 0.078847 | 0.33413 | 9.5572 | 40.5003
04/19/22 | 20 | 36-UWSIF-UT Glut 1-20220028.17 | 36-UWSIF-UT Glut 1- | 0.761 | 39.664 | 103.276 |  |  | 0.072197 | 0.30701 | 9.4871 | 40.3431
04/19/22 | 21 | 39-UWSIF-UW Glut 2-20220028.13 | 39-UWSIF-UW Glut 2- | 0.752 | 43.052 | 111.37 |  |  | 0.078281 | 0.33039 | 10.4098 | 43.9355
04/19/22 | 22 | 39-UWSIF-UW Glut 2-20220028.14 | 39-UWSIF-UW Glut 2- | 0.741 | 42.068 | 109.663 |  |  | 0.076514 | 0.32546 | 10.3258 | 43.9222
04/19/22 | 23 | 315-UWSIF-Chitin20220028.15 | 315-UWSIF-Chitin | 0.814 | 33.418 | 118.802 |  |  | 0.060979 | 0.35187 | 7.4912 | 43.2268
04/19/22 | 24 | 315-UWSIF-Chitin20220028.16 | 315-UWSIF-Chitin | 0.816 | 33.264 | 117.915 |  |  | 0.060702 | 0.3493 | 7.439 | 42.8068